## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2023/231019.seq/629.smp/MSDevice.xraw
Channel | TIC
Injection Information: |  |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Data Vault | ChromeleonLocal |  |  | 95 | 6517585.5 | 50-200% of mass 174 | 95/174 | 110.1495 | 50 | 200 | TRUE
Injection | H2O-V2 |  |  | 96 | 437327.25 | 5 to 9% of m/z 95 | 96/95 | 6.71 | 5 | 9 | TRUE
Injection Number | 5 |  |  | 173 | 7160.7881 | <2% of m/z 174 | 173/174 | 0.12102 | 0 | 2 | TRUE
Position | 0 |  |  | 174 | 5917036.5 | 50-200% of mass 95 | 174/95 | 90.7857 | 50 | 200 | TRUE
Comment |  |  |  | 175 | 400648.4375 | 5 to 9% of m/z 174 | 175/174 | 6.7711 | 5 | 9 | TRUE
Processing Method | VOC-9 |  |  | 176 | 5702602 | 95 to 105% of m/z 174 | 176/174 | 96.376 | 95 | 105 | TRUE
Instrument Method | VOC-25 |  |  | 177 | 340678.0625 | 5 to 10% of m/z 176 | 177/176 | 5.9741 | 5 | 10 | TRUE
Type | Unknown
Status | Finished
Injection Date | 2023-10-19
Injection Time | 11:42
Injection Volume (µl) | 20 | *N/A for P&T
Dilution Factor | 1
Weight | 1
Raw Data Information:
Data Points | 240
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9808
Retention Time (min) | 9.8334
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3285
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0349
Base Peak Intensity | 6517585.5
Total Ion Count | 28050568.8947
Raw Data: |  | Rounded:
Mass (m/z) | Value (counts) | Mass (m/z) | Value (counts)
35 | 480.0278 | 35 | 480.0278
36 | 33035.3008 | 36 | 33035.3008
37.1 | 187116.2188 | 37 | 187116.2188
38 | 147679.5312 | 38 | 147679.5312
39.1 | 65492.8789 | 39 | 65492.8789
40 | 45675.1875 | 40 | 45675.1875
41.1 | 4524.0684 | 41 | 4524.0684
42 | 2417.8284 | 42 | 2417.8284
43 | 3117.2642 | 43 | 3117.2642
44 | 68298.1875 | 44 | 68298.1875
45.1 | 35913.7422 | 45 | 35913.7422
46.1 | 2918.6831 | 46 | 2918.6831
47 | 73790.25 | 47 | 73790.25
47.9 | 26332.9902 | 48 | 26332.9902
49 | 159434.5781 | 49 | 159434.5781
50 | 822107.125 | 50 | 822107.125
51 | 237777.3594 | 51 | 237777.3594
52.1 | 9044.1768 | 52 | 9044.1768
53.1 | 955.5417 | 53 | 955.5417
54.2 | 613.6002 | 54 | 613.6002
55 | 7116.1084 | 55 | 7116.1084
56 | 41917.8945 | 56 | 41917.8945
57 | 90379.5938 | 57 | 90379.5938
58.1 | 3593.2354 | 58 | 3593.2354
58.7 | 473.1745 | 59 | 473.1745
60 | 33261.2852 | 60 | 33261.2852
61 | 172754.875 | 61 | 172754.875
62 | 161497.9844 | 62 | 161497.9844
63 | 116070.9375 | 63 | 116070.9375
64 | 11199.1895 | 64 | 11199.1895
65.1 | 3119.1594 | 65 | 3119.1594
65.9 | 1168.3223 | 66 | 1168.3223
67 | 8278.4297 | 67 | 8278.4297
68 | 423395.4375 | 68 | 423395.4375
69 | 373906.7188 | 69 | 373906.7188
70 | 26171.8359 | 70 | 26171.8359
71 | 779.4811 | 71 | 779.4811
72 | 18254.1621 | 72 | 18254.1621
73 | 163934.0469 | 73 | 163934.0469
74 | 653876.0625 | 74 | 653876.0625
75 | 2447144.75 | 75 | 2447144.75
76 | 195019.0938 | 76 | 195019.0938
77.1 | 28998.0234 | 77 | 28998.0234
78 | 21272.2324 | 78 | 21272.2324
78.9 | 36814.5625 | 79 | 36814.5625
80 | 13382.3164 | 80 | 13382.3164
80.9 | 40136.9844 | 81 | 40136.9844
81.9 | 7024.5503 | 82 | 7024.5503
83 | 1918.9893 | 83 | 1918.9893
83.5 | 338.3937 | 84 | 338.3937
84.2 | 449.4079 | 84 | 449.4079
85.2 | 75.5131 | 85 | 75.5131
86.1 | 8658.0518 | 86 | 8658.0518
87 | 251434.3281 | 87 | 251434.3281
88 | 256694.5781 | 88 | 256694.5781
89.9 | 127.1239 | 90 | 127.1239
90.9 | 12276.0684 | 91 | 12276.0684
92 | 115657.4531 | 92 | 115657.4531
93 | 191220.6094 | 93 | 191220.6094
94 | 523744.1875 | 94 | 523744.1875
95 | 6517585.5 | 95 | 6517585.5
96 | 437327.25 | 96 | 437327.25
97 | 12590.5059 | 97 | 12590.5059
98.1 | 720.7899 | 98 | 720.7899
99 | 0.2024 | 99 | 0.2024
100.1 | 184.7249 | 100 | 184.7249
101 | 108.9153 | 101 | 108.9153
102.6 | 644.9257 | 103 | 644.9257
103.3 | 427.2199 | 103 | 427.2199
103.9 | 11819.6777 | 104 | 11819.6777
104.9 | 6146.3735 | 105 | 6146.3735
105.9 | 13143.6289 | 106 | 13143.6289
107 | 1649.2659 | 107 | 1649.2659
108.2 | 409.7735 | 108 | 409.7735
109.2 | 182.6485 | 109 | 182.6485
109.9 | 591.1666 | 110 | 591.1666
110.7 | 1706.8429 | 111 | 1706.8429
111.4 | 573.3417 | 111 | 573.3417
111.9 | 1268.3096 | 112 | 1268.3096
112.8 | 833.399 | 113 | 833.399
113.5 | 96.2228 | 114 | 96.2228
114.9 | 3757.7327 | 115 | 3757.7327
115.9 | 9737.127 | 116 | 9737.127
117 | 22471.0391 | 117 | 22471.0391
118 | 8939.5332 | 118 | 8939.5332
118.9 | 17564.1973 | 119 | 17564.1973
119.8 | 318.556 | 120 | 318.556
120.4 | 452.1699 | 120 | 452.1699
121.5 | 527.6892 | 122 | 527.6892
122.2 | 308.5503 | 122 | 308.5503
122.9 | 503.6487 | 123 | 503.6487
124 | 1548.3159 | 124 | 1548.3159
125 | 1315.3225 | 125 | 1315.3225
126 | 1473.7032 | 126 | 1473.7032
127 | 1433.3472 | 127 | 1433.3472
127.9 | 14348.9219 | 128 | 14348.9219
129 | 5507.3218 | 129 | 5507.3218
129.9 | 13323.6797 | 130 | 13323.6797
130.9 | 6009.6216 | 131 | 6009.6216
131.9 | 769.0589 | 132 | 769.0589
133.5 | 906.201 | 134 | 906.201
134.9 | 5065.1094 | 135 | 5065.1094
136.1 | 401.927 | 136 | 401.927
136.9 | 5213.8726 | 137 | 5213.8726
138 | 315.7957 | 138 | 315.7957
138.8 | 1403.8077 | 139 | 1403.8077
140 | 1438.2109 | 140 | 1438.2109
140.9 | 29660.5625 | 141 | 29660.5625
141.8 | 3933.2498 | 142 | 3933.2498
142.9 | 30352.7676 | 143 | 30352.7676
143.8 | 1378.2764 | 144 | 1378.2764
144.8 | 2626.7803 | 145 | 2626.7803
145.9 | 5778.7529 | 146 | 5778.7529
146.9 | 3504.8994 | 147 | 3504.8994
147.9 | 10000.0996 | 148 | 10000.0996
148.9 | 2872.8533 | 149 | 2872.8533
149.9 | 2438.6873 | 150 | 2438.6873
150.8 | 159.0164 | 151 | 159.0164
151.9 | 1548.2887 | 152 | 1548.2887
153 | 3880.781 | 153 | 3880.781
154 | 3114.5215 | 154 | 3114.5215
154.9 | 11946.2139 | 155 | 11946.2139
156 | 1434.3798 | 156 | 1434.3798
156.9 | 7566.9507 | 157 | 7566.9507
157.9 | 1055.2666 | 158 | 1055.2666
159 | 4537.2202 | 159 | 4537.2202
159.9 | 415.6466 | 160 | 415.6466
160.9 | 4150.5508 | 161 | 4150.5508
161.4 | 62.0053 | 161 | 62.0053
162 | 78.9946 | 162 | 78.9946
163.2 | 551.7443 | 163 | 551.7443
165.1 | 314.7862 | 165 | 314.7862
166.2 | 126.2613 | 166 | 126.2613
166.9 | 289.2597 | 167 | 289.2597
167.7 | 143.0075 | 168 | 143.0075
168.4 | 266.881 | 168 | 266.881
168.9 | 505.1684 | 169 | 505.1684
169.6 | 330.2804 | 170 | 330.2804
170.2 | 542.3459 | 170 | 542.3459
171.3 | 1519.182 | 171 | 1519.182
172 | 5233.8535 | 172 | 5233.8535
172.6 | 7160.7881 | 173 | 7160.7881
173.9 | 5917036.5 | 174 | 5917036.5
175 | 400648.4375 | 175 | 400648.4375
175.9 | 5702602 | 176 | 5702602
177 | 340678.0625 | 177 | 340678.0625
178 | 10713.6338 | 178 | 10713.6338
179 | 85.395 | 179 | 85.395
179.6 | 0 | 180 | 0
180.2 | 274.4447 | 180 | 274.4447
180.9 | 0 | 181 | 0
181.5 | 1.8e-05 | 182 | 1.8e-05
182.3 | 0.0004 | 182 | 0.0004
182.9 | 0 | 183 | 0
183.5 | 0 | 184 | 0
184.2 | 0 | 184 | 0
185.6 | 0 | 186 | 0
186.2 | 42.8578 | 186 | 42.8578
187.2 | 106.7078 | 187 | 106.7078
188.2 | 0 | 188 | 0
188.7 | 86.604 | 189 | 86.604
190 | 121.9875 | 190 | 121.9875
190.7 | 508.6107 | 191 | 508.6107
192 | 177.8241 | 192 | 177.8241
192.8 | 76.4303 | 193 | 76.4303
193.7 | 2.3276 | 194 | 2.3276
194.5 | 0 | 195 | 0
195.9 | 84.3155 | 196 | 84.3155
196.4 | 0.005895 | 196 | 0.005895
197.2 | 74.3841 | 197 | 74.3841
198.1 | 0 | 198 | 0
198.7 | 0 | 199 | 0
199.6 | 0 | 200 | 0
200.8 | 0.004445 | 201 | 0.004445
202 | 40.7819 | 202 | 40.7819
203.1 | 3e-05 | 203 | 3e-05
203.7 | 0.000289 | 204 | 0.000289
204.8 | 4.8916 | 205 | 4.8916
205.4 | 0 | 205 | 0
207.1 | 2336.4553 | 207 | 2336.4553
208.2 | 283.0523 | 208 | 283.0523
209.1 | 151.5377 | 209 | 151.5377
209.8 | 164.2599 | 210 | 164.2599
210.6 | 4.2e-05 | 211 | 4.2e-05
211.4 | 0 | 211 | 0
212.1 | 74.2264 | 212 | 74.2264
213.2 | 0 | 213 | 0
213.7 | 0 | 214 | 0
214.8 | 0 | 215 | 0
215.4 | 0 | 215 | 0
216 | 0 | 216 | 0
216.9 | 0 | 217 | 0
218.1 | 91.7804 | 218 | 91.7804
218.7 | 0.013261 | 219 | 0.013261
219.5 | 0 | 220 | 0
220.1 | 0.000265 | 220 | 0.000265
221.3 | 102.2976 | 221 | 102.2976
222 | 200.0225 | 222 | 200.0225
223 | 0.00021 | 223 | 0.00021
223.9 | 0 | 224 | 0
225.1 | 0 | 225 | 0
226.1 | 8.8e-05 | 226 | 8.8e-05
226.7 | 0 | 227 | 0
227.6 | 0 | 228 | 0
228.5 | 0 | 229 | 0
229.2 | 0 | 229 | 0
229.9 | 0 | 230 | 0
231.3 | 231.7047 | 231 | 231.7047
232.1 | 107.9983 | 232 | 107.9983
233 | 109.2173 | 233 | 109.2173
234 | 4.9472 | 234 | 4.9472
234.6 | 0 | 235 | 0
235.3 | 38.1866 | 235 | 38.1866
236.2 | 0.033007 | 236 | 0.033007
237.1 | 0 | 237 | 0
238.8 | 0.018339 | 239 | 0.018339
239.5 | 0.000272 | 240 | 0.000272
240.5 | 0 | 241 | 0
241.1 | 0 | 241 | 0
241.8 | 0.2696 | 242 | 0.2696
242.3 | 3.418 | 242 | 3.418
243.3 | 0 | 243 | 0
244.8 | 0 | 245 | 0
245.7 | 2.6911 | 246 | 2.6911
246.3 | 0 | 246 | 0
247.3 | 1e-05 | 247 | 1e-05
248 | 2.8e-05 | 248 | 2.8e-05
249.1 | 52.9047 | 249 | 52.9047
250.3 | 4.1e-05 | 250 | 4.1e-05
251.2 | 315.1255 | 251 | 315.1255
252.1 | 0 | 252 | 0
253.5 | 3.3149 | 254 | 3.3149
254.5 | 0 | 255 | 0
255.1 | 0 | 255 | 0
256.1 | 0 | 256 | 0
256.7 | 0 | 257 | 0
257.7 | 0 | 258 | 0
258.8 | 2.7447 | 259 | 2.7447
259.3 | 0 | 259 | 0
260 | 121.3069
258.7 | 4e-06
259.9 | 229.5571
257.3 | 0.004411
258.3 | 181.4352
259 | 0.043369
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[8 empty] | TV | [7 empty] | Actual | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
Name | RT | Window | Quant | Conf1 | Conf2 | IS |  | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | MRL | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | RSD<20% | R^2>0.99 | Total Analytes | 78
Chloromethane (methyl chloride) | 1.464 | 0.050 AN | 50 | 52 | 49 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8181 | 1.1589 | 1.8917 | 4.5922 | 19.7375 | 100.0735 | #DIV/0! | 115.89 | 94.585 | 91.844 | 98.6875 | 100.0735 | Chloromethane (methyl chloride) | 1.46 | Lin, Avg | 5 | 1.0335 | 0.99996 | TRUE | TRUE | Failed | 0 | 0
Chloroethene (vinyl chloride) | 1.553 | 0.050 AN | 62 | 64 | 61 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6227 | 0.9769 | 1.7516 | 4.4724 | 20.0012 | 100.0307 | 124.54 | 97.69 | 87.58 | 89.448 | 100.006 | 100.0307 | Chloroethene (vinyl chloride) | 1.55 | Lin, Avg | 6 | 1.2532 | 0.99995 | TRUE | TRUE | Allowance | 7.8
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7536 | 1 | 2 | 5 | 20 | 100 | #DIV/0! | 100 | 100 | 100 | 100 | 100 | Bromomethane (methyl bromide) | 1.83 | P-P | 5 | 0 | 1 | TRUE | TRUE | Pass? | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6247 | 0.9442 | 1.6725 | 4.5078 | 19.2934 | 100.1724 | 124.94 | 94.42 | 83.625 | 90.156 | 96.467 | 100.1724 | Chloroethane (ethyl chloride) | 1.94 | Lin, Avg | 6 | 1.9978 | 0.99988 | TRUE | TRUE
Trichlorofluoromethane | 2.182 | 0.050 AN | 101 | 103 | 105 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.513 | 0.8815 | 1.6805 | 4.2772 | 19.7917 | 100.0853 | 102.6 | 88.15 | 84.025 | 85.544 | 98.9585 | 100.0853 | Trichlorofluoromethane | 2.19 | Lin, Avg | 6 | 1.751 | 0.99991 | TRUE | TRUE
Diethyl ether | 2.505 | 0.080 AN | 59 | 74 | 45 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8618 | 1.0197 | 1.826 | 4.9486 | 19.5703 | 100.0918 | #DIV/0! | 101.97 | 91.3 | 98.972 | 97.8515 | 100.0918 | Diethyl ether | 2.5 | Lin, Avg | 5 | 0.9332 | 0.99997 | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6132 | 0.9245 | 1.442 | 4.4148 | 19.7851 | 100.0836 | 122.64 | 92.45 | 72.1 | 88.296 | 98.9255 | 100.0836 | 1,1-Dichloroethene | 2.73 | Lin, Avg | 6 | 1.7959 | 0.99991 | TRUE | TRUE
Acetone | 2.852 | 0.088 AN | 43 | 58 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1.8 | 3.6 | 9 | 36 | 180 | ppb | 1.8 | 1.8516 | 1.651 | 3.0661 | 8.0419 | 34.6498 | 180.3301 | #DIV/0! | 91.7222 | 85.1694 | 89.3544 | 96.2494 | 100.1834 | Acetone | 2.82 | Lin, Avg | 5 | 1.9505 | 0.99987 | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | Internal Pentafluorobenzene [IS1] | P-P | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5 | 1 | 2 | 5 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Iodomethane | 2.88 | P-P | 6 | 0 | 1 | TRUE | TRUE
Carbon disulfide | 2.961 | 0.050 AN | 76 | 78 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6427 | 0.976 | 1.7319 | 4.5063 | 19.9561 | 100.0383 | 128.54 | 97.6 | 86.595 | 90.126 | 99.7805 | 100.0383 | Carbon disulfide | 2.95 | Lin, Avg | 6 | 1.2236 | 0.99996 | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5826 | 0.856 | 1.5414 | 4.0683 | 18.2517 | 100.4064 | 116.52 | 85.6 | 77.07 | 81.366 | 91.2585 | 100.4064 | 3-Chloropropene (allyl chloride) | 3.19 | Lin, Avg | 6 | 4.4405 | 0.99945 | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9511 | 1.138 | 1.9638 | 4.9762 | 19.9538 | 100.0098 | #DIV/0! | 113.8 | 98.19 | 99.524 | 99.769 | 100.0098 | Methylene chloride (DCM) | 3.36 | Lin, Avg | 5 | 0.2971 | 1 | TRUE | TRUE
trans-1,2-Dichloroethene | 3.678 | 0.100 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6315 | 0.8919 | 1.6547 | 4.2235 | 18.8901 | 100.2681 | 126.3 | 89.19 | 82.735 | 84.47 | 94.4505 | 100.2681 | trans-1,2-Dichloroethene | 3.68 | Lin, Avg | 6 | 3.0393 | 0.99974 | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9056 | 0.9534 | 1.7822 | 4.4929 | 18.5067 | 100.3288 | #DIV/0! | 95.34 | 89.11 | 89.858 | 92.5335 | 100.3288 | Methyl tert-butyl ether (MTBE) | 3.69 | Lin, Avg | 5 | 3.2252 | 0.99964 | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7655 | 1.0124 | 1.779 | 4.6233 | 19.8728 | 100.0486 | #DIV/0! | 101.24 | 88.95 | 92.466 | 99.364 | 100.0486 | 1,1-Dichloroethane | 4.19 | Lin, Avg | 5 | 0.8985 | 0.99997 | TRUE | TRUE
2,2-Dichloropropane | 4.819 | 0.060 AN | 77 | 41 | 79 | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5371 | 0.8169 | 1.4374 | 3.7934 | 17.8048 | 100.5123 | 107.42 | 81.69 | 71.87 | 75.868 | 89.024 | 100.5123 | 2,2-Dichloropropane | 4.81 | Lin, Avg | 6 | 5.6386 | 0.99912 | TRUE | TRUE
cis-1,2-Dichloroethene | 4.823 | 0.060 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7303 | 0.9357 | 1.6635 | 4.3072 | 18.6963 | 100.3027 | #DIV/0! | 93.57 | 83.175 | 86.144 | 93.4815 | 100.3027 | cis-1,2-Dichloroethene | 4.82 | Lin, Avg | 5 | 3.0687 | 0.99967 | TRUE | TRUE
2-Butanone (MEK) | 4.851 | 0.060 AN | 43 | 72 | 57 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1.8 | 3.6 | 9 | 36 | 180 | ppb | 1.8 | 1.5522 | 1.6182 | 2.8322 | 8.3443 | 34.7511 | 180.2997 | #DIV/0! | 89.9 | 78.6722 | 92.7144 | 96.5308 | 100.1665 | 2-Butanone (MEK) | 4.84 | Lin, Avg | 5 | 1.8037 | 0.99988 | TRUE | TRUE
Methyl acrylate | 4.939 | 0.060 AN | 55 | 85 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7454 | 0.8232 | 1.4302 | 4.3353 | 18.808 | 100.2848 | #DIV/0! | 82.32 | 71.51 | 86.706 | 94.04 | 100.2848 | Methyl acrylate | 4.94 | Lin, Avg | 5 | 3.0166 | 0.99968 | TRUE | TRUE
Bromochloromethane | 5.073 | 0.060 AN | 49 | 130 | 128 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9276 | 1.0091 | 1.77 | 4.6552 | 18.9539 | 100.231 | #DIV/0! | 100.91 | 88.5 | 93.104 | 94.7695 | 100.231 | Methacrylonitrile | 5.06 | Lin, Avg | 5 | 4.4117 | 0.99933 | TRUE | TRUE
Methacrylonitrile | 5.076 | 0.060 AN | 67 | 52 | 40 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7563 | 0.6713 | 1.4771 | 4.3604 | 18.0301 | 100.4397 | #DIV/0! | 67.13 | 73.855 | 87.208 | 90.1505 | 100.4397 | Bromochloromethane | 5.07 | Lin, Avg | 5 | 2.268 | 0.99982 | TRUE | TRUE
Tetrahydrofuran | 5.09 | 0.060 AN | 42 | 72 | 71 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8376 | 0.812 | 1.7061 | 4.7701 | 18.8566 | 100.2479 | #DIV/0! | 81.2 | 85.305 | 95.402 | 94.283 | 100.2479 | Tetrahydrofuran | 5.08 | Lin, Avg | 5 | 2.4573 | 0.99979 | TRUE | TRUE
Trichloromethane (chloroform) | 5.198 | 0.060 AN | 83 | 85 | 47 | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8238 | 1.0432 | 1.898 | 4.8213 | 19.8036 | 100.0498 | #DIV/0! | 104.32 | 94.9 | 96.426 | 99.018 | 100.0498 | Trichloromethane (chloroform) | 5.2 | Lin, Avg | 5 | 0.572 | 0.99999 | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5999 | 0.9064 | 1.7452 | 4.671 | 20.677 | 99.8866 | 119.98 | 90.64 | 87.26 | 93.42 | 103.385 | 99.8866 | 1,1,1-Trichloroethane | 5.34 | Lin, Avg | 6 | 1.7006 | 0.99991 | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 19.7372 | 19.5638 | 19.9286 | 19.9486 | 20.3044 | 20.5174 | 98.686 | 97.819 | 99.643 | 99.743 | 101.522 | 102.587 | Dibromofluoromethane [SS1] | 5.36 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Pentafluorobenzene [IS1] | 5.43 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.476 | 0.040 AN | 56 | 41 | 43 | Internal 1,4-Difluorobenzene [IS2] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.583 | 0.8497 | 1.5665 | 4.2049 | 18.3812 | 100.3733 | 116.6 | 84.97 | 78.325 | 84.098 | 91.906 | 100.3733 | 1-Chlorobutane (butyl chloride) | 5.48 | Lin, Avg | 6 | 4.0471 | 0.99954 | TRUE | TRUE
Carbon tetrachloride | 5.493 | 0.040 AN | 119 | 121 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5111 | 0.7362 | 1.4539 | 3.8412 | 18.242 | 100.423 | 102.22 | 73.62 | 72.695 | 76.824 | 91.21 | 100.423 | Carbon tetrachloride | 5.49 | Lin, Avg | 6 | 4.7825 | 0.99936 | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | Internal 1,4-Difluorobenzene [IS2] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6185 | 0.9061 | 1.6194 | 4.3669 | 19.3162 | 100.1764 | 123.7 | 90.61 | 80.97 | 87.338 | 96.581 | 100.1764 | 1,1-Dichloropropene | 5.51 | Lin, Avg | 6 | 2.1825 | 0.99986 | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7148 | 0.9295 | 1.762 | 4.6472 | 20.1736 | 99.9873 | 142.96 | 92.95 | 88.1 | 92.944 | 100.868 | 99.9873 | Benzene | 5.7 | Lin, Avg | 6 | 1.0722 | 0.99997 | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.1169 | 1.2162 | 2.1882 | 5.8832 | 22.5145 | 99.447 | #DIV/0! | 121.62 | 109.41 | 117.664 | 112.5725 | 99.447 | 1,2-Dichloroethane | 5.77 | Lin, Avg | 5 | 5.2132 | 0.99893 | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.167 | 0.080 AN | 114 | 88 | 63 | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Difluorobenzene [IS2] | 6.17 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7968 | 1.064 | 1.9598 | 5.128 | 21.1554 | 99.7627 | #DIV/0! | 106.4 | 97.99 | 102.56 | 105.777 | 99.7627 | Trichloroethene | 6.39 | Lin, Avg | 5 | 2.3027 | 0.9998 | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8664 | 1.0363 | 1.8868 | 5.0194 | 20.3628 | 99.9284 | #DIV/0! | 103.63 | 94.34 | 100.388 | 101.814 | 99.9284 | 1,2-Dichloropropane | 6.64 | Lin, Avg | 5 | 0.7582 | 0.99998 | TRUE | TRUE
Dibromomethane | 6.715 | 0.060 AN | 174 | 93 | 95 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.0878 | 1.0759 | 1.9635 | 5.3734 | 20.431 | 99.8951 | #DIV/0! | 107.59 | 98.175 | 107.468 | 102.155 | 99.8951 | Dibromomethane | 6.72 | Lin, Avg | 5 | 1.1377 | 0.99995 | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8858 | 0.8398 | 1.6071 | 4.485 | 18.2894 | 100.3773 | #DIV/0! | 83.98 | 80.355 | 89.7 | 91.447 | 100.3773 | Methyl methacrylate | 6.74 | Lin, Avg | 5 | 3.7311 | 0.99952 | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.863 | 1.0001 | 1.8061 | 4.9065 | 20.4297 | 99.9226 | #DIV/0! | 100.01 | 90.305 | 98.13 | 102.1485 | 99.9226 | Bromodichloromethane | 6.92 | Lin, Avg | 5 | 0.9503 | 0.99997 | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9667 | 1.0897 | 1.8088 | 4.5342 | 18.7203 | 100.2822 | #DIV/0! | 108.97 | 90.44 | 90.684 | 93.6015 | 100.2822 | 2-Nitropropane | 7.14 | Lin, Avg | 5 | 2.7815 | 0.99973 | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7066 | 0.8345 | 1.5167 | 4.189 | 17.7678 | 100.4973 | 141.32 | 83.45 | 75.835 | 83.78 | 88.839 | 100.4973 | cis-1,3-Dichloropropene | 7.36 | Lin, Avg | 6 | 5.3196 | 0.99921 | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.518 | 0.049 AN | 43 | 58 | 41 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1.8 | 3.6 | 9 | 36 | 180 | ppb | 1.8 | 1.662 | 1.7037 | 3.1872 | 8.564 | 34.4038 | 180.3503 | #DIV/0! | 94.65 | 88.5333 | 95.1556 | 95.5661 | 100.1946 | 4-Methyl-2-pentanone (MIBK) | 7.52 | Lin, Avg | 5 | 1.9101 | 0.99987 | TRUE | TRUE
Toluene-d8 [SS2] | 7.607 | 0.058 AN | 98 | 100 | 70 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 19.5749 | 19.6155 | 20.0527 | 20.0628 | 20.435 | 20.259 | 97.8745 | 98.0775 | 100.2635 | 100.314 | 102.175 | 101.295 | Toluene-d8 [SS2] | 7.61 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Toluene | 7.67 | 0.047 AN | 91 | 92 | 65 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6825 | 0.9436 | 1.7434 | 4.6028 | 20.1798 | 99.9887 | 136.5 | 94.36 | 87.17 | 92.056 | 100.899 | 99.9887 | Toluene | 7.67 | Lin, Avg | 6 | 1.1326 | 0.99996 | TRUE | TRUE
trans-1,3-Dichloropropene | 7.92 | 0.060 AN | 75 | 39 | 77 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9203 | 1 | 2 | 5 | 20 | 100 | #DIV/0! | 100 | 100 | 100 | 100 | 100 | trans-1,3-Dichloropropene | 7.92 | P-P | 5 | 0 | 1 | TRUE | TRUE
Ethyl methacrylate | 7.99 | 0.060 AN | 69 | 41 | 99 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8727 | 0.9118 | 1.5823 | 4.3625 | 17.9893 | 100.4433 | #DIV/0! | 91.18 | 79.115 | 87.25 | 89.9465 | 100.4433 | Ethyl methacrylate | 7.99 | Lin, Avg | 5 | 4.3845 | 0.99934 | TRUE | TRUE
1,1,2-Trichloroethane | 8.098 | 0.060 AN | 97 | 83 | 99 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.0099 | 1.091 | 1.9384 | 5.2513 | 20.9797 | 99.7918 | #DIV/0! | 109.1 | 96.92 | 105.026 | 104.8985 | 99.7918 | 1,1,2-Trichloroethane | 8.1 | Lin, Avg | 5 | 2.0116 | 0.99985 | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9048 | 1.215 | 2.2003 | 5.8044 | 23.1312 | 99.3274 | #DIV/0! | 121.5 | 110.015 | 116.088 | 115.656 | 99.3274 | Tetrachloroethene | 8.15 | Lin, Avg | 5 | 6.294 | 0.99842 | TRUE | TRUE
1,3-Dichloropropane | 8.24 | 0.060 AN | 76 | 41 | 78 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9909 | 1.0533 | 1.946 | 5.1326 | 19.841 | 100.0257 | #DIV/0! | 105.33 | 97.3 | 102.652 | 99.205 | 100.0257 | 1,3-Dichloropropane | 8.24 | Lin, Avg | 5 | 0.4336 | 0.99999 | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1.8 | 3.6 | 9 | 36 | 180 | ppb | 1.8 | 1.7229 | 1.661 | 3.1633 | 9.0459 | 34.4965 | 180.3085 | #DIV/0! | 92.2778 | 87.8694 | 100.51 | 95.8236 | 100.1714 | 2-Hexanone | 8.31 | Lin, Avg | 5 | 1.7519 | 0.99989 | TRUE | TRUE
Dibromochloromethane | 8.418 | 0.050 AN | 129 | 127 | 131 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8304 | 0.914 | 1.658 | 4.4716 | 18.9845 | 100.2372 | #DIV/0! | 91.4 | 82.9 | 89.432 | 94.9225 | 100.2372 | Dibromochloromethane | 8.43 | Lin, Avg | 5 | 2.4177 | 0.99979 | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.517 | 0.049 AN | 107 | 109 | 93 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.0689 | 1.1128 | 2.0033 | 5.4255 | 21.2386 | 99.7298 | #DIV/0! | 111.28 | 100.165 | 108.51 | 106.193 | 99.7298 | 1,2-Dibromoethane (EDB) | 8.52 | Lin, Avg | 5 | 2.5906 | 0.99975 | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.915 | 0.040 AF | 117 | 82 | 52 | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Chlorobenzene-d5 [IS3] | 8.91 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Chlorobenzene | 8.937 | 0.040 AN | 112 | 77 | 114 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7474 | 0.9731 | 1.7681 | 4.7203 | 20.0022 | 100.0172 | 149.48 | 97.31 | 88.405 | 94.406 | 100.011 | 100.0172 | Chlorobenzene | 8.93 | Lin, Avg | 6 | 0.9223 | 0.99997 | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.01 | 0.040 AN | 131 | 133 | 117 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7121 | 0.8564 | 1.6045 | 4.3987 | 19.7074 | 100.0969 | 142.42 | 85.64 | 80.225 | 87.974 | 98.537 | 100.0969 | 1,1,1,2-Tetrachloroethane | 9.01 | Lin, Avg | 6 | 1.7353 | 0.99991 | TRUE | TRUE
Ethylbenzene | 9.018 | 0.040 AN | 91 | 106 | 51 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7182 | 1.0175 | 1.829 | 4.8002 | 21.3537 | 99.7414 | 143.64 | 101.75 | 91.45 | 96.004 | 106.7685 | 99.7414 | Ethylbenzene | 9.02 | Lin, Avg | 6 | 2.9433 | 0.99974 | TRUE | TRUE
m/p-Xylene | 9.125 | 0.060 AN | 91 | 106 | 105 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7652 | 1.0635 | 1.8776 | 5.1302 | 22.8814 | 99.4177 | 153.04 | 106.35 | 93.88 | 102.604 | 114.407 | 99.4177 | m/p-Xylene | 9.13 | Lin, Avg | 6 | 6.052 | 0.99884 | TRUE | TRUE
o-Xylene | 9.423 | 0.060 AN | 91 | 106 | 105 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7334 | 1.0458 | 1.8143 | 4.8434 | 21.3149 | 99.7469 | 146.68 | 104.58 | 90.715 | 96.868 | 106.5745 | 99.7469 | o-Xylene | 9.43 | Lin, Avg | 6 | 2.8626 | 0.99975 | TRUE | TRUE
Styrene | 9.439 | 0.060 AN | 104 | 78 | 103 | Internal Chlorobenzene-d5 [IS3] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8129 | 1.0293 | 1.8896 | 5.1285 | 22.1896 | 99.5576 | #DIV/0! | 102.93 | 94.48 | 102.57 | 110.948 | 99.5576 | Styrene | 9.44 | Lin, Avg | 5 | 4.3153 | 0.99929 | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.808 | 0.8296 | 1.4823 | 4.3072 | 18.1353 | 100.4196 | #DIV/0! | 82.96 | 74.115 | 86.144 | 90.6765 | 100.4196 | Bromoform | 9.57 | Lin, Avg | 5 | 4.2037 | 0.99939 | TRUE | TRUE
Isopropylbenzene (cumene) | 9.726 | 0.060 AN | 105 | 120 | 79 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6172 | 0.8589 | 1.6101 | 4.5804 | 20.645 | 99.9006 | 123.44 | 85.89 | 80.505 | 91.608 | 103.225 | 99.9006 | Isopropylbenzene (cumene) | 9.71 | Lin, Avg | 6 | 1.8574 | 0.9999 | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.835 | 0.060 AN | 95 | 174 | 176 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 19.9777 | 20.2899 | 19.928 | 19.9004 | 19.8704 | 20.0336 | 99.8885 | 101.4495 | 99.64 | 99.502 | 99.352 | 100.168 | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Bromobenzene | 9.936 | 0.040 AN | 77 | 156 | 158 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.805 | 0.9925 | 1.7504 | 4.6758 | 19.2982 | 100.1616 | #DIV/0! | 99.25 | 87.52 | 93.516 | 96.491 | 100.1616 | Bromobenzene | 9.93 | Lin, Avg | 5 | 1.632 | 0.9999 | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.944 | 0.040 AN | 83 | 85 | 95 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7113 | 0.7323 | 1.323 | 3.9113 | 17.4906 | 100.5715 | 142.26 | 73.23 | 66.15 | 78.226 | 87.453 | 100.5715 | 1,1,2,2-Tetrachloroethane | 9.95 | Lin, Avg | 6 | 6.2284 | 0.99893 | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.969 | 0.040 AN | 75 | 53 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.046 | 1.043 | 1.9049 | 5.2617 | 20.195 | 99.9494 | #DIV/0! | 104.3 | 95.245 | 105.234 | 100.975 | 99.9494 | 1,2,3-Trichloropropane (TCP) | 9.98 | Lin, Avg | 5 | 0.6746 | 0.99998 | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.975 | 0.040 AN | 77 | 110 | 61 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 1.0267 | 1.01 | 1.7826 | 5.1337 | 20.019 | 99.9938 | #DIV/0! | 101 | 89.13 | 102.674 | 100.095 | 99.9938 | trans-1,4-Dichloro-2-butene | 9.98 | Lin, Avg | 5 | 0.5007 | 0.99999 | TRUE | TRUE
n-Propylbenzene | 10.016 | 0.026 AN | 91 | 120 | 65 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6356 | 0.9239 | 1.6958 | 4.781 | 21.0069 | 99.8157 | 127.12 | 92.39 | 84.79 | 95.62 | 105.0345 | 99.8157 | n-Propylbenzene | 10.02 | Lin, Avg | 6 | 2.293 | 0.99984 | TRUE | TRUE
2-Chlorotoluene | 10.071 | 0.031 AN | 91 | 126 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7641 | 1.0349 | 1.8366 | 5.0198 | 21.6486 | 99.6709 | 152.82 | 103.49 | 91.83 | 100.396 | 108.243 | 99.6709 | 2-Chlorotoluene | 10.08 | Lin, Avg | 6 | 3.5303 | 0.99962 | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.154 | 0.040 AN | 105 | 120 | 119 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6386 | 0.9064 | 1.6656 | 4.8169 | 21.8626 | 99.6436 | 127.72 | 90.64 | 83.28 | 96.338 | 109.313 | 99.6436 | 1,3,5-Trimethylbenzene | 10.15 | Lin, Avg | 6 | 4.022 | 0.99951 | TRUE | TRUE
4-Chlorotoluene | 10.168 | 0.040 AN | 91 | 126 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7315 | 0.9905 | 1.7695 | 5.042 | 21.6965 | 99.6622 | 146.3 | 99.05 | 88.475 | 100.84 | 108.4825 | 99.6622 | 4-Chlorotoluene | 10.17 | Lin, Avg | 6 | 3.6376 | 0.99959 | TRUE | TRUE
tert-Butylbenzene | 10.357 | 0.040 AN | 119 | 91 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5768 | 0.8284 | 1.5057 | 4.2772 | 19.367 | 100.174 | 115.36 | 82.84 | 75.285 | 85.544 | 96.835 | 100.174 | tert-Butylbenzene | 10.37 | Lin, Avg | 6 | 2.3511 | 0.99984 | TRUE | TRUE
Pentachloroethane | 10.381 | 0.030 AN | 167 | 130 | 132 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8378 | 1 | 2 | 5 | 20 | 100 | #DIV/0! | 100 | 100 | 100 | 100 | 100 | Pentachloroethane | 10.39 | P-P | 5 | 0 | 1 | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.411 | 0.030 AN | 105 | 120 | 77 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6824 | 0.9288 | 1.7245 | 4.9581 | 21.1605 | 99.7753 | 136.48 | 92.88 | 86.225 | 99.162 | 105.8025 | 99.7753 | 1,2,4-Trimethylbenzene | 10.41 | Lin, Avg | 6 | 2.5543 | 0.9998 | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.521 | 0.037 AN | 105 | 134 | 91 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6008 | 0.9163 | 1.6637 | 4.7383 | 21.3549 | 99.7492 | 120.16 | 91.63 | 83.185 | 94.766 | 106.7745 | 99.7492 | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | Lin, Avg | 6 | 3.0118 | 0.99973 | TRUE | TRUE
1,3-Dichlorobenzene | 10.603 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8452 | 1.0716 | 1.9256 | 5.2758 | 21.5899 | 99.669 | #DIV/0! | 107.16 | 96.28 | 105.516 | 107.9495 | 99.669 | 1,3-Dichlorobenzene | 10.61 | Lin, Avg | 5 | 3.1856 | 0.99962 | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6191 | 0.9024 | 1.6734 | 4.83 | 21.6672 | 99.682 | 123.82 | 90.24 | 83.67 | 96.6 | 108.336 | 99.682 | 4-Isopropyltoluene (p-cymene) | 10.63 | Lin, Avg | 6 | 3.6162 | 0.9996 | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.656 | 0.036 AN | 152 | 150 |  | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1,4-Dichlorobenzene | 10.679 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8753 | 1.1032 | 1.9613 | 5.4033 | 21.3308 | 99.7134 | #DIV/0! | 110.32 | 98.065 | 108.066 | 106.654 | 99.7134 | 1,4-Dichlorobenzene | 10.68 | Lin, Avg | 5 | 2.7489 | 0.99971 | TRUE | TRUE
n-Butylbenzene | 10.9 | 0.040 AN | 91 | 92 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5689 | 0.8548 | 1.5146 | 4.4937 | 20.8933 | 99.8575 | 113.78 | 85.48 | 75.73 | 89.874 | 104.4665 | 99.8575 | n-Butylbenzene | 10.91 | Lin, Avg | 6 | 2.4196 | 0.99983 | TRUE | TRUE
1,2-Dichlorobenzene | 10.914 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.9221 | 1.0611 | 1.9031 | 5.3164 | 21.9978 | 99.5859 | #DIV/0! | 106.11 | 95.155 | 106.328 | 109.989 | 99.5859 | 1,2-Dichlorobenzene | 10.92 | Lin, Avg | 5 | 3.9807 | 0.99939 | TRUE | TRUE
Hexachloroethane | 11.093 | 0.040 AN | 117 | 119 | 201 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.5 | 1 | 2 | 5 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Hexachloroethane | 11.1 | P-P | 6 | 0 | 1 | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.447 | 0.040 AN | 157 | 155 | 75 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7911 | 0.7436 | 1.4345 | 4.3072 | 18.2716 | 100.3942 | #DIV/0! | 74.36 | 71.725 | 86.144 | 91.358 | 100.3942 | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | Lin, Avg | 5 | 3.9993 | 0.99945 | TRUE | TRUE
Nitrobenzene | 11.566 | 0.040 AN | 77 | 51 | 123 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7174 | 0.7495 | 1.4445 | 3.8271 | 17.6732 | 100.5365 | 143.48 | 74.95 | 72.225 | 76.542 | 88.366 | 100.5365 | Nitrobenzene | 11.57 | Lin, Avg | 6 | 5.8797 | 0.99904 | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.966 | 0.040 AN | 180 | 182 | 145 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.6834 | 0.8605 | 1.5019 | 4.3374 | 17.7614 | 100.4913 | 136.68 | 86.05 | 75.095 | 86.748 | 88.807 | 100.4913 | 1,2,4-Trichlorobenzene | 11.97 | Lin, Avg | 6 | 5.2284 | 0.99923 | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.066 | 0.040 AN | 225 | 227 | 223 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 0.5 | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.595 | 0.8484 | 1.4705 | 4.454 | 18.3898 | 100.361 | 119 | 84.84 | 73.525 | 89.08 | 91.949 | 100.361 | Hexachloro-1,3-butadiene | 12.06 | Lin, Avg | 6 | 3.8844 | 0.99957 | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.8157 | 0.8202 | 1.487 | 4.4095 | 18.2856 | 100.3845 | #DIV/0! | 82.02 | 74.35 | 88.19 | 91.428 | 100.3845 | Naphthalene | 12.15 | Lin, Avg | 5 | 3.8514 | 0.99949 | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.273 | 0.040 AN | 180 | 182 | 145 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 1 | 2 | 5 | 20 | 100 | ppb | 1 | 0.7719 | 0.8663 | 1.56 | 4.3059 | 17.8999 | 100.4649 | #DIV/0! | 86.63 | 78 | 86.118 | 89.4995 | 100.4649 | 1,2,3-Trichlorobenzene | 12.29 | Lin, Avg | 5 | 4.6094 | 0.99927 | TRUE | TRUE

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 1.8
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 1.8
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 1
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | n.a.
Pentafluorobenzene [IS1] | n.a.
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | n.a.
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 1
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 1.8
Toluene-d8 [SS2] | n.a.
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 1.8
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | n.a.
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 1
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | n.a.
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | n.a.
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 1
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.46 | 213 | 0.01 | 0.087 | Not confirmed | 50 | 52 | 31.71 | 33.73 | Confirmed | 49 | 9.01 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 31.1 | n.a. | Not confirmed | 61 | 7.37 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.84 | 726 | 0.05 | 0.11 | Confirmed | 94 | 96 | 92.8 | 82.08 | Confirmed | 93 | 18.28 | 16.47 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 31.17 | n.a. | Not confirmed | 49 | 22.55 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | 2.18 | 31 | 0 | 0.006 | Not confirmed | 101 | 103 | 63.97 | n.a. | Not confirmed | 105 | 9.43 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 73.64 | n.a. | Not confirmed | 45 | 71.41 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 61 | 96 | 76.66 | n.a. | Not confirmed | 98 | 49.44 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 33.08 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | Iodomethane | 2.89 | 406 | 0.03 | 0.103 | Not confirmed | 142 | 127 | 29.47 | 22.1 | Confirmed | 141 | 12.1 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon disulfide | 2.96 | 803 | 0.05 | 0.066 | Not confirmed | 76 | 78 | 8.53 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 50.39 | n.a. | Not confirmed | 76 | 34.85 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methylene chloride (DCM) | 3.35 | 254 | 0.02 | 0.047 | Not confirmed | 49 | 84 | 85.32 | 81.61 | Confirmed | 86 | 53.17 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | trans-1,2-Dichloroethene | 3.68 | 151 | 0.01 | 0.026 | Confirmed | 61 | 96 | 74.1 | 61.94 | Confirmed | 98 | 46.54 | 33.83 | Confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 26.71 | n.a. | Not confirmed | 57 | 22.34 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 65 | 31.43 | n.a. | Not confirmed | 83 | 11.93 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 76.36 | n.a. | Not confirmed | 79 | 31.16 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | cis-1,2-Dichloroethene | 4.82 | 78 | 0.01 | 0.012 | Confirmed | 61 | 96 | 79.58 | 71.87 | Confirmed | 98 | 50.7 | 61.9 | Confirmed
1.8 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 26.8 | n.a. | Not confirmed | 57 | 8.71 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 16.43 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 33.66 | n.a. | Not confirmed | 40 | 41.59 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromochloromethane | 5.08 | 39 | 0 | 0.01 | Not confirmed | 49 | 130 | 87.46 | 92 | Confirmed | 128 | 66.82 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 39.18 | n.a. | Not confirmed | 71 | 39.07 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 65.26 | n.a. | Not confirmed | 47 | 18.4 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 99 | 63.31 | n.a. | Not confirmed | 61 | 39.67 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.36 | 69639 | 4.71 | 19.892 | Confirmed | 113 | 111 | 101.16 | 104.37 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.43 | 157728 | 10.68 | 20 | Confirmed | 168 | 99 | 44.54 | 44.28 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 56 | 41 | 57.62 | n.a. | Not confirmed | 43 | 26.17 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon tetrachloride | n.a. | n.a. | n.a. | n.a. | Not confirmed | 119 | 121 | 31.46 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,1-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 77 | 31.06 | n.a. | Not confirmed | 110 | 44.53 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Benzene | 5.71 | 237 | 0.02 | 0.013 | Not confirmed | 78 | 77 | 23.38 | n.a. | Not confirmed | 52 | 14.64 | 31.17 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | 5.78 | 59 | 0 | 0.014 | Not confirmed | 62 | 64 | 31.86 | n.a. | Not confirmed | 49 | 31.03 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.17 | 243100 | 16.46 | 20 | Confirmed | 114 | 88 | 17.05 | 17.13 | Confirmed | 63 | 16.81 | 16.76 | Confirmed
1 | TRUE | FALSE | FALSE | Trichloroethene | 6.39 | 161 | 0.01 | 0.03 | Confirmed | 130 | 132 | 97.52 | 97.15 | Confirmed | 95 | 89.89 | 98.26 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 69.37 | n.a. | Not confirmed | 41 | 43.58 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.73 | 42 | 0 | 0.013 | Not confirmed | 174 | 93 | 81.5 | 118.26 | Not confirmed | 95 | 68.91 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 84.46 | n.a. | Not confirmed | 39 | 46.2 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 64.36 | n.a. | Not confirmed | 47 | 16.04 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 81.84 | n.a. | Not confirmed | 39 | 25.82 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 39 | 44.86 | n.a. | Not confirmed | 77 | 32.49 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 40.35 | n.a. | Not confirmed | 41 | 23.25 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.61 | 307433 | 20.81 | 20.607 | Confirmed | 98 | 100 | 65.28 | 63.93 | Confirmed | 70 | 10.31 | 10.15 | Confirmed
1 | TRUE | FALSE | FALSE | Toluene | 7.67 | 1351 | 0.09 | 0.068 | Confirmed | 91 | 92 | 58.02 | 57.96 | Confirmed | 65 | 10.97 | 10.2 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,3-Dichloropropene | 7.93 | 151 | 0.01 | 0.038 | Not confirmed | 75 | 39 | 44.91 | n.a. | Not confirmed | 77 | 31.89 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 69 | 41 | 64.81 | n.a. | Not confirmed | 99 | 26.63 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | 8.1 | 54 | 0 | 0.013 | Not confirmed | 97 | 83 | 84.16 | n.a. | Not confirmed | 99 | 62.29 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.16 | 243 | 0.02 | 0.032 | Confirmed | 166 | 164 | 77.08 | 75.46 | Confirmed | 129 | 67.48 | 66.41 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3-Dichloropropane | 8.24 | 107 | 0.01 | 0.016 | Not confirmed | 76 | 41 | 64.98 | n.a. | Not confirmed | 78 | 32.57 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 55.72 | n.a. | Not confirmed | 57 | 19.29 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | 8.42 | 29 | 0 | 0.006 | Not confirmed | 129 | 127 | 76.76 | 42.96 | Not confirmed | 131 | 23.36 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 94.98 | n.a. | Not confirmed | 93 | 3.9 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.91 | 230486 | 15.6 | 20 | Confirmed | 117 | 82 | 53.61 | 53.57 | Confirmed | 52 | 13.31 | 12.73 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.94 | 534 | 0.04 | 0.04 | Not confirmed | 112 | 77 | 60.85 | 282.93 | Not confirmed | 114 | 31.83 | 29.09 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 131 | 133 | 96.02 | n.a. | Not confirmed | 117 | 82.72 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Ethylbenzene | 9.02 | 1361 | 0.09 | 0.07 | Confirmed | 91 | 106 | 37.89 | 42.2 | Confirmed | 51 | 8.74 | 8.57 | Confirmed
1 | TRUE | FALSE | FALSE | m/p-Xylene | 9.13 | 3018 | 0.2 | 0.095 | Confirmed | 91 | 106 | 55.32 | 54.04 | Confirmed | 105 | 21.95 | 21.51 | Confirmed
1 | TRUE | FALSE | FALSE | o-Xylene | 9.43 | 1165 | 0.08 | 0.067 | Confirmed | 91 | 106 | 52.77 | 51.54 | Confirmed | 105 | 25.68 | 26.89 | Confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.45 | 1182 | 0.08 | 0.083 | Confirmed | 104 | 78 | 49.97 | 49.02 | Confirmed | 103 | 52.14 | 48.07 | Confirmed
1 | TRUE | FALSE | FALSE | Bromoform | 9.57 | 35 | 0 | 0.01 | Confirmed | 173 | 171 | 50.06 | 55.54 | Confirmed | 175 | 47.63 | 45.87 | Confirmed
1 | TRUE | FALSE | FALSE | Isopropylbenzene (cumene) | 9.71 | 578 | 0.04 | 0.028 | Confirmed | 105 | 120 | 31.21 | 29.97 | Confirmed | 79 | 14.85 | 14.9 | Confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 116631 | 7.9 | 20.912 | Confirmed | 95 | 174 | 88.43 | 89.12 | Confirmed | 176 | 85.06 | 85.94 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | 9.94 | 476 | 0.03 | 0.237 | Not confirmed | 77 | 110 | 82.83 | 6.94 | Not confirmed | 61 | 55.97 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.94 | 476 | 0.03 | 0.056 | Confirmed | 77 | 156 | 72.68 | 75.41 | Confirmed | 158 | 71.4 | 77.56 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | 9.95 | 60 | 0 | 0.011 | Not confirmed | 83 | 85 | 65.36 | n.a. | Not confirmed | 95 | 14.99 | 97.58 | Not confirmed
1 | TRUE | TRUE | FALSE | trans-1,4-Dichloro-2-butene | 9.98 | 92 | 0.01 | 0.015 | Not confirmed | 75 | 53 | 19.65 | n.a. | Not confirmed | 89 | 12 | 32.48 | Confirmed
1 | TRUE | FALSE | FALSE | n-Propylbenzene | 10.02 | 1497 | 0.1 | 0.063 | Confirmed | 91 | 120 | 28.56 | 28.63 | Confirmed | 65 | 9.75 | 8.78 | Confirmed
1 | TRUE | FALSE | FALSE | 2-Chlorotoluene | 10.08 | 712 | 0.05 | 0.052 | Confirmed | 91 | 126 | 39.64 | 40.26 | Confirmed | 89 | 17.23 | 14.55 | Confirmed
1 | TRUE | FALSE | FALSE | 1,3,5-Trimethylbenzene | 10.15 | 750 | 0.05 | 0.039 | Confirmed | 105 | 120 | 52.59 | 49.44 | Confirmed | 119 | 11.93 | 10.32 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Chlorotoluene | 10.17 | 1433 | 0.1 | 0.086 | Confirmed | 91 | 126 | 35.46 | 38.66 | Confirmed | 89 | 10.88 | 11.01 | Confirmed
1 | TRUE | FALSE | FALSE | tert-Butylbenzene | 10.37 | 400 | 0.03 | 0.022 | Confirmed | 119 | 91 | 63.34 | 71.9 | Confirmed | 134 | 25.06 | 27.33 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 167 | 130 | 53.8 | n.a. | Not confirmed | 132 | 62.91 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,4-Trimethylbenzene | 10.41 | 1224 | 0.08 | 0.064 | Not confirmed | 105 | 120 | 50.81 | 48.69 | Confirmed | 77 | 10.34 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 1033 | 0.07 | 0.045 | Confirmed | 105 | 134 | 22.93 | 23.62 | Confirmed | 91 | 14.28 | 16.65 | Confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.61 | 1319 | 0.09 | 0.124 | Confirmed | 146 | 148 | 63.81 | 59.44 | Confirmed | 111 | 39.53 | 37.19 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.64 | 1105 | 0.07 | 0.057 | Confirmed | 119 | 91 | 25.02 | 27.69 | Confirmed | 134 | 31.11 | 30.03 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 136105 | 9.21 | 20 | Confirmed | 152 | 115 | 57.32 | 60.06 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.68 | 1382 | 0.09 | 0.126 | Not confirmed | 146 | 148 | 63.43 | 81.82 | Confirmed | 111 | 44.06 | 132.4 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.91 | 1994 | 0.14 | 0.107 | Confirmed | 91 | 92 | 52.7 | 55.13 | Confirmed | 134 | 31.86 | 30.91 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.92 | 895 | 0.06 | 0.081 | Confirmed | 146 | 148 | 62.65 | 63.31 | Confirmed | 111 | 40.2 | 41.73 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 96.49 | n.a. | Not confirmed | 201 | 98.24 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 157 | 155 | 76.06 | n.a. | Not confirmed | 75 | 77.06 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 49.41 | n.a. | Not confirmed | 123 | 54.32 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.98 | 1362 | 0.09 | 0.141 | Confirmed | 180 | 182 | 96.18 | 101.16 | Confirmed | 145 | 30.56 | 29.42 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.07 | 399 | 0.03 | 0.097 | Confirmed | 225 | 227 | 62.56 | 59.39 | Confirmed | 223 | 61.61 | 69.31 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.15 | 2171 | 0.15 | 0.083 | Confirmed | 128 | 127 | 12.21 | 12.79 | Confirmed | 129 | 10.62 | 11.08 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.29 | 1134 | 0.08 | 0.118 | Confirmed | 180 | 182 | 96.06 | 94.75 | Confirmed | 145 | 31.93 | 29.89 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2023
231019
Peak Name | Amount
ppb
First Injection | H2O-R1 | H2O-R2 | H2O-V1 | H2O-V2 | CAL1-0.5 | CAL2-1 | CAL3-2 | CAL4-5 | CAL5-20 | CAL6-100 | H2O-V3 | H2O-V4 | H2O-V5 | H2O-V6 | MB | 10-9-23-JI-01 | 10-9-23-JI-02 | 10-9-23-JI-03 | 10-9-23-JI-04 | 10-3-23-JI-05 | UR-Tank | CCV | H2O-V7
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8181 | 1.1589 | 1.8917 | 4.5922 | 19.7375 | 100.0735 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7575 | n.a./n.r.
Chloroethene (vinyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6227 | 0.9769 | 1.7516 | 4.4724 | 20.0012 | 100.0307 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9034 | n.a./n.r.
Bromomethane (methyl bromide) | n.a./n.r. | 0.0485 | 0.0364 | 0.0426 | 0.7536 | 1 | 2 | 5 | 20 | 100 | 0.2485 | 0.1603 | 0.1082 | 0.1095 | 0.07 | 0.1046 | 0.0884 | 0.1062 | 0.1007 | 0.0964 | 0.0858 | 11.4699 | 0.0838
Chloroethane (ethyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6247 | 0.9442 | 1.6725 | 4.5078 | 19.2934 | 100.1724 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1004 | n.a./n.r.
Trichlorofluoromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.513 | 0.8815 | 1.6805 | 4.2772 | 19.7917 | 100.0853 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.529 | n.a./n.r.
Diethyl ether | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8618 | 1.0197 | 1.826 | 4.9486 | 19.5703 | 100.0918 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6936 | n.a./n.r.
1,1-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6132 | 0.9245 | 1.442 | 4.4148 | 19.7851 | 100.0836 | 0.1079 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3804 | n.a./n.r.
Acetone | 0.5443 | 0.5741 | n.a./n.r. | n.a./n.r. | 1.8516 | 1.651 | 3.0661 | 8.0419 | 34.6498 | 180.3301 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 6.1333 | 13.3334 | 6.1811 | 8.242 | 9.4437 | 10.9961 | 16.5524 | n.a./n.r.
Iodomethane | 0.0947 | 0.067 | n.a./n.r. | n.a./n.r. | 0.5 | 1 | 2 | 5 | 20 | 100 | 0.3133 | 0.1713 | 0.1292 | n.a./n.r. | 0.0802 | 0.1162 | 0.0899 | 0.1011 | n.a./n.r. | 0.1194 | 0.1034 | 8.175 | 0.1264
Carbon disulfide | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6427 | 0.976 | 1.7319 | 4.5063 | 19.9561 | 100.0383 | 0.4745 | 0.1688 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6981 | n.a./n.r.
3-Chloropropene (allyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5826 | 0.856 | 1.5414 | 4.0683 | 18.2517 | 100.4064 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6718 | n.a./n.r.
Methylene chloride (DCM) | 0.0759 | 0.09 | 0.0439 | 0.0545 | 0.9511 | 1.138 | 1.9638 | 4.9762 | 19.9538 | 100.0098 | 0.1085 | 0.0471 | n.a./n.r. | n.a./n.r. | 0.1176 | n.a./n.r. | 0.035 | n.a./n.r. | n.a./n.r. | 0.0224 | n.a./n.r. | 9.8915 | n.a./n.r.
trans-1,2-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6315 | 0.8919 | 1.6547 | 4.2235 | 18.8901 | 100.2681 | 0.1998 | 0.0584 | 0.0372 | 0.026 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7792 | 0.022
Methyl tert-butyl ether (MTBE) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9056 | 0.9534 | 1.7822 | 4.4929 | 18.5067 | 100.3288 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1644 | n.a./n.r.
1,1-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7655 | 1.0124 | 1.779 | 4.6233 | 19.8728 | 100.0486 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7215 | n.a./n.r.
2,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5371 | 0.8169 | 1.4374 | 3.7934 | 17.8048 | 100.5123 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1702 | n.a./n.r.
cis-1,2-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7303 | 0.9357 | 1.6635 | 4.3072 | 18.6963 | 100.3027 | 0.1619 | 0.0341 | 0.0196 | 0.0118 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8982 | 0.018
2-Butanone (MEK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.5522 | 1.6182 | 2.8322 | 8.3443 | 34.7511 | 180.2997 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5525 | 17.5358 | n.a./n.r.
Methyl acrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7454 | 0.8232 | 1.4302 | 4.3353 | 18.808 | 100.2848 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4634 | n.a./n.r.
Bromochloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9276 | 1.0091 | 1.77 | 4.6552 | 18.9539 | 100.231 | 0.1105 | 0.0318 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0107 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3861 | n.a./n.r.
Methacrylonitrile | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7563 | 0.6713 | 1.4771 | 4.3604 | 18.0301 | 100.4397 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0411 | n.a./n.r.
Tetrahydrofuran | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8376 | 0.812 | 1.7061 | 4.7701 | 18.8566 | 100.2479 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 49.3241 | 9.5562 | n.a./n.r.
Trichloromethane (chloroform) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8238 | 1.0432 | 1.898 | 4.8213 | 19.8036 | 100.0498 | 0.0843 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.3979 | 40.0372 | 40.1379 | 43.2032 | 37.951 | 63.0649 | 9.9549 | 0.0314
1,1,1-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5999 | 0.9064 | 1.7452 | 4.671 | 20.677 | 99.8866 | 0.0421 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9119 | n.a./n.r.
Dibromofluoromethane [SS1] | 20.2339 | 20.2886 | 19.6386 | 20.3122 | 19.7372 | 19.5638 | 19.9286 | 19.9486 | 20.3044 | 20.5174 | 20.328 | 20.3929 | 19.846 | 19.8921 | 19.8735 | 20.7408 | 21.066 | 20.6819 | 20.9672 | 20.2363 | 20.7001 | 20.7487 | 20.3772
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.583 | 0.8497 | 1.5665 | 4.2049 | 18.3812 | 100.3733 | 0.1961 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5915 | n.a./n.r.
Carbon tetrachloride | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5111 | 0.7362 | 1.4539 | 3.8412 | 18.242 | 100.423 | 0.0648 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6537 | n.a./n.r.
1,1-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6185 | 0.9061 | 1.6194 | 4.3669 | 19.3162 | 100.1764 | 0.3118 | 0.0804 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1934 | n.a./n.r.
Benzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7148 | 0.9295 | 1.762 | 4.6472 | 20.1736 | 99.9873 | 0.1078 | 0.0289 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4819 | n.a./n.r.
1,2-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.1169 | 1.2162 | 2.1882 | 5.8832 | 22.5145 | 99.447 | 0.1478 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.1626 | n.a./n.r.
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7968 | 1.064 | 1.9598 | 5.128 | 21.1554 | 99.7627 | 0.3283 | 0.086 | 0.0381 | 0.0295 | 0.0168 | n.a./n.r. | n.a./n.r. | 0.0091 | 0.0089 | 0.0083 | 0.0066 | 10.101 | 0.0332
1,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8664 | 1.0363 | 1.8868 | 5.0194 | 20.3628 | 99.9284 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.015 | n.a./n.r.
Dibromomethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0878 | 1.0759 | 1.9635 | 5.3734 | 20.431 | 99.8951 | 0.1926 | 0.0607 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.2639 | 0.0156
Methyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8858 | 0.8398 | 1.6071 | 4.485 | 18.2894 | 100.3773 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8775 | n.a./n.r.
Bromodichloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.863 | 1.0001 | 1.8061 | 4.9065 | 20.4297 | 99.9226 | 0.1056 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.8607 | 15.3096 | 14.3678 | 16.6781 | 14.5524 | 10.7962 | 10.0024 | 0.021
2-Nitropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9667 | 1.0897 | 1.8088 | 4.5342 | 18.7203 | 100.2822 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4227 | n.a./n.r.
cis-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7066 | 0.8345 | 1.5167 | 4.189 | 17.7678 | 100.4973 | 0.2342 | 0.0579 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4695 | n.a./n.r.
4-Methyl-2-pentanone (MIBK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.662 | 1.7037 | 3.1872 | 8.564 | 34.4038 | 180.3503 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 17.1255 | n.a./n.r.
Toluene-d8 [SS2] | 19.9106 | 19.7945 | 19.7396 | 19.4992 | 19.5749 | 19.6155 | 20.0527 | 20.0628 | 20.435 | 20.259 | 20.4092 | 19.5398 | 20.3235 | 20.6071 | 19.8445 | 20.2983 | 19.2744 | 19.3537 | 19.8682 | 19.602 | 20.3922 | 20.3001 | 19.8031
Toluene | n.a./n.r. | n.a./n.r. | 0.0485 | 0.0468 | 0.6825 | 0.9436 | 1.7434 | 4.6028 | 20.1798 | 99.9887 | 0.3309 | 0.1066 | 0.0727 | 0.0678 | 0.0476 | 0.0696 | 0.0466 | 0.081 | 0.0642 | 0.057 | 0.0395 | 9.4481 | 0.0808
trans-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9203 | 1 | 2 | 5 | 20 | 100 | 0.4141 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5857 | n.a./n.r.
Ethyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8727 | 0.9118 | 1.5823 | 4.3625 | 17.9893 | 100.4433 | 0.1145 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0977 | n.a./n.r.
1,1,2-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0099 | 1.091 | 1.9384 | 5.2513 | 20.9797 | 99.7918 | 0.1577 | 0.0333 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0602 | n.a./n.r.
Tetrachloroethene | 0.0087 | 0.0064 | n.a./n.r. | n.a./n.r. | 0.9048 | 1.215 | 2.2003 | 5.8044 | 23.1312 | 99.3274 | 0.4553 | 0.106 | 0.0475 | 0.0318 | 0.0161 | 0.0148 | 0.0134 | 0.0125 | 0.0109 | 0.0094 | 0.0082 | 11.3251 | 0.067
1,3-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9909 | 1.0533 | 1.946 | 5.1326 | 19.841 | 100.0257 | 0.1902 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9821 | n.a./n.r.
2-Hexanone | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.7229 | 1.661 | 3.1633 | 9.0459 | 34.4965 | 180.3085 | 0.1983 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 17.0445 | n.a./n.r.
Dibromochloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8304 | 0.914 | 1.658 | 4.4716 | 18.9845 | 100.2372 | 0.1447 | 0.0321 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 2.9912 | 4.6887 | 4.3569 | 5.639 | 4.724 | 3.3381 | 9.2133 | n.a./n.r.
1,2-Dibromoethane (EDB) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0689 | 1.1128 | 2.0033 | 5.4255 | 21.2386 | 99.7298 | 0.2859 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.7421 | n.a./n.r.
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7474 | 0.9731 | 1.7681 | 4.7203 | 20.0022 | 100.0172 | 0.531 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6663 | n.a./n.r.
1,1,1,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7121 | 0.8564 | 1.6045 | 4.3987 | 19.7074 | 100.0969 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2427 | n.a./n.r.
Ethylbenzene | n.a./n.r. | n.a./n.r. | 0.0417 | 0.0355 | 0.7182 | 1.0175 | 1.829 | 4.8002 | 21.3537 | 99.7414 | 0.6483 | 0.1641 | 0.0873 | 0.0702 | 0.0538 | n.a./n.r. | n.a./n.r. | 0.0137 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1458 | 0.1161
m/p-Xylene | 0.0187 | 0.0141 | 0.052 | 0.0538 | 0.7652 | 1.0635 | 1.8776 | 5.1302 | 22.8814 | 99.4177 | 0.7204 | 0.2004 | 0.1126 | 0.0952 | 0.068 | 0.045 | 0.0274 | 0.0228 | 0.0226 | 0.0216 | 0.022 | 10.612 | 0.1372
o-Xylene | n.a./n.r. | n.a./n.r. | 0.0432 | 0.0404 | 0.7334 | 1.0458 | 1.8143 | 4.8434 | 21.3149 | 99.7469 | 0.4417 | 0.1234 | 0.0769 | 0.0666 | 0.0463 | 0.0314 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9677 | 0.1141
Styrene | 0.0149 | n.a./n.r. | 0.0314 | 0.0281 | 0.8129 | 1.0293 | 1.8896 | 5.1285 | 22.1896 | 99.5576 | 0.7237 | 0.1934 | 0.0954 | 0.0826 | 0.0529 | 0.0329 | 0.0236 | 0.0203 | n.a./n.r. | 0.0163 | 0.0233 | 10.4751 | 0.1169
Bromoform | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.808 | 0.8296 | 1.4823 | 4.3072 | 18.1353 | 100.4196 | 0.186 | 0.0382 | 0.0137 | 0.0103 | n.a./n.r. | 0.1983 | 0.3563 | 0.3643 | 0.5133 | 0.3984 | 0.4159 | 8.6367 | 0.0235
Isopropylbenzene (cumene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6172 | 0.8589 | 1.6101 | 4.5804 | 20.645 | 99.9006 | 0.6019 | 0.1145 | 0.0418 | 0.0276 | 0.0146 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3542 | 0.0845
1-Bromo-4-fluorobenzene (BFB) [SS3] | 21.2796 | 20.4677 | 20.5291 | 20.2574 | 19.9777 | 20.2899 | 19.928 | 19.9004 | 19.8704 | 20.0336 | 19.6341 | 19.918 | 20.5278 | 20.9124 | 20.9821 | 20.3644 | 20.1906 | 19.663 | 20.571 | 19.9259 | 19.8618 | 20.0524 | 20.2653
Bromobenzene | 0.0119 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.805 | 0.9925 | 1.7504 | 4.6758 | 19.2982 | 100.1616 | 0.7106 | 0.1971 | 0.0815 | 0.0558 | 0.0375 | 0.0325 | 0.0261 | 0.0198 | 0.02 | 0.0129 | 0.0134 | 9.3812 | 0.0902
1,1,2,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7113 | 0.7323 | 1.323 | 3.9113 | 17.4906 | 100.5715 | 0.1592 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5121 | n.a./n.r.
1,2,3-Trichloropropane (TCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.046 | 1.043 | 1.9049 | 5.2617 | 20.195 | 99.9494 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1419 | n.a./n.r.
trans-1,4-Dichloro-2-butene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0267 | 1.01 | 1.7826 | 5.1337 | 20.019 | 99.9938 | 0.2992 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7519 | 0.0352
n-Propylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6356 | 0.9239 | 1.6958 | 4.781 | 21.0069 | 99.8157 | 1.0764 | 0.2672 | 0.0951 | 0.0632 | 0.0375 | 0.0342 | n.a./n.r. | 0.0193 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.565 | 0.1215
2-Chlorotoluene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7641 | 1.0349 | 1.8366 | 5.0198 | 21.6486 | 99.6709 | 0.7757 | 0.1898 | 0.0657 | 0.0519 | 0.0301 | 0.027 | 0.0219 | n.a./n.r. | 0.0113 | n.a./n.r. | n.a./n.r. | 10.0145 | 0.1051
1,3,5-Trimethylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6386 | 0.9064 | 1.6656 | 4.8169 | 21.8626 | 99.6436 | 0.7248 | 0.1528 | 0.053 | 0.0393 | 0.0247 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9721 | 0.0953
4-Chlorotoluene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7315 | 0.9905 | 1.7695 | 5.042 | 21.6965 | 99.6622 | 1.0131 | 0.2871 | 0.1132 | 0.0864 | 0.0427 | 0.043 | 0.0355 | 0.0253 | 0.0226 | 0.0226 | n.a./n.r. | 10.0131 | 0.1219
tert-Butylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5768 | 0.8284 | 1.5057 | 4.2772 | 19.367 | 100.174 | 0.5188 | 0.0934 | 0.0309 | 0.0219 | n.a./n.r. | 0.0087 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0388 | 0.07
Pentachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8378 | 1 | 2 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.6779 | n.a./n.r.
1,2,4-Trimethylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6824 | 0.9288 | 1.7245 | 4.9581 | 21.1605 | 99.7753 | 0.8686 | 0.2021 | 0.0778 | n.a./n.r. | 0.0389 | 0.0439 | 0.0328 | 0.0292 | 0.025 | n.a./n.r. | n.a./n.r. | 9.6086 | 0.1153
1-Methylpropylbenzene (sec-butylbenzene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6008 | 0.9163 | 1.6637 | 4.7383 | 21.3549 | 99.7492 | 1.0632 | 0.2231 | 0.0716 | 0.0449 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7231 | n.a./n.r.
1,3-Dichlorobenzene | 0.0222 | 0.014 | n.a./n.r. | 0.0079 | 0.8452 | 1.0716 | 1.9256 | 5.2758 | 21.5899 | 99.669 | 1.2338 | 0.3634 | 0.1534 | 0.124 | 0.0695 | 0.0619 | 0.0511 | 0.0346 | 0.0314 | 0.0287 | 0.0291 | 10.3564 | 0.1595
4-Isopropyltoluene (p-cymene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6191 | 0.9024 | 1.6734 | 4.83 | 21.6672 | 99.682 | 1.1825 | 0.274 | 0.0825 | 0.0566 | 0.0311 | 0.024 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9425 | 0.1189
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8753 | 1.1032 | 1.9613 | 5.4033 | 21.3308 | 99.7134 | 1.2668 | 0.4158 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.2971 | n.a./n.r.
n-Butylbenzene | 0.0206 | 0.0139 | n.a./n.r. | 0.0077 | 0.5689 | 0.8548 | 1.5146 | 4.4937 | 20.8933 | 99.8575 | 1.5421 | 0.4892 | 0.1596 | 0.1074 | 0.0659 | 0.051 | 0.043 | 0.0338 | 0.0268 | 0.0215 | 0.0221 | 9.16 | 0.1469
1,2-Dichlorobenzene | 0.0325 | 0.0184 | 0.012 | 0.0122 | 0.9221 | 1.0611 | 1.9031 | 5.3164 | 21.9978 | 99.5859 | 0.9989 | 0.2696 | 0.1062 | 0.0813 | 0.0486 | 0.0425 | 0.0409 | 0.0282 | 0.0232 | 0.0238 | 0.027 | 10.4924 | 0.1433
Hexachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5 | 1 | 2 | 5 | 20 | 100 | 0.2567 | 0.0425 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6422 | n.a./n.r.
1,2-Dibromo-3-chloropropane (DBCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7911 | 0.7436 | 1.4345 | 4.3072 | 18.2716 | 100.3942 | 0.3872 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7781 | n.a./n.r.
Nitrobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7174 | 0.7495 | 1.4445 | 3.8271 | 17.6732 | 100.5365 | 0.9086 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.4798 | n.a./n.r.
1,2,4-Trichlorobenzene | 0.0231 | 0.0162 | 0.0089 | 0.0105 | 0.6834 | 0.8605 | 1.5019 | 4.3374 | 17.7614 | 100.4913 | 1.5541 | 0.4698 | 0.1764 | 0.141 | 0.0834 | 0.0732 | 0.0547 | 0.0437 | 0.0424 | 0.037 | 0.0354 | 8.3658 | 0.1798
Hexachloro-1,3-butadiene | 0.0147 | 0.0104 | n.a./n.r. | 0.0049 | 0.595 | 0.8484 | 1.4705 | 4.454 | 18.3898 | 100.361 | 1.3014 | 0.3955 | 0.1346 | 0.0974 | 0.0546 | 0.0452 | 0.0345 | 0.0248 | 0.0245 | 0.016 | 0.019 | 8.5813 | 0.1116
Naphthalene | n.a./n.r. | 0.0099 | n.a./n.r. | n.a./n.r. | 0.8157 | 0.8202 | 1.487 | 4.4095 | 18.2856 | 100.3845 | 1.2611 | 0.3074 | 0.1083 | 0.083 | 0.049 | 0.0392 | 0.0335 | 0.0273 | 0.0253 | 0.0203 | 0.0217 | 8.6111 | 0.1684
1,2,3-Trichlorobenzene | 0.0214 | 0.0121 | 0.0068 | n.a./n.r. | 0.7719 | 0.8663 | 1.56 | 4.3059 | 17.8999 | 100.4649 | 1.4969 | 0.4126 | 0.1368 | 0.1183 | 0.0673 | 0.0558 | 0.0446 | 0.0338 | 0.0294 | 0.0238 | 0.0248 | 8.4557 | 0.1784

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library
| No. | min | SI |  |  | SI |  |  | SI
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC

## CCV
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? | <70% highlighted |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 159439 | 5.42 | 161264 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.17 | 253249 | 6.16 | 244452 | TRUE | TRUE | 97.58 | 10 | TRUE | Chloromethane (methyl chloride) | 1.46 | 24249 | 0.17 | 9.758 | Confirmed | 50 | 52 | 31.71 | 32.32 | Confirmed | 49 | 9.01 | 9.96 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 237201 | 8.91 | 230817 | TRUE | TRUE | 99.03 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.55 | 39918 | 0.27 | 9.903 | Confirmed | 62 | 64 | 31.1 | 31.94 | Confirmed | 61 | 7.37 | 7.67 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 144546 | 10.66 | 138670 | TRUE | TRUE | 114.7 | 10 | TRUE | Bromomethane (methyl bromide) | 1.83 | 53387 | 0.36 | 11.47 | Confirmed | 94 | 96 | 92.8 | 97.37 | Confirmed | 93 | 18.28 | 20.53 | Confirmed
[7 empty] | 91 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.94 | 28936 | 0.2 | 9.1 | Confirmed | 64 | 66 | 31.17 | 32.83 | Confirmed | 49 | 22.55 | 24.18 | Confirmed
[7 empty] | 95.29 | 10 | TRUE | Trichlorofluoromethane | 2.19 | 48462 | 0.33 | 9.529 | Confirmed | 101 | 103 | 63.97 | 65.09 | Confirmed | 105 | 9.43 | 10.62 | Confirmed
Total Analytes | 81 | [5 empty] | 96.94 | 10 | TRUE | Diethyl ether | 2.5 | 38271 | 0.26 | 9.694 | Confirmed | 59 | 74 | 73.64 | 74.61 | Confirmed | 45 | 71.41 | 72.95 | Confirmed
Failed | 1 | [5 empty] | 93.8 | 10 | TRUE | 1,1-Dichloroethene | 2.73 | 50018 | 0.34 | 9.38 | Confirmed | 61 | 96 | 76.66 | 76.93 | Confirmed | 98 | 49.44 | 49.88 | Confirmed
Allowance | 16.2 | [5 empty] | 91.9556 | 18 | TRUE | Acetone | 2.82 | 24199 | 0.17 | 16.552 | Confirmed | 43 | 58 | 33.08 | 40.27 | Confirmed | n.a. | n.a. | n.a. | n.a.
Pass? | TRUE | [5 empty] | 81.75 | 10 | TRUE | Iodomethane | 2.88 | 41130 | 0.28 | 8.175 | Confirmed | 142 | 127 | 29.47 | 29.44 | Confirmed | 141 | 12.1 | 11.59 | Confirmed
[7 empty] | 96.98 | 10 | TRUE | Carbon disulfide | 2.95 | 118531 | 0.81 | 9.698 | Confirmed | 76 | 78 | 8.53 | 8.84 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 86.72 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.19 | 60037 | 0.41 | 8.672 | Confirmed | 41 | 39 | 50.39 | 54.02 | Confirmed | 76 | 34.85 | 36.17 | Confirmed
[7 empty] | 98.91 | 10 | TRUE | Methylene chloride (DCM) | 3.36 | 54034 | 0.37 | 9.891 | Confirmed | 49 | 84 | 85.32 | 86.59 | Confirmed | 86 | 53.17 | 55.33 | Confirmed
[7 empty] | 87.79 | 10 | TRUE | trans-1,2-Dichloroethene | 3.68 | 51749 | 0.35 | 8.779 | Confirmed | 61 | 96 | 74.1 | 79.48 | Confirmed | 98 | 46.54 | 49.83 | Confirmed
[7 empty] | 91.64 | 10 | TRUE | Methyl tert-butyl ether (MTBE) | 3.69 | 111593 | 0.76 | 9.164 | Confirmed | 73 | 41 | 26.71 | 26.81 | Confirmed | 57 | 22.34 | 23.27 | Confirmed
[7 empty] | 97.21 | 10 | TRUE | 1,1-Dichloroethane | 4.19 | 68766 | 0.47 | 9.721 | Confirmed | 63 | 65 | 31.43 | 30.93 | Confirmed | 83 | 11.93 | 12.23 | Confirmed
[7 empty] | 81.7 | 10 | TRUE | 2,2-Dichloropropane | 4.82 | 37631 | 0.26 | 8.17 | Confirmed | 77 | 41 | 76.36 | 76.6 | Confirmed | 79 | 31.16 | 32.49 | Confirmed
[7 empty] | 88.98 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 59132 | 0.4 | 8.898 | Confirmed | 61 | 96 | 79.58 | 80.03 | Confirmed | 98 | 50.7 | 51.83 | Confirmed
[7 empty] | 97.4222 | 18 | TRUE | 2-Butanone (MEK) | 4.84 | 41895 | 0.29 | 17.536 | Confirmed | 43 | 72 | 26.8 | 27.05 | Confirmed | 57 | 8.71 | 8.36 | Confirmed
[7 empty] | 94.63 | 10 | TRUE | Methyl acrylate | 4.93 | 41339 | 0.28 | 9.463 | Confirmed | 55 | 85 | 16.43 | 16.15 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 90.41 | 10 | TRUE | Methacrylonitrile | 5.06 | 22788 | 0.16 | 9.041 | Confirmed | 67 | 52 | 33.66 | 32.15 | Confirmed | 40 | 41.59 | 38.16 | Confirmed
[7 empty] | 93.86 | 10 | TRUE | Bromochloromethane | 5.07 | 37773 | 0.26 | 9.386 | Confirmed | 49 | 130 | 87.46 | 88.64 | Confirmed | 128 | 66.82 | 69.33 | Confirmed
[7 empty] | 95.56 | 10 | TRUE | Tetrahydrofuran | 5.08 | 17699 | 0.12 | 9.556 | Confirmed | 42 | 72 | 39.18 | 41.05 | Confirmed | 71 | 39.07 | 40.98 | Confirmed
[7 empty] | 99.55 | 10 | TRUE | Trichloromethane (chloroform) | 5.2 | 66436 | 0.45 | 9.955 | Confirmed | 83 | 85 | 65.26 | 65.33 | Confirmed | 47 | 18.4 | 18.89 | Confirmed
[7 empty] | 99.12 | 10 | TRUE | 1,1,1-Trichloroethane | 5.34 | 50254 | 0.34 | 9.912 | Confirmed | 97 | 99 | 63.31 | 63.23 | Confirmed | 61 | 39.67 | 39.04 | Confirmed
[7 empty] | 103.745 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.36 | 73425 | 0.5 | 20.749 | Confirmed | 113 | 111 | 101.16 | 100.94 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 159439 | 1.09 | 20 | Confirmed | 168 | 99 | 44.54 | 44.88 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 85.91 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 71278 | 0.49 | 8.591 | Confirmed | 56 | 41 | 57.62 | 57.83 | Confirmed | 43 | 26.17 | 26.35 | Confirmed
[7 empty] | 86.54 | 10 | TRUE | Carbon tetrachloride | 5.49 | 39672 | 0.27 | 8.654 | Confirmed | 119 | 121 | 31.46 | 30.99 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 91.93 | 10 | TRUE | 1,1-Dichloropropene | 5.51 | 49863 | 0.34 | 9.193 | Confirmed | 75 | 77 | 31.06 | 31.1 | Confirmed | 110 | 44.53 | 45.76 | Confirmed
[7 empty] | 94.82 | 10 | TRUE | Benzene | 5.7 | 177301 | 1.21 | 9.482 | Confirmed | 78 | 77 | 23.38 | 23.42 | Confirmed | 52 | 14.64 | 14.96 | Confirmed
[7 empty] | 111.63 | 10 | TRUE | 1,2-Dichloroethane | 5.77 | 47713 | 0.33 | 11.163 | Confirmed | 62 | 64 | 31.86 | 31.76 | Confirmed | 49 | 31.03 | 31.1 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.17 | 253249 | 1.73 | 20 | Confirmed | 114 | 88 | 17.05 | 16.72 | Confirmed | 63 | 16.81 | 16.41 | Confirmed
[7 empty] | 101.01 | 10 | TRUE | Trichloroethene | 6.39 | 56543 | 0.39 | 10.101 | Confirmed | 130 | 132 | 97.52 | 97.75 | Confirmed | 95 | 89.89 | 88.97 | Confirmed
[7 empty] | 100.15 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 45086 | 0.31 | 10.015 | Confirmed | 63 | 62 | 69.37 | 68.5 | Confirmed | 41 | 43.58 | 43.32 | Confirmed
[7 empty] | 102.64 | 10 | TRUE | Dibromomethane | 6.72 | 34986 | 0.24 | 10.264 | Confirmed | 174 | 93 | 81.5 | 80.81 | Confirmed | 95 | 68.91 | 69.07 | Confirmed
[7 empty] | 88.78 | 10 | TRUE | Methyl methacrylate | 6.74 | 35842 | 0.24 | 8.878 | Confirmed | 41 | 69 | 84.46 | 86.61 | Confirmed | 39 | 46.2 | 46.38 | Confirmed
[7 empty] | 100.02 | 10 | TRUE | Bromodichloromethane | 6.92 | 49191 | 0.34 | 10.002 | Confirmed | 83 | 85 | 64.36 | 64.04 | Confirmed | 47 | 16.04 | 15.51 | Confirmed
[7 empty] | 94.23 | 10 | TRUE | 2-Nitropropane | 7.15 | 10371 | 0.07 | 9.423 | Confirmed | 43 | 41 | 81.84 | 86.2 | Confirmed | 39 | 25.82 | 27.98 | Confirmed
[7 empty] | 84.69 | 10 | TRUE | cis-1,3-Dichloropropene | 7.35 | 55806 | 0.38 | 8.469 | Confirmed | 75 | 39 | 44.86 | 43.3 | Confirmed | 77 | 32.49 | 32.07 | Confirmed
[7 empty] | 95.1389 | 18 | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.52 | 92929 | 0.63 | 17.125 | Confirmed | 43 | 58 | 40.35 | 40.58 | Confirmed | 41 | 23.25 | 23.98 | Confirmed
[7 empty] | 101.5 | 20 | TRUE | Toluene-d8 [SS2] | 7.61 | 311677 | 2.13 | 20.3 | Confirmed | 98 | 100 | 65.28 | 65.29 | Confirmed | 70 | 10.31 | 10.22 | Confirmed
[7 empty] | 94.48 | 10 | TRUE | Toluene | 7.67 | 193822 | 1.32 | 9.448 | Confirmed | 91 | 92 | 58.02 | 59.77 | Confirmed | 65 | 10.97 | 11.01 | Confirmed
[7 empty] | 95.86 | 10 | TRUE | trans-1,3-Dichloropropene | 7.92 | 42820 | 0.29 | 9.586 | Confirmed | 75 | 39 | 44.91 | 44.73 | Confirmed | 77 | 31.89 | 32.37 | Confirmed
[7 empty] | 90.98 | 10 | TRUE | Ethyl methacrylate | 7.99 | 53537 | 0.37 | 9.098 | Confirmed | 69 | 41 | 64.81 | 65.27 | Confirmed | 99 | 26.63 | 27.19 | Confirmed
[7 empty] | 100.6 | 10 | TRUE | 1,1,2-Trichloroethane | 8.1 | 43640 | 0.3 | 10.06 | Confirmed | 97 | 83 | 84.16 | 84.09 | Confirmed | 99 | 62.29 | 62.69 | Confirmed
[7 empty] | 113.25 | 10 | TRUE | Tetrachloroethene | 8.15 | 88990 | 0.61 | 11.325 | Confirmed | 166 | 164 | 77.08 | 77.8 | Confirmed | 129 | 67.48 | 67.31 | Confirmed
[7 empty] | 99.82 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 69182 | 0.47 | 9.982 | Confirmed | 76 | 41 | 64.98 | 63.56 | Confirmed | 78 | 32.57 | 31.43 | Confirmed
[7 empty] | 94.6889 | 18 | TRUE | 2-Hexanone | 8.31 | 62580 | 0.43 | 17.044 | Confirmed | 43 | 58 | 55.72 | 58.45 | Confirmed | 57 | 19.29 | 20.38 | Confirmed
[7 empty] | 92.13 | 10 | TRUE | Dibromochloromethane | 8.43 | 42043 | 0.29 | 9.213 | Confirmed | 129 | 127 | 76.76 | 76.91 | Confirmed | 131 | 23.36 | 23.93 | Confirmed
[7 empty] | 107.42 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.52 | 43968 | 0.3 | 10.742 | Confirmed | 107 | 109 | 94.98 | 92.22 | Confirmed | 93 | 3.9 | 4.33 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 237201 | 1.62 | 20 | Confirmed | 117 | 82 | 53.61 | 54.25 | Confirmed | 52 | 13.31 | 13.51 | Confirmed
[7 empty] | 96.66 | 10 | TRUE | Chlorobenzene | 8.93 | 132213 | 0.9 | 9.666 | Confirmed | 112 | 77 | 60.85 | 56.23 | Confirmed | 114 | 31.83 | 31.59 | Confirmed
[7 empty] | 92.43 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 38040 | 0.26 | 9.243 | Confirmed | 131 | 133 | 96.02 | 95.86 | Confirmed | 117 | 82.72 | 75.48 | Confirmed
[7 empty] | 101.46 | 10 | TRUE | Ethylbenzene | 9.02 | 202423 | 1.38 | 10.146 | Confirmed | 91 | 106 | 37.89 | 37.98 | Confirmed | 51 | 8.74 | 8.63 | Confirmed
[7 empty] | 106.12 | 10 | TRUE | m/p-Xylene | 9.13 | 346181 | 2.36 | 10.612 | Confirmed | 91 | 106 | 55.32 | 55.4 | Confirmed | 105 | 21.95 | 22.02 | Confirmed
[7 empty] | 99.68 | 10 | TRUE | o-Xylene | 9.43 | 179453 | 1.22 | 9.968 | Confirmed | 91 | 106 | 52.77 | 52.52 | Confirmed | 105 | 25.68 | 25.72 | Confirmed
[7 empty] | 104.75 | 10 | TRUE | Styrene | 9.44 | 154284 | 1.05 | 10.475 | Confirmed | 104 | 78 | 49.97 | 50.71 | Confirmed | 103 | 52.14 | 52.94 | Confirmed
[7 empty] | 86.37 | 10 | TRUE | Bromoform | 9.57 | 30960 | 0.21 | 8.637 | Confirmed | 173 | 171 | 50.06 | 51.58 | Confirmed | 175 | 47.63 | 48.5 | Confirmed
[7 empty] | 93.54 | 10 | TRUE | Isopropylbenzene (cumene) | 9.71 | 208344 | 1.42 | 9.354 | Confirmed | 105 | 120 | 31.21 | 31.62 | Confirmed | 79 | 14.85 | 15.3 | Confirmed
[7 empty] | 100.26 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 118771 | 0.81 | 20.052 | Confirmed | 95 | 174 | 88.43 | 87.39 | Confirmed | 176 | 85.06 | 87.09 | Confirmed
[7 empty] | 93.81 | 10 | TRUE | Bromobenzene | 9.93 | 85092 | 0.58 | 9.381 | Confirmed | 77 | 156 | 72.68 | 72.92 | Confirmed | 158 | 71.4 | 69.72 | Confirmed
[7 empty] | 85.12 | 10 | TRUE | 1,1,2,2-Tetrachloroethane | 9.95 | 48822 | 0.33 | 8.512 | Confirmed | 83 | 85 | 65.36 | 63.66 | Confirmed | 95 | 14.99 | 13.59 | Confirmed
[7 empty] | 101.42 | 10 | TRUE | 1,2,3-Trichloropropane (TCP) | 9.98 | 21611 | 0.15 | 10.142 | Confirmed | 77 | 110 | 82.83 | 82.98 | Confirmed | 61 | 55.97 | 58.64 | Confirmed
[7 empty] | 97.52 | 10 | TRUE | trans-1,4-Dichloro-2-butene | 9.98 | 63121 | 0.43 | 9.752 | Confirmed | 75 | 53 | 19.65 | 19.54 | Confirmed | 89 | 12 | 11.88 | Confirmed
[7 empty] | 95.65 | 10 | TRUE | n-Propylbenzene | 10.02 | 240560 | 1.64 | 9.565 | Confirmed | 91 | 120 | 28.56 | 28.88 | Confirmed | 65 | 9.75 | 9.79 | Confirmed
[7 empty] | 100.14 | 10 | TRUE | 2-Chlorotoluene | 10.08 | 145756 | 0.99 | 10.014 | Confirmed | 91 | 126 | 39.64 | 40.18 | Confirmed | 89 | 17.23 | 17.43 | Confirmed
[7 empty] | 99.72 | 10 | TRUE | 1,3,5-Trimethylbenzene | 10.15 | 202189 | 1.38 | 9.972 | Confirmed | 105 | 120 | 52.59 | 51.74 | Confirmed | 119 | 11.93 | 11.53 | Confirmed
[7 empty] | 100.13 | 10 | TRUE | 4-Chlorotoluene | 10.17 | 176413 | 1.2 | 10.013 | Confirmed | 91 | 126 | 35.46 | 35.81 | Confirmed | 89 | 10.88 | 11.23 | Confirmed
[7 empty] | 90.39 | 10 | TRUE | tert-Butylbenzene | 10.37 | 175060 | 1.19 | 9.039 | Confirmed | 119 | 91 | 63.34 | 63.2 | Confirmed | 134 | 25.06 | 24.76 | Confirmed
[7 empty] | 106.78 | 10 | TRUE | Pentachloroethane | 10.39 | 9140 | 0.06 | 10.678 | Confirmed | 167 | 130 | 53.8 | 51.42 | Confirmed | 132 | 62.91 | 52.55 | Confirmed
[7 empty] | 96.09 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.41 | 194069 | 1.32 | 9.609 | Confirmed | 105 | 120 | 50.81 | 51.83 | Confirmed | 77 | 10.34 | 10.68 | Confirmed
[7 empty] | 97.23 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 237736 | 1.62 | 9.723 | Confirmed | 105 | 134 | 22.93 | 22.68 | Confirmed | 91 | 14.28 | 14.24 | Confirmed
[7 empty] | 103.56 | 10 | TRUE | 1,3-Dichlorobenzene | 10.61 | 117008 | 0.8 | 10.356 | Confirmed | 146 | 148 | 63.81 | 63.42 | Confirmed | 111 | 39.53 | 39.28 | Confirmed
[7 empty] | 99.43 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.63 | 206344 | 1.41 | 9.943 | Confirmed | 119 | 91 | 25.02 | 25.41 | Confirmed | 134 | 31.11 | 30.81 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 144546 | 0.99 | 20 | Confirmed | 152 | 115 | 57.32 | 49.51 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 102.97 | 10 | TRUE | 1,4-Dichlorobenzene | 10.68 | 119730 | 0.82 | 10.297 | Confirmed | 146 | 148 | 63.43 | 63.76 | Confirmed | 111 | 44.06 | 40.01 | Confirmed
[7 empty] | 91.6 | 10 | TRUE | n-Butylbenzene | 10.91 | 180678 | 1.23 | 9.16 | Confirmed | 91 | 92 | 52.7 | 52.78 | Confirmed | 134 | 31.86 | 32.55 | Confirmed
[7 empty] | 104.92 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 122593 | 0.84 | 10.492 | Confirmed | 146 | 148 | 62.65 | 61.64 | Confirmed | 111 | 40.2 | 39.61 | Confirmed
[7 empty] | 96.42 | 10 | TRUE | Hexachloroethane | 11.1 | 21239 | 0.14 | 9.642 | Confirmed | 117 | 119 | 96.49 | 97.38 | Confirmed | 201 | 98.24 | 98.29 | Confirmed
[7 empty] | 87.78 | 10 | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.44 | 14691 | 0.1 | 8.778 | Confirmed | 157 | 155 | 76.06 | 77.3 | Confirmed | 75 | 77.06 | 77.4 | Confirmed
[7 empty] | 74.8 | 10 | FALSE | Nitrobenzene | 11.57 | 3247 | 0.02 | 7.48 | Confirmed | 77 | 51 | 49.41 | 49.59 | Confirmed | 123 | 54.32 | 56.8 | Confirmed
[7 empty] | 83.66 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.97 | 85785 | 0.59 | 8.366 | Confirmed | 180 | 182 | 96.18 | 94.92 | Confirmed | 145 | 30.56 | 30.77 | Confirmed
[7 empty] | 85.81 | 10 | TRUE | Hexachloro-1,3-butadiene | 12.06 | 37324 | 0.25 | 8.581 | Confirmed | 225 | 227 | 62.56 | 62.75 | Confirmed | 223 | 61.61 | 61.82 | Confirmed
[7 empty] | 86.11 | 10 | TRUE | Naphthalene | 12.14 | 239156 | 1.63 | 8.611 | Confirmed | 128 | 127 | 12.21 | 12.61 | Confirmed | 129 | 10.62 | 10.61 | Confirmed
[7 empty] | 84.56 | 10 | TRUE | 1,2,3-Trichlorobenzene | 12.29 | 86066 | 0.59 | 8.456 | Confirmed | 180 | 182 | 96.06 | 96.1 | Confirmed | 145 | 31.93 | 32.2 | Confirmed