## IS-RSD
[8 empty] | Peak Name | Peak Area
counts*min
Total Ion Count | [7 empty] | First Injection | Startup | H2O-V1 | H2O-V2 | H2O-V3 | H2O-V4 | H2O-V5 (TypeII) | H2O-V6 | H2O-R1 | H2O-R2 | H2O-R3
RSD | RSD (1-3) | RSD (2-4) | RSD (3-5) | RSD (4-6) | RSD (5-7) | RSD (6-8) | RSD (7-9) | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC
5.4129 | 8.0694 | 8.077 | 2.6787 | 2.8688 | 2.3717 | 2.9281 | 1.7046 | Pentafluorobenzene [IS1] | n.a. | 1133521.767 | 1006864.121 | 1179220.292 | 1134191.537 | 1120652.159 | 1073555.283 | 1116903.028 | 1137123.737 | 1155637.696
4.7799 | 7.3453 | 7.236 | 3.0489 | 1.7564 | 1.7242 | 3.1415 | 1.8898 | 1,4-Difluorobenzene [IS2] | n.a. | 1504572.293 | 1345255.39 | 1550562.235 | 1492613.018 | 1460223.201 | 1441790.719 | 1491727.734 | 1535309.783 | 1546156.775
6.5903 | 9.1746 | 9.1019 | 6.9798 | 4.729 | 5.388 | 3.1966 | 2.4185 | Chlorobenzene-d5 [IS3] | n.a. | 1853959.587 | 1616515.498 | 1934436.062 | 1842803.304 | 1683374.3 | 1809560.157 | 1872630.937 | 1929078.462 | 1964835.278
11.2996 | 13.3311 | 12.9363 | 14.9326 | 11.7022 | 12.2913 | 2.0853 | 1.744 | 1,4-Dichlorobenzene-d4 [IS4] | n.a. | 1934533.84 | 1751129.632 | 2273077.29 | 2054507.478 | 1681543.632 | 2093403.499 | 2103854.661 | 2174810.807 | 2159995.401
[8 empty] | Peak Name | Peak Area
counts*min
Quantitation Ion | [7 empty] | First Injection | Startup | H2O-V1 | H2O-V2 | H2O-V3 | H2O-V4 | H2O-V5 (TypeII) | H2O-V6 | H2O-R1 | H2O-R2 | H2O-R3
RSD | RSD (1-3) | RSD (2-4) | RSD (3-5) | RSD (4-6) | RSD (5-7) | RSD (6-8) | RSD (7-9) | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
5.7569 | 8.7527 | 8.7542 | 2.9136 | 2.5295 | 1.8761 | 2.472 | 2.9466 | Pentafluorobenzene [IS1] | n.a. | 444663.917 | 390857.911 | 464036.422 | 444718.722 | 439074.933 | 423442.192 | 435034.098 | 444898.679 | 461121.045
4.4342 | 6.7702 | 6.7158 | 2.7304 | 1.808 | 1.7682 | 3.5052 | 2.04 | 1,4-Difluorobenzene [IS2] | n.a. | 713823.392 | 646474.499 | 737801.053 | 710647.249 | 699853.666 | 685468.43 | 710050.454 | 735258.49 | 736204.345
6.2478 | 8.6508 | 8.5289 | 6.9745 | 4.4228 | 5.5524 | 2.6796 | 2.668 | Chlorobenzene-d5 [IS3] | n.a. | 682840.172 | 603663.73 | 716067.396 | 671703.058 | 622630.157 | 674001.652 | 694066.832 | 711102.59 | 732013.814
12.0286 | 14.2362 | 13.8222 | 15.6446 | 12.4707 | 12.6915 | 1.4647 | 1.5691 | 1,4-Dichlorobenzene-d4 [IS4] | n.a. | 356761.637 | 319931.575 | 422910.729 | 380584.356 | 308258.87 | 390410.287 | 384518.866 | 395950.014 | 394137.145
High RSD from vials might be due to multiple injections from the same vial
RSD from resevoir appears better

## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2023/231031.seq/206.smp/MSDevice.xraw
Channel | TIC
Injection Information: |  |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Data Vault | ChromeleonLocal |  |  | 95 | 20167550 | 50-200% of mass 174 | 95/174 | 131.0667 | 50 | 200 | TRUE
Injection | H2O-R3 |  |  | 96 | 1304864 | 5 to 9% of m/z 95 | 96/95 | 6.4701 | 5 | 9 | TRUE
Injection Number | 14 |  |  | 173 | 111488.4297 | <2% of m/z 174 | 173/174 | 0.72455 | 0 | 2 | TRUE
Position | 0 |  |  | 174 | 15387236 | 50-200% of mass 95 | 174/95 | 76.297 | 50 | 200 | TRUE
Comment |  |  |  | 175 | 1035747.1875 | 5 to 9% of m/z 174 | 175/174 | 6.7312 | 5 | 9 | TRUE
Processing Method | VOC-9 |  |  | 176 | 15758876 | 95 to 105% of m/z 174 | 176/174 | 102.4152 | 95 | 105 | TRUE
Instrument Method | VOC-25 |  |  | 177 | 977295.375 | 5 to 10% of m/z 176 | 177/176 | 6.2016 | 5 | 10 | TRUE
Type | Unknown
Status | Running
Injection Date | 2023-10-31
Injection Time | 14:06
Injection Volume (µl) | 20 | *N/A for P&T
Dilution Factor | 1
Weight | 1
Raw Data Information:
Data Points | 234
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9807
Retention Time (min) | 9.8359
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3286
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0244
Base Peak Intensity | 20167550
Total Ion Count | 82617859.9071
Raw Data: |  | Rounded:
Mass (m/z) | Value (counts) | Mass (m/z) | Value (counts)
35.1 | 1062.9114 | 35 | 1062.9114
36.1 | 88328.2656 | 36 | 88328.2656
37 | 485795.7188 | 37 | 485795.7188
38 | 443340.5312 | 38 | 443340.5312
39.1 | 173831.1719 | 39 | 173831.1719
40 | 103634.3594 | 40 | 103634.3594
41 | 8026.0879 | 41 | 8026.0879
42.1 | 4929.3359 | 42 | 4929.3359
43 | 3563.4253 | 43 | 3563.4253
44 | 163135.5938 | 44 | 163135.5938
45 | 108227.9453 | 45 | 108227.9453
46 | 7509.7378 | 46 | 7509.7378
47.1 | 234809.5469 | 47 | 234809.5469
48 | 72164.0547 | 48 | 72164.0547
49 | 465980.2188 | 49 | 465980.2188
50 | 2456162.5 | 50 | 2456162.5
51.1 | 729375.25 | 51 | 729375.25
52 | 32606.8086 | 52 | 32606.8086
52.8 | 2231.3975 | 53 | 2231.3975
53.3 | 1026.7639 | 53 | 1026.7639
54.2 | 900.1281 | 54 | 900.1281
55 | 28455.0781 | 55 | 28455.0781
56 | 128960.6953 | 56 | 128960.6953
57 | 288811.1875 | 57 | 288811.1875
58.1 | 9742.293 | 58 | 9742.293
59.3 | 1447.166 | 59 | 1447.166
60 | 99831.3359 | 60 | 99831.3359
61 | 510623.1875 | 61 | 510623.1875
62 | 514109.375 | 62 | 514109.375
63 | 400475.3438 | 63 | 400475.3438
64.1 | 37522.0625 | 64 | 37522.0625
65.1 | 14278.1846 | 65 | 14278.1846
66.1 | 2619.3835 | 66 | 2619.3835
67 | 29410.7012 | 67 | 29410.7012
68 | 1420311.5 | 68 | 1420311.5
69 | 1258526.75 | 69 | 1258526.75
70 | 104299.4453 | 70 | 104299.4453
71.1 | 3148.1162 | 71 | 3148.1162
72 | 66487.0625 | 72 | 66487.0625
73 | 550697.5 | 73 | 550697.5
74 | 2212770.25 | 74 | 2212770.25
75 | 8171616.5 | 75 | 8171616.5
76.1 | 678846.875 | 76 | 678846.875
77 | 108444.8672 | 77 | 108444.8672
78 | 77981.2734 | 78 | 77981.2734
78.9 | 124180.3438 | 79 | 124180.3438
80 | 54771.0703 | 80 | 54771.0703
80.9 | 148321.9531 | 81 | 148321.9531
81.9 | 27568.8027 | 82 | 27568.8027
82.9 | 4727.0908 | 83 | 4727.0908
83.7 | 469.3996 | 84 | 469.3996
84.3 | 166.1154 | 84 | 166.1154
84.9 | 239.14 | 85 | 239.14
86 | 13693.4814 | 86 | 13693.4814
87 | 805072.375 | 87 | 805072.375
88 | 758009.9375 | 88 | 758009.9375
88.8 | 15590.0449 | 89 | 15590.0449
90 | 1319.269 | 90 | 1319.269
90.8 | 21363.7285 | 91 | 21363.7285
92 | 366425.875 | 92 | 366425.875
93 | 557672.125 | 93 | 557672.125
94 | 1642019.25 | 94 | 1642019.25
95 | 20167550 | 95 | 20167550
96 | 1304864 | 96 | 1304864
97.1 | 39899.0195 | 97 | 39899.0195
98.2 | 2621.8098 | 98 | 2621.8098
99 | 437.947 | 99 | 437.947
99.9 | 75.7497 | 100 | 75.7497
100.5 | 110.0593 | 101 | 110.0593
101.2 | 405.3799 | 101 | 405.3799
103 | 5088.7705 | 103 | 5088.7705
103.9 | 43760.043 | 104 | 43760.043
105 | 16968.7324 | 105 | 16968.7324
105.9 | 31839.9238 | 106 | 31839.9238
106.9 | 11489.4014 | 107 | 11489.4014
107.9 | 800.0591 | 108 | 800.0591
108.6 | 0 | 109 | 0
109.9 | 2417.8706 | 110 | 2417.8706
110.9 | 5204.4111 | 111 | 5204.4111
112 | 3195.5088 | 112 | 3195.5088
113 | 3885.8215 | 113 | 3885.8215
114 | 812.368 | 114 | 812.368
114.9 | 11048.9961 | 115 | 11048.9961
115.9 | 30078.2695 | 116 | 30078.2695
116.9 | 65709.5781 | 117 | 65709.5781
118 | 39115.4023 | 118 | 39115.4023
118.9 | 54383.1914 | 119 | 54383.1914
120.1 | 1082.5376 | 120 | 1082.5376
120.9 | 1195.9094 | 121 | 1195.9094
122 | 3022.7151 | 122 | 3022.7151
122.9 | 2070.0476 | 123 | 2070.0476
123.9 | 8909.8896 | 124 | 8909.8896
124.9 | 3464.792 | 125 | 3464.792
126 | 4924.4473 | 126 | 4924.4473
126.9 | 3114.9712 | 127 | 3114.9712
128 | 40453.2461 | 128 | 40453.2461
129 | 20333.6387 | 129 | 20333.6387
129.9 | 43217.5977 | 130 | 43217.5977
130.9 | 12958.6221 | 131 | 12958.6221
132 | 2906.4207 | 132 | 2906.4207
133 | 1106.186 | 133 | 1106.186
133.8 | 1542.8763 | 134 | 1542.8763
134.9 | 17530.9219 | 135 | 17530.9219
135.9 | 2628.3567 | 136 | 2628.3567
136.9 | 16023.4658 | 137 | 16023.4658
138 | 820.2974 | 138 | 820.2974
138.9 | 3154.6577 | 139 | 3154.6577
139.9 | 7465.4639 | 140 | 7465.4639
140.9 | 92979.4922 | 141 | 92979.4922
141.9 | 6065.3379 | 142 | 6065.3379
142.9 | 88801.3203 | 143 | 88801.3203
144 | 4806.2515 | 144 | 4806.2515
144.9 | 8391.8467 | 145 | 8391.8467
145.9 | 20264.6895 | 146 | 20264.6895
146.9 | 7432.313 | 147 | 7432.313
147.9 | 32087.1523 | 148 | 32087.1523
149 | 9967.6582 | 149 | 9967.6582
149.9 | 8327.4287 | 150 | 8327.4287
151 | 743.8942 | 151 | 743.8942
151.9 | 4960.8345 | 152 | 4960.8345
153 | 9049.543 | 153 | 9049.543
153.9 | 8843.6885 | 154 | 8843.6885
154.9 | 38474.4648 | 155 | 38474.4648
156 | 4863.3208 | 156 | 4863.3208
156.9 | 25101.0371 | 157 | 25101.0371
158 | 2538.5598 | 158 | 2538.5598
158.9 | 12891.4092 | 159 | 12891.4092
160 | 1148.7529 | 160 | 1148.7529
160.8 | 10241.7295 | 161 | 10241.7295
161.9 | 738.0129 | 162 | 738.0129
162.4 | 282.2885 | 162 | 282.2885
164.4 | 289.1107 | 164 | 289.1107
166 | 596.0144 | 166 | 596.0144
167.1 | 343.9621 | 167 | 343.9621
167.7 | 395.5647 | 168 | 395.5647
168.3 | 88.2604 | 168 | 88.2604
169.8 | 816.5031 | 170 | 816.5031
171 | 1089.4939 | 171 | 1089.4939
172 | 25313.002 | 172 | 25313.002
173 | 111488.4297 | 173 | 111488.4297
173.9 | 15387236 | 174 | 15387236
175 | 1035747.1875 | 175 | 1035747.1875
175.9 | 15758876 | 176 | 15758876
177 | 977295.375 | 177 | 977295.375
178 | 19036.4746 | 178 | 19036.4746
178.9 | 1090.3656 | 179 | 1090.3656
180.3 | 0.053215 | 180 | 0.053215
181 | 2.1e-05 | 181 | 2.1e-05
182.3 | 1.6e-05 | 182 | 1.6e-05
183 | 0 | 183 | 0
183.6 | 0.003275 | 184 | 0.003275
185 | 91.7151 | 185 | 91.7151
186.1 | 0 | 186 | 0
186.8 | 0.010969 | 187 | 0.010969
187.5 | 8.8e-05 | 188 | 8.8e-05
188 | 0.10523 | 188 | 0.10523
190 | 173.1493 | 190 | 173.1493
191 | 127.1672 | 191 | 127.1672
192 | 5.1286 | 192 | 5.1286
192.9 | 1211.2577 | 193 | 1211.2577
194 | 325.5101 | 194 | 325.5101
195 | 214.2272 | 195 | 214.2272
196 | 0.16827 | 196 | 0.16827
196.9 | 63.6832 | 197 | 63.6832
197.6 | 1e-06 | 198 | 1e-06
198.4 | 0 | 198 | 0
199.3 | 0.024765 | 199 | 0.024765
200.6 | 61.0218 | 201 | 61.0218
201.2 | 0 | 201 | 0
202.2 | 0.080066 | 202 | 0.080066
202.7 | 0.061508 | 203 | 0.061508
203.4 | 37.5567 | 203 | 37.5567
204.7 | 243.2038 | 205 | 243.2038
206 | 0.33368 | 206 | 0.33368
207.1 | 5191.0786 | 207 | 5191.0786
208.1 | 691.5928 | 208 | 691.5928
209 | 165.0159 | 209 | 165.0159
209.9 | 871.4249 | 210 | 871.4249
210.7 | 112.7899 | 211 | 112.7899
212.2 | 0.000119 | 212 | 0.000119
213.6 | 189.1193 | 214 | 189.1193
215 | 171.9062 | 215 | 171.9062
215.8 | 2.4509 | 216 | 2.4509
216.8 | 60.1941 | 217 | 60.1941
218 | 16.1386 | 218 | 16.1386
218.8 | 0.070336 | 219 | 0.070336
219.5 | 0.088646 | 220 | 0.088646
220.2 | 35.7903 | 220 | 35.7903
221.1 | 0.003274 | 221 | 0.003274
222.5 | 78.9653 | 223 | 78.9653
223.8 | 0 | 224 | 0
224.5 | 0 | 225 | 0
225.6 | 4.3117 | 226 | 4.3117
226.2 | 0.00262 | 226 | 0.00262
227 | 0.000254 | 227 | 0.000254
227.7 | 0 | 228 | 0
228.8 | 0.070316 | 229 | 0.070316
229.4 | 0.011002 | 229 | 0.011002
230.3 | 1.5e-05 | 230 | 1.5e-05
230.9 | 0 | 231 | 0
232 | 99.9201 | 232 | 99.9201
232.5 | 0.019425 | 233 | 0.019425
233.1 | 29.8289 | 233 | 29.8289
234.4 | 1e-06 | 234 | 1e-06
234.9 | 94.4587 | 235 | 94.4587
235.6 | 0.033521 | 236 | 0.033521
236.2 | 2.0714 | 236 | 2.0714
236.7 | 162.7879 | 237 | 162.7879
238.1 | 41.6433 | 238 | 41.6433
239 | 146.718 | 239 | 146.718
240 | 0 | 240 | 0
240.7 | 0 | 241 | 0
241.2 | 0 | 241 | 0
241.8 | 0 | 242 | 0
243.5 | 1.3e-05 | 244 | 1.3e-05
244.1 | 0 | 244 | 0
245 | 2.5e-05 | 245 | 2.5e-05
245.6 | 0 | 246 | 0
246.3 | 0.097737 | 246 | 0.097737
247.4 | 6.6107 | 247 | 6.6107
248.3 | 7e-06 | 248 | 7e-06
248.8 | 282.8799 | 249 | 282.8799
250.1 | 191.27 | 250 | 191.27
250.8 | 63.8807 | 251 | 63.8807
251.7 | 0.1908 | 252 | 0.1908
252.2 | 427.5085 | 252 | 427.5085
252.9 | 0.000724 | 253 | 0.000724
253.6 | 0.10696 | 254 | 0.10696
254.8 | 0 | 255 | 0
256.4 | 0.016663 | 256 | 0.016663
256.9 | 9e-06 | 257 | 9e-06
258 | 104.332 | 258 | 104.332
258.9 | 1e-06 | 259 | 1e-06
259.9 | 0.012631 | 260 | 0.012631
252.4 | 0 | 252 | 0
253.7 | 17.7301 | 254 | 17.7301
254.6 | 0 | 255 | 0
255.9 | 0.011016 | 256 | 0.011016
256.8 | 0 | 257 | 0
258.1 | 0
259.1 | 109.026
259.9 | 229.5571
257.3 | 0.004411
258.3 | 181.4352
259 | 0.043369
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[12 empty] | TV | [6 empty] | Measured | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
Name | RT | Window | Quant | Conf1 | Conf2 | Window |  | Channel | Eval. Type | IS | Cal.Type | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | RSD<20% | R^2>0.99
Chloromethane (methyl chloride) | 1.464 | 0.050 AN | 50 | 52 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Chloroethene (vinyl chloride) | 1.553 | 0.050 AN | 62 | 64 | 61 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 |  | ppb | [18 empty] | TRUE | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Trichlorofluoromethane | 2.182 | 0.050 AN | 101 | 103 | 105 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Diethyl ether | 2.505 | 0.080 AN | 59 | 74 | 45 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Acetone | 2.852 | 0.088 AN | 43 | 58 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1.8 | 9 | 18 | 36 | 90 | 180 | ppb | [18 empty] | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Carbon disulfide | 2.961 | 0.050 AN | 76 | 78 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
trans-1,2-Dichloroethene | 3.678 | 0.100 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
2,2-Dichloropropane | 4.819 | 0.060 AN | 77 | 41 | 79 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
cis-1,2-Dichloroethene | 4.823 | 0.060 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
2-Butanone (MEK) | 4.851 | 0.060 AN | 43 | 72 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1.8 | 9 | 18 | 36 | 90 | 180 | ppb | [18 empty] | TRUE | TRUE
Methyl acrylate | 4.939 | 0.060 AN | 55 | 85 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Bromochloromethane | 5.073 | 0.060 AN | 49 | 130 | 128 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Methacrylonitrile | 5.076 | 0.060 AN | 67 | 52 | 40 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Tetrahydrofuran | 5.09 | 0.060 AN | 42 | 72 | 71 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Trichloromethane (chloroform) | 5.198 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.476 | 0.040 AN | 56 | 41 | 43 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Carbon tetrachloride | 5.493 | 0.040 AN | 119 | 121 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.167 | 0.080 AN | 114 | 88 | 63 | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Dibromomethane | 6.715 | 0.060 AN | 174 | 93 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.518 | 0.049 AN | 43 | 58 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1.8 | 9 | 18 | 36 | 90 | 180 | ppb | [18 empty] | TRUE | TRUE
Toluene-d8 [SS2] | 7.607 | 0.058 AN | 98 | 100 | 70 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
Toluene | 7.67 | 0.047 AN | 91 | 92 | 65 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
trans-1,3-Dichloropropene | 7.92 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Ethyl methacrylate | 7.99 | 0.060 AN | 69 | 41 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1,2-Trichloroethane | 8.098 | 0.060 AN | 97 | 83 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,3-Dichloropropane | 8.24 | 0.060 AN | 76 | 41 | 78 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1.8 | 9 | 18 | 36 | 90 | 180 | ppb | [18 empty] | TRUE | TRUE
Dibromochloromethane | 8.418 | 0.050 AN | 129 | 127 | 131 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.517 | 0.049 AN | 107 | 109 | 93 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.915 | 0.040 AF | 117 | 82 | 52 | 0.1 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
Chlorobenzene | 8.937 | 0.040 AN | 112 | 77 | 114 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.01 | 0.040 AN | 131 | 133 | 117 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Ethylbenzene | 9.018 | 0.040 AN | 91 | 106 | 51 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
m/p-Xylene | 9.125 | 0.060 AN | 91 | 106 | 105 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
o-Xylene | 9.423 | 0.060 AN | 91 | 106 | 105 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Styrene | 9.439 | 0.060 AN | 104 | 78 | 103 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Isopropylbenzene (cumene) | 9.726 | 0.060 AN | 105 | 120 | 79 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.835 | 0.060 AN | 95 | 174 | 176 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
Bromobenzene | 9.936 | 0.040 AN | 77 | 156 | 158 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.944 | 0.040 AN | 83 | 85 | 95 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.969 | 0.040 AN | 75 | 53 | 89 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.975 | 0.040 AN | 77 | 110 | 61 | 0.05 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
n-Propylbenzene | 10.016 | 0.026 AN | 91 | 120 | 65 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
2-Chlorotoluene | 10.071 | 0.031 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.154 | 0.040 AN | 105 | 120 | 119 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
4-Chlorotoluene | 10.168 | 0.040 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
tert-Butylbenzene | 10.357 | 0.040 AN | 119 | 91 | 134 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Pentachloroethane | 10.381 | 0.030 AN | 167 | 130 | 132 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.411 | 0.030 AN | 105 | 120 | 77 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.521 | 0.037 AN | 105 | 134 | 91 | 0.08 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,3-Dichlorobenzene | 10.603 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.656 | 0.036 AN | 152 | 115 |  | 0.07 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | [18 empty] | TRUE | TRUE
1,4-Dichlorobenzene | 10.679 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
n-Butylbenzene | 10.9 | 0.040 AN | 91 | 92 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2-Dichlorobenzene | 10.914 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Hexachloroethane | 11.093 | 0.040 AN | 117 | 119 | 201 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.447 | 0.040 AN | 157 | 155 | 75 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Nitrobenzene | 11.566 | 0.040 AN | 77 | 51 | 123 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.966 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.066 | 0.040 AN | 225 | 227 | 223 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.273 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 5 | 10 | 20 | 50 | 100 | ppb | [18 empty] | TRUE | TRUE

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.45 | 318 | 0.01 | 0.069 | Not confirmed | 50 | 52 | 32.43 | 47.11 | Confirmed | 49 | 9.42 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 31.79 | n.a. | Not confirmed | 61 | 7.66 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.83 | 2747 | 0.07 | 0.218 | Confirmed | 94 | 96 | 94.19 | 80.57 | Confirmed | 93 | 20.21 | 25.04 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 31.92 | n.a. | Not confirmed | 49 | 22.26 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | 2.18 | 29 | 0 | 0.005 | Not confirmed | 101 | 103 | 64.79 | n.a. | Not confirmed | 105 | 10.31 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 78.08 | n.a. | Not confirmed | 45 | 69.16 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | 2.74 | 68 | 0 | 0.01 | Not confirmed | 61 | 96 | 72.34 | n.a. | Not confirmed | 98 | 45.55 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 42.06 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | FALSE | FALSE | TRUE | Iodomethane | 2.9 | 1674 | 0.04 | 0.559 | Confirmed | 142 | 127 | 29.81 | 27.38 | Confirmed | 141 | 12.12 | 6.28 | Confirmed
1 | TRUE | TRUE | FALSE | Carbon disulfide | 2.95 | 3061 | 0.08 | 0.087 | Not confirmed | 76 | 78 | 8.34 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 50.8 | n.a. | Not confirmed | 76 | 40.18 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Methylene chloride (DCM) | 3.35 | 842 | 0.02 | 0.058 | Confirmed | 49 | 84 | 91.76 | 100.95 | Confirmed | 86 | 58.92 | 60.36 | Confirmed
1 | TRUE | FALSE | FALSE | trans-1,2-Dichloroethene | 3.68 | 318 | 0.01 | 0.02 | Confirmed | 61 | 96 | 74.28 | 58.61 | Confirmed | 98 | 46.43 | 59.54 | Confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 24.57 | n.a. | Not confirmed | 57 | 21.97 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 65 | 31.62 | n.a. | Not confirmed | 83 | 12.12 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 68.94 | n.a. | Not confirmed | 79 | 31.62 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,2-Dichloroethene | 4.83 | 126 | 0 | 0.007 | Not confirmed | 61 | 96 | 78.03 | n.a. | Not confirmed | 98 | 49.32 | 64.58 | Confirmed
1.8 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 30.88 | n.a. | Not confirmed | 57 | 8.64 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 16.59 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Bromochloromethane | 5.07 | 126 | 0 | 0.012 | Confirmed | 49 | 130 | 85.41 | 97.09 | Confirmed | 128 | 66.39 | 83.27 | Confirmed
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 30.8 | n.a. | Not confirmed | 40 | 36.05 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 45.8 | n.a. | Not confirmed | 71 | 45.7 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 64.22 | n.a. | Not confirmed | 47 | 16.92 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 99 | 63.11 | n.a. | Not confirmed | 61 | 41.54 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.36 | 200195 | 5.27 | 19.625 | Confirmed | 113 | 111 | 102.64 | 104.35 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.43 | 409442 | 10.78 | 20 | Confirmed | 168 | 99 | 50.71 | 50.04 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 56 | 41 | 53.9 | n.a. | Not confirmed | 43 | 24.71 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon tetrachloride | n.a. | n.a. | n.a. | n.a. | Not confirmed | 119 | 121 | 31.42 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,1-Dichloropropene | 5.51 | 384 | 0.01 | 0.056 | Not confirmed | 75 | 77 | 31.18 | n.a. | Not confirmed | 110 | 40.06 | 25.78 | Confirmed
1 | TRUE | TRUE | FALSE | Benzene | 5.71 | 368 | 0.01 | 0.007 | Not confirmed | 78 | 77 | 23.98 | n.a. | Not confirmed | 52 | 14.15 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 31.94 | n.a. | Not confirmed | 49 | 28.93 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.17 | 654250 | 17.23 | 20 | Confirmed | 114 | 88 | 18.07 | 18.19 | Confirmed | 63 | 18.4 | 18.45 | Confirmed
1 | TRUE | TRUE | FALSE | Trichloroethene | 6.38 | 231 | 0.01 | 0.018 | Not confirmed | 130 | 132 | 95.47 | 60.58 | Not confirmed | 95 | 95.99 | 93.94 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 68.71 | n.a. | Not confirmed | 41 | 38.08 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.72 | 68 | 0 | 0.008 | Not confirmed | 174 | 93 | 92.51 | 141.76 | Not confirmed | 95 | 80.32 | 129.39 | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 97.5 | n.a. | Not confirmed | 39 | 43.16 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 63.78 | n.a. | Not confirmed | 47 | 14.61 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 87.57 | n.a. | Not confirmed | 39 | 27.3 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 39 | 36.74 | n.a. | Not confirmed | 77 | 31.57 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 43.17 | n.a. | Not confirmed | 41 | 22.65 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.61 | 853253 | 22.47 | 20.071 | Confirmed | 98 | 100 | 63.49 | 63.9 | Confirmed | 70 | 11.11 | 11.36 | Confirmed
1 | TRUE | FALSE | FALSE | Toluene | 7.68 | 1495 | 0.04 | 0.027 | Confirmed | 91 | 92 | 56.84 | 65.6 | Confirmed | 65 | 11.51 | 11.17 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,3-Dichloropropene | 7.92 | 162 | 0 | 0.02 | Not confirmed | 75 | 39 | 36.89 | n.a. | Not confirmed | 77 | 31.71 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 69 | 41 | 56.16 | n.a. | Not confirmed | 99 | 24.48 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 83 | 87.07 | n.a. | Not confirmed | 99 | 62.03 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.15 | 318 | 0.01 | 0.021 | Confirmed | 166 | 164 | 78.37 | 83.88 | Confirmed | 129 | 72.11 | 71.26 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3-Dichloropropane | 8.24 | 96 | 0 | 0.005 | Not confirmed | 76 | 41 | 56.65 | n.a. | Not confirmed | 78 | 31.88 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 60.85 | n.a. | Not confirmed | 57 | 21.01 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 129 | 127 | 77.06 | n.a. | Not confirmed | 131 | 23.53 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 94.97 | n.a. | Not confirmed | 93 | 4.29 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.91 | 612482 | 16.13 | 20 | Confirmed | 117 | 82 | 58.75 | 59.19 | Confirmed | 52 | 14.04 | 13.23 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.93 | 646 | 0.02 | 0.018 | Not confirmed | 112 | 77 | 64.75 | 601.57 | Not confirmed | 114 | 31.51 | 22.45 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 131 | 133 | 94.33 | n.a. | Not confirmed | 117 | 80.87 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Ethylbenzene | 9.02 | 2117 | 0.06 | 0.039 | Confirmed | 91 | 106 | 35.95 | 29.79 | Confirmed | 51 | 8.38 | 13.3 | Confirmed
1 | TRUE | FALSE | FALSE | m/p-Xylene | 9.12 | 4704 | 0.12 | 0.053 | Confirmed | 91 | 106 | 52.8 | 53.42 | Confirmed | 105 | 21.46 | 22.29 | Confirmed
1 | TRUE | FALSE | FALSE | o-Xylene | 9.43 | 2250 | 0.06 | 0.046 | Confirmed | 91 | 106 | 50.24 | 50.54 | Confirmed | 105 | 25.22 | 26.53 | Confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.45 | 1285 | 0.03 | 0.033 | Confirmed | 104 | 78 | 54.13 | 60.83 | Confirmed | 103 | 53 | 37.94 | Confirmed
1 | TRUE | FALSE | FALSE | Bromoform | 9.57 | 54 | 0 | 0.007 | Confirmed | 173 | 171 | 51.65 | 57.04 | Confirmed | 175 | 48.32 | 50.3 | Confirmed
1 | TRUE | TRUE | FALSE | Isopropylbenzene (cumene) | 9.71 | 805 | 0.02 | 0.03 | Not confirmed | 105 | 120 | 29.46 | 31.65 | Confirmed | 79 | 15.73 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 319134 | 8.4 | 21.947 | Confirmed | 95 | 174 | 77.01 | 76.86 | Confirmed | 176 | 74.73 | 74.32 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | 9.94 | 614 | 0.02 | 0.112 | Not confirmed | 77 | 110 | 75.42 | n.a. | Not confirmed | 61 | 55.64 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.94 | 614 | 0.02 | 0.026 | Confirmed | 77 | 156 | 61.5 | 58.74 | Confirmed | 158 | 59.81 | 61.19 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 64.94 | n.a. | Not confirmed | 95 | 13.61 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | trans-1,4-Dichloro-2-butene | 9.97 | 258 | 0.01 | 0.016 | Not confirmed | 75 | 53 | 18.06 | n.a. | Not confirmed | 89 | 11.39 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | n-Propylbenzene | 10.02 | 1533 | 0.04 | 0.025 | Not confirmed | 91 | 120 | 26.68 | 26.89 | Confirmed | 65 | 10.18 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 2-Chlorotoluene | 10.07 | 680 | 0.02 | 0.018 | Confirmed | 91 | 126 | 37.43 | 50.87 | Confirmed | 89 | 17.41 | 28.79 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3,5-Trimethylbenzene | 10.15 | 1013 | 0.03 | 0.021 | Not confirmed | 105 | 120 | 50.57 | 53.4 | Confirmed | 119 | 11.57 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 4-Chlorotoluene | 10.16 | 1276 | 0.03 | 0.03 | Confirmed | 91 | 126 | 33.24 | 34.45 | Confirmed | 89 | 11.37 | 9.92 | Confirmed
1 | TRUE | FALSE | FALSE | tert-Butylbenzene | 10.37 | 656 | 0.02 | 0.031 | Confirmed | 119 | 91 | 64.28 | 79.59 | Confirmed | 134 | 23.78 | 25.33 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 167 | 130 | 53.76 | n.a. | Not confirmed | 132 | 53.97 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,4-Trimethylbenzene | 10.41 | 1376 | 0.04 | 0.053 | Not confirmed | 105 | 120 | 47.98 | 52.83 | Confirmed | 77 | 11.1 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.52 | 1296 | 0.03 | 0.046 | Not confirmed | 105 | 134 | 21.44 | 22.36 | Confirmed | 91 | 14.93 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.6 | 1056 | 0.03 | 0.041 | Confirmed | 146 | 148 | 63.15 | 68.09 | Confirmed | 111 | 41.81 | 47.06 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.63 | 1138 | 0.03 | 0.048 | Confirmed | 119 | 91 | 26.67 | 33.49 | Confirmed | 134 | 29.6 | 27.23 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 312284 | 8.22 | 20 | Confirmed | 152 | 115 | 61.11 | 61.72 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.67 | 1386 | 0.04 | 0.053 | Not confirmed | 146 | 148 | 63.73 | 102.23 | Not confirmed | 111 | 44.46 | 265.08 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.91 | 1927 | 0.05 | 0.087 | Confirmed | 91 | 92 | 52.46 | 48.39 | Confirmed | 134 | 29.07 | 27.38 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.92 | 975 | 0.03 | 0.036 | Confirmed | 146 | 148 | 62.68 | 70.05 | Confirmed | 111 | 42.47 | 40.68 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 98.01 | n.a. | Not confirmed | 201 | 89.25 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 157 | 155 | 76.98 | n.a. | Not confirmed | 75 | 88.27 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 44.57 | n.a. | Not confirmed | 123 | 47.29 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.97 | 1653 | 0.04 | 0.134 | Confirmed | 180 | 182 | 95.79 | 87.86 | Confirmed | 145 | 32.18 | 29.15 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.06 | 472 | 0.01 | 0.105 | Confirmed | 225 | 227 | 63.97 | 63.67 | Confirmed | 223 | 63.08 | 63.21 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.14 | 3611 | 0.1 | 0.094 | Confirmed | 128 | 127 | 12.79 | 11.59 | Confirmed | 129 | 10.56 | 10.53 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.28 | 1525 | 0.04 | 0.097 | Confirmed | 180 | 182 | 94.54 | 80.76 | Confirmed | 145 | 33.51 | 29.04 | Confirmed
*Iodomethane failure due to Optima methanol used in needle rinse - will be switched to P&T grade

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 1.8
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 1.8
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 1
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | n.a.
Pentafluorobenzene [IS1] | n.a.
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | n.a.
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 1
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 1.8
Toluene-d8 [SS2] | n.a.
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 1.8
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | n.a.
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 1
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | n.a.
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | n.a.
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 1
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## CCV1
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? | <70% highlighted |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 426503 | 5.43 | 448577 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.17 | 678596 | 6.17 | 756018 | TRUE | TRUE | 90.22 | 10 | TRUE | Chloromethane (methyl chloride) | 1.46 | 38904 | 0.11 | 9.022 | Confirmed | 50 | 52 | 32.43 | 31.88 | Confirmed | 49 | 9.42 | 9.63 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 646689 | 8.92 | 726359 | TRUE | TRUE | 88.71 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.55 | 75516 | 0.21 | 8.871 | Confirmed | 62 | 64 | 31.79 | 32.49 | Confirmed | 61 | 7.66 | 7.14 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 366888 | 10.66 | 432454 | TRUE | TRUE | 97.46 | 10 | TRUE | Bromomethane (methyl bromide) | 1.83 | 128342 | 0.36 | 9.746 | Confirmed | 94 | 96 | 94.19 | 96.34 | Confirmed | 93 | 20.21 | 20.71 | Confirmed
[7 empty] | 88.16 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.95 | 68284 | 0.19 | 8.816 | Confirmed | 64 | 66 | 31.92 | 32.11 | Confirmed | 49 | 22.26 | 20.67 | Confirmed
[7 empty] | 83.3 | 10 | TRUE | Trichlorofluoromethane | 2.19 | 99235 | 0.28 | 8.33 | Confirmed | 101 | 103 | 64.79 | 65.3 | Confirmed | 105 | 10.31 | 9.62 | Confirmed
Total Analytes | 81 | [5 empty] | 88.01 | 10 | TRUE | Diethyl ether | 2.5 | 96416 | 0.27 | 8.801 | Confirmed | 59 | 74 | 78.08 | 76.33 | Confirmed | 45 | 69.16 | 68.57 | Confirmed
Failed | 1 | [5 empty] | 85.5 | 10 | TRUE | 1,1-Dichloroethene | 2.74 | 115313 | 0.33 | 8.55 | Confirmed | 61 | 96 | 72.34 | 72.81 | Confirmed | 98 | 45.55 | 46.61 | Confirmed
[7 empty] | 87.8222 | 18 | TRUE | Acetone | 2.82 | 56545 | 0.16 | 15.808 | Confirmed | 43 | 58 | 42.06 | 35.98 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 81.43 | 10 | TRUE | Iodomethane | 2.89 | 69436 | 0.2 | 8.143 | Confirmed | 142 | 127 | 29.81 | 30.93 | Confirmed | 141 | 12.12 | 12.62 | Confirmed
[7 empty] | 85.78 | 10 | TRUE | Carbon disulfide | 2.96 | 280805 | 0.8 | 8.578 | Confirmed | 76 | 78 | 8.34 | 8.74 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 85.78 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.2 | 129918 | 0.37 | 8.578 | Confirmed | 41 | 39 | 50.8 | 52.05 | Confirmed | 76 | 40.18 | 41.6 | Confirmed
[7 empty] | 87.38 | 10 | TRUE | Methylene chloride (DCM) | 3.36 | 130944 | 0.37 | 8.738 | Confirmed | 49 | 84 | 91.76 | 93.01 | Confirmed | 86 | 58.92 | 59.29 | Confirmed
[7 empty] | 89.87 | 10 | TRUE | trans-1,2-Dichloroethene | 3.68 | 130335 | 0.37 | 8.987 | Confirmed | 61 | 96 | 74.28 | 72.98 | Confirmed | 98 | 46.43 | 46.77 | Confirmed
[7 empty] | 84.89 | 10 | TRUE | Methyl tert-butyl ether (MTBE) | 3.69 | 267858 | 0.76 | 8.489 | Confirmed | 73 | 41 | 24.57 | 24.57 | Confirmed | 57 | 21.97 | 22.6 | Confirmed
[7 empty] | 92.28 | 10 | TRUE | 1,1-Dichloroethane | 4.19 | 177869 | 0.51 | 9.228 | Confirmed | 63 | 65 | 31.62 | 31.52 | Confirmed | 83 | 12.12 | 12.03 | Confirmed
[7 empty] | 71.05 | 10 | FALSE | 2,2-Dichloropropane | 4.82 | 72393 | 0.21 | 7.105 | Confirmed | 77 | 41 | 68.94 | 70.34 | Confirmed | 79 | 31.62 | 32.07 | Confirmed
[7 empty] | 89.75 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 147818 | 0.42 | 8.975 | Confirmed | 61 | 96 | 78.03 | 77.76 | Confirmed | 98 | 49.32 | 49.37 | Confirmed
[7 empty] | 82.6833 | 18 | TRUE | 2-Butanone (MEK) | 4.83 | 87253 | 0.25 | 14.883 | Confirmed | 43 | 72 | 30.88 | 27.38 | Confirmed | 57 | 8.64 | 8.46 | Confirmed
[7 empty] | 80.17 | 10 | TRUE | Methyl acrylate | 4.93 | 86041 | 0.24 | 8.017 | Confirmed | 55 | 85 | 16.59 | 17.47 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 84.07 | 10 | TRUE | Methacrylonitrile | 5.06 | 54852 | 0.16 | 8.407 | Confirmed | 67 | 52 | 30.8 | 31.32 | Confirmed | 40 | 36.05 | 34.18 | Confirmed
[7 empty] | 90.1 | 10 | TRUE | Bromochloromethane | 5.06 | 91337 | 0.26 | 9.01 | Confirmed | 49 | 130 | 85.41 | 89.48 | Confirmed | 128 | 66.39 | 69.59 | Confirmed
[7 empty] | 84.32 | 10 | TRUE | Tetrahydrofuran | 5.08 | 37727 | 0.11 | 8.432 | Confirmed | 42 | 72 | 45.8 | 46.48 | Confirmed | 71 | 45.7 | 46.35 | Confirmed
[7 empty] | 92.94 | 10 | TRUE | Trichloromethane (chloroform) | 5.2 | 171347 | 0.49 | 9.294 | Confirmed | 83 | 85 | 64.22 | 65.9 | Confirmed | 47 | 16.92 | 16.59 | Confirmed
[7 empty] | 97.85 | 10 | TRUE | 1,1,1-Trichloroethane | 5.33 | 111669 | 0.32 | 9.785 | Confirmed | 97 | 99 | 63.11 | 63.08 | Confirmed | 61 | 41.54 | 40.07 | Confirmed
[7 empty] | 94.65 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.36 | 201155 | 0.57 | 18.93 | Confirmed | 113 | 111 | 102.64 | 104.11 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 426503 | 1.21 | 20 | Confirmed | 168 | 99 | 50.71 | 48.13 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 92.04 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 156122 | 0.44 | 9.204 | Confirmed | 56 | 41 | 53.9 | 54.16 | Confirmed | 43 | 24.71 | 25.34 | Confirmed
[7 empty] | 90.88 | 10 | TRUE | Carbon tetrachloride | 5.49 | 81623 | 0.23 | 9.088 | Confirmed | 119 | 121 | 31.42 | 32.35 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 92.93 | 10 | TRUE | 1,1-Dichloropropene | 5.51 | 114262 | 0.32 | 9.293 | Confirmed | 75 | 77 | 31.18 | 31.04 | Confirmed | 110 | 40.06 | 41.5 | Confirmed
[7 empty] | 97.61 | 10 | TRUE | Benzene | 5.7 | 446777 | 1.27 | 9.761 | Confirmed | 78 | 77 | 23.98 | 23.91 | Confirmed | 52 | 14.15 | 13.87 | Confirmed
[7 empty] | 93.44 | 10 | TRUE | 1,2-Dichloroethane | 5.77 | 121276 | 0.34 | 9.344 | Confirmed | 62 | 64 | 31.94 | 32.13 | Confirmed | 49 | 28.93 | 29.95 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.17 | 678596 | 1.93 | 20 | Confirmed | 114 | 88 | 18.07 | 18.09 | Confirmed | 63 | 18.4 | 17.96 | Confirmed
[7 empty] | 101.71 | 10 | TRUE | Trichloroethene | 6.38 | 119004 | 0.34 | 10.171 | Confirmed | 130 | 132 | 95.47 | 96.73 | Confirmed | 95 | 95.99 | 91.57 | Confirmed
[7 empty] | 97.78 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 114838 | 0.33 | 9.778 | Confirmed | 63 | 62 | 68.71 | 68.74 | Confirmed | 41 | 38.08 | 37.09 | Confirmed
[7 empty] | 102.47 | 10 | TRUE | Dibromomethane | 6.72 | 79840 | 0.23 | 10.247 | Confirmed | 174 | 93 | 92.51 | 92.49 | Confirmed | 95 | 80.32 | 78.4 | Confirmed
[7 empty] | 91.46 | 10 | TRUE | Methyl methacrylate | 6.74 | 73378 | 0.21 | 9.146 | Confirmed | 41 | 69 | 97.5 | 97.81 | Confirmed | 39 | 43.16 | 43.19 | Confirmed
[7 empty] | 98.23 | 10 | TRUE | Bromodichloromethane | 6.92 | 123502 | 0.35 | 9.823 | Confirmed | 83 | 85 | 63.78 | 64.59 | Confirmed | 47 | 14.61 | 14.33 | Confirmed
[7 empty] | 88.59 | 10 | TRUE | 2-Nitropropane | 7.14 | 19711 | 0.06 | 8.859 | Confirmed | 43 | 41 | 87.57 | 83.34 | Confirmed | 39 | 27.3 | 26.43 | Confirmed
[7 empty] | 95.46 | 10 | TRUE | cis-1,3-Dichloropropene | 7.35 | 137578 | 0.39 | 9.546 | Confirmed | 75 | 39 | 36.74 | 35.97 | Confirmed | 77 | 31.57 | 31.37 | Confirmed
[7 empty] | 91.2389 | 18 | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.51 | 197440 | 0.56 | 16.423 | Confirmed | 43 | 58 | 43.17 | 42.03 | Confirmed | 41 | 22.65 | 22.38 | Confirmed
[7 empty] | 97.14 | 20 | TRUE | Toluene-d8 [SS2] | 7.6 | 872062 | 2.48 | 19.428 | Confirmed | 98 | 100 | 63.49 | 63.79 | Confirmed | 70 | 11.11 | 11.08 | Confirmed
[7 empty] | 99.63 | 10 | TRUE | Toluene | 7.67 | 466483 | 1.33 | 9.963 | Confirmed | 91 | 92 | 56.84 | 56.05 | Confirmed | 65 | 11.51 | 10.92 | Confirmed
[7 empty] | 95 | 10 | TRUE | trans-1,3-Dichloropropene | 7.92 | 105746 | 0.3 | 9.5 | Confirmed | 75 | 39 | 36.89 | 35.76 | Confirmed | 77 | 31.71 | 31.38 | Confirmed
[7 empty] | 96.73 | 10 | TRUE | Ethyl methacrylate | 7.99 | 134714 | 0.38 | 9.673 | Confirmed | 69 | 41 | 56.16 | 55.42 | Confirmed | 99 | 24.48 | 23.71 | Confirmed
[7 empty] | 96.75 | 10 | TRUE | 1,1,2-Trichloroethane | 8.1 | 104780 | 0.3 | 9.675 | Confirmed | 97 | 83 | 87.07 | 88.06 | Confirmed | 99 | 62.03 | 63.76 | Confirmed
[7 empty] | 104.7 | 10 | TRUE | Tetrachloroethene | 8.15 | 168826 | 0.48 | 10.47 | Confirmed | 166 | 164 | 78.37 | 77.63 | Confirmed | 129 | 72.11 | 72.67 | Confirmed
[7 empty] | 97.74 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 177659 | 0.5 | 9.774 | Confirmed | 76 | 41 | 56.65 | 54.92 | Confirmed | 78 | 31.88 | 32.4 | Confirmed
[7 empty] | 91.3778 | 18 | TRUE | 2-Hexanone | 8.31 | 135210 | 0.38 | 16.448 | Confirmed | 43 | 58 | 60.85 | 59.57 | Confirmed | 57 | 21.01 | 20.96 | Confirmed
[7 empty] | 99.62 | 10 | TRUE | Dibromochloromethane | 8.42 | 98150 | 0.28 | 9.962 | Confirmed | 129 | 127 | 77.06 | 77.93 | Confirmed | 131 | 23.53 | 24.17 | Confirmed
[7 empty] | 100.61 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.51 | 104812 | 0.3 | 10.061 | Confirmed | 107 | 109 | 94.97 | 93.37 | Confirmed | 93 | 4.29 | 4.38 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 646689 | 1.84 | 20 | Confirmed | 117 | 82 | 58.75 | 55.63 | Confirmed | 52 | 14.04 | 12.98 | Confirmed
[7 empty] | 97.88 | 10 | TRUE | Chlorobenzene | 8.93 | 303664 | 0.86 | 9.788 | Confirmed | 112 | 77 | 64.75 | 61.17 | Confirmed | 114 | 31.51 | 31.44 | Confirmed
[7 empty] | 98.46 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 88823 | 0.25 | 9.846 | Confirmed | 131 | 133 | 94.33 | 92.68 | Confirmed | 117 | 80.87 | 75.61 | Confirmed
[7 empty] | 100.02 | 10 | TRUE | Ethylbenzene | 9.02 | 486827 | 1.38 | 10.002 | Confirmed | 91 | 106 | 35.95 | 35.62 | Confirmed | 51 | 8.38 | 8.29 | Confirmed
[7 empty] | 99.84 | 10 | TRUE | m/p-Xylene | 9.13 | 834697 | 2.37 | 9.984 | Confirmed | 91 | 106 | 52.8 | 51.66 | Confirmed | 105 | 21.46 | 21.42 | Confirmed
[7 empty] | 101.1 | 10 | TRUE | o-Xylene | 9.42 | 446103 | 1.27 | 10.11 | Confirmed | 91 | 106 | 50.24 | 50.07 | Confirmed | 105 | 25.22 | 25.46 | Confirmed
[7 empty] | 98.1 | 10 | TRUE | Styrene | 9.44 | 366855 | 1.04 | 9.81 | Confirmed | 104 | 78 | 54.13 | 53.3 | Confirmed | 103 | 53 | 53.52 | Confirmed
[7 empty] | 100.29 | 10 | TRUE | Bromoform | 9.57 | 68214 | 0.19 | 10.029 | Confirmed | 173 | 171 | 51.65 | 51.62 | Confirmed | 175 | 48.32 | 49.29 | Confirmed
[7 empty] | 100.86 | 10 | TRUE | Isopropylbenzene (cumene) | 9.7 | 480625 | 1.37 | 10.086 | Confirmed | 105 | 120 | 29.46 | 30.24 | Confirmed | 79 | 15.73 | 15.85 | Confirmed
[7 empty] | 97.97 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.83 | 334740 | 0.95 | 19.594 | Confirmed | 95 | 174 | 77.01 | 76.4 | Confirmed | 176 | 74.73 | 76.01 | Confirmed
[7 empty] | 101.52 | 10 | TRUE | Bromobenzene | 9.93 | 216746 | 0.62 | 10.152 | Confirmed | 77 | 156 | 61.5 | 64.83 | Confirmed | 158 | 59.81 | 62.43 | Confirmed
[7 empty] | 96.7 | 10 | TRUE | 1,1,2,2-Tetrachloroethane | 9.94 | 130719 | 0.37 | 9.67 | Confirmed | 83 | 85 | 64.94 | 65.05 | Confirmed | 95 | 13.61 | 13.57 | Confirmed
[7 empty] | 98.39 | 10 | TRUE | 1,2,3-Trichloropropane (TCP) | 9.98 | 52422 | 0.15 | 9.839 | Confirmed | 77 | 110 | 75.42 | 78.77 | Confirmed | 61 | 55.64 | 56.24 | Confirmed
[7 empty] | 103.72 | 10 | TRUE | trans-1,4-Dichloro-2-butene | 9.98 | 161141 | 0.46 | 10.372 | Confirmed | 75 | 53 | 18.06 | 16.61 | Confirmed | 89 | 11.39 | 10.79 | Confirmed
[7 empty] | 101.9 | 10 | TRUE | n-Propylbenzene | 10.01 | 577631 | 1.64 | 10.19 | Confirmed | 91 | 120 | 26.68 | 26.3 | Confirmed | 65 | 10.18 | 9.95 | Confirmed
[7 empty] | 101.59 | 10 | TRUE | 2-Chlorotoluene | 10.07 | 366590 | 1.04 | 10.159 | Confirmed | 91 | 126 | 37.43 | 37.25 | Confirmed | 89 | 17.41 | 17.26 | Confirmed
[7 empty] | 104.45 | 10 | TRUE | 1,3,5-Trimethylbenzene | 10.15 | 466840 | 1.33 | 10.445 | Confirmed | 105 | 120 | 50.57 | 50.14 | Confirmed | 119 | 11.57 | 11.75 | Confirmed
[7 empty] | 103.42 | 10 | TRUE | 4-Chlorotoluene | 10.16 | 441674 | 1.26 | 10.342 | Confirmed | 91 | 126 | 33.24 | 32.56 | Confirmed | 89 | 11.37 | 11.05 | Confirmed
[7 empty] | 99.29 | 10 | TRUE | tert-Butylbenzene | 10.37 | 402078 | 1.14 | 9.929 | Confirmed | 119 | 91 | 64.28 | 65.64 | Confirmed | 134 | 23.78 | 23.97 | Confirmed
[7 empty] | 83.65 | 10 | TRUE | Pentachloroethane | 10.39 | 27786 | 0.08 | 8.365 | Confirmed | 167 | 130 | 53.76 | 54.46 | Confirmed | 132 | 53.97 | 57.84 | Confirmed
[7 empty] | 103.6 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.41 | 476137 | 1.35 | 10.36 | Confirmed | 105 | 120 | 47.98 | 48.3 | Confirmed | 77 | 11.1 | 11.18 | Confirmed
[7 empty] | 101.2 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.52 | 549772 | 1.56 | 10.12 | Confirmed | 105 | 134 | 21.44 | 22.18 | Confirmed | 91 | 14.93 | 15.08 | Confirmed
[7 empty] | 104.59 | 10 | TRUE | 1,3-Dichlorobenzene | 10.6 | 267684 | 0.76 | 10.459 | Confirmed | 146 | 148 | 63.15 | 62.35 | Confirmed | 111 | 41.81 | 41.32 | Confirmed
[7 empty] | 104.23 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.63 | 473923 | 1.35 | 10.423 | Confirmed | 119 | 91 | 26.67 | 26.64 | Confirmed | 134 | 29.6 | 29.43 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 366888 | 1.04 | 20 | Confirmed | 152 | 115 | 61.11 | 53.05 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 104.69 | 10 | TRUE | 1,4-Dichlorobenzene | 10.67 | 275461 | 0.78 | 10.469 | Confirmed | 146 | 148 | 63.73 | 62.36 | Confirmed | 111 | 44.46 | 42.3 | Confirmed
[7 empty] | 102.63 | 10 | TRUE | n-Butylbenzene | 10.91 | 435578 | 1.24 | 10.263 | Confirmed | 91 | 92 | 52.46 | 51.95 | Confirmed | 134 | 29.07 | 28.55 | Confirmed
[7 empty] | 104.61 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 283595 | 0.81 | 10.461 | Confirmed | 146 | 148 | 62.68 | 62.6 | Confirmed | 111 | 42.47 | 42.39 | Confirmed
[7 empty] | 102.33 | 10 | TRUE | Hexachloroethane | 11.1 | 52124 | 0.15 | 10.233 | Confirmed | 117 | 119 | 98.01 | 95.27 | Confirmed | 201 | 89.25 | 85.99 | Confirmed
[7 empty] | 100.13 | 10 | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.44 | 32319 | 0.09 | 10.013 | Confirmed | 157 | 155 | 76.98 | 77.44 | Confirmed | 75 | 88.27 | 86.21 | Confirmed
[7 empty] | 91.19 | 10 | TRUE | Nitrobenzene | 11.57 | 8018 | 0.02 | 9.119 | Confirmed | 77 | 51 | 44.57 | 44.64 | Confirmed | 123 | 47.29 | 47.86 | Confirmed
[7 empty] | 104.26 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.97 | 182442 | 0.52 | 10.426 | Confirmed | 180 | 182 | 95.79 | 96.21 | Confirmed | 145 | 32.18 | 31.67 | Confirmed
[7 empty] | 100.41 | 10 | TRUE | Hexachloro-1,3-butadiene | 12.06 | 74975 | 0.21 | 10.041 | Confirmed | 225 | 227 | 63.97 | 63.71 | Confirmed | 223 | 63.08 | 62.81 | Confirmed
[7 empty] | 101.5 | 10 | TRUE | Naphthalene | 12.14 | 525258 | 1.49 | 10.15 | Confirmed | 128 | 127 | 12.79 | 12.71 | Confirmed | 129 | 10.56 | 10.99 | Confirmed
[7 empty] | 92.45 | 10 | TRUE | 1,2,3-Trichlorobenzene | 12.28 | 184592 | 0.52 | 9.245 | Confirmed | 180 | 182 | 94.54 | 96.19 | Confirmed | 145 | 33.51 | 33.4 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2023
231031
Peak Name | Amount
ppb
First Injection | CAL1-1 | CAL2-5 | CAL3-10 | CAL4-20 | CAL5-50 | CAL6-100 | H2O-V7 | H2O-V8 | H2O-V9 | 231031-Holmes | 10-23-23-JI-01 | 10-24-KDP-1 | 10-24-KDP-2 | 10-24-KDP-3 | 10-24-KDP-4 | 10-24-KDP-5 | 10-24-KDP-6 | 10-24-KDP-7 | 10-24-KDP-8 | 10-25-KDP-1 | 10-25-KDP-2 | 10-25-KDP-3 | 10-25-KDP-4 | 10-25-KDP-5 | 10-25-KDP-6 | 10-27-23-JI-01 | 10-27-23-JI-02 | 10-28-JES-01 | 10-28-JES-02 | 10-28-JES-03 | 10-28-JES-04 | 10-28-JES-05 | 10-28-JES-06 | 10-28-JES-07 | 10-28-JES-08 | CCV1 | H2O-V10
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0225 | n.a./n.r.
Chloroethene (vinyl chloride) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8713 | n.a./n.r.
Bromomethane (methyl bromide) | 1 | 5 | 10 | 20 | 50 |  | 0.494 | 0.3238 | 0.2183 | 0.1478 | 0.1414 | 0.1579 | 0.1586 | 0.1444 | 0.1629 | n.a./n.r. | 0.1322 | 0.1293 | 0.1435 | 0.1059 | 0.1086 | 0.1288 | 0.1305 | 0.1258 | n.a./n.r. | 0.1236 | 0.1475 | 0.1472 | 0.1549 | 0.1069 | 0.0934 | 0.1251 | 0.1049 | 0.1297 | 0.1171 | 9.7458 | 0.1265
Chloroethane (ethyl chloride) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8161 | n.a./n.r.
Trichlorofluoromethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3305 | n.a./n.r.
Diethyl ether | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8014 | n.a./n.r.
1,1-Dichloroethene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0639 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5495 | n.a./n.r.
Acetone | 1.8 | 9 | 18 | 36 | 90 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.1234 | 5.4863 | 8.2125 | 6.3227 | 9.6142 | 4.6097 | 2.9607 | 7.7057 | 0.879 | 7.5162 | 6.6768 | n.a./n.r. | 1.8573 | 12.099 | 3.6982 | 3.2077 | 4.529 | 3.8156 | 3.2691 | 4.086 | 1.9697 | 4.65 | 3.8897 | 3.0989 | 3.1387 | 15.8082 | n.a./n.r.
Iodomethane | 1 | 5 | 10 | 20 | 50 | 100 | 1.2615 | 1.0058 | 0.5591 | 0.5284 | 0.4067 | 0.3939 | 0.3298 | 0.265 | 0.3684 | 0.3606 | 0.2458 | 0.2295 | 0.2233 | 0.2065 | 0.1933 | 0.2642 | n.a./n.r. | n.a./n.r. | 0.2322 | 0.2647 | 0.2269 | 0.2758 | 0.2586 | n.a./n.r. | 0.1913 | 0.1314 | 0.1925 | 0.1712 | 0.2024 | 8.1427 | 0.3173
Carbon disulfide | 1 | 5 | 10 | 20 | 50 | 100 | 0.3597 | 0.2497 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1832 | 0.1807 | n.a./n.r. | 2.9292 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5784 | n.a./n.r.
3-Chloropropene (allyl chloride) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5775 | n.a./n.r.
Methylene chloride (DCM) | 1 | 5 | 10 | 20 | 50 | 100 | 0.1134 | 0.1061 | 0.0735 | 0.0655 | 0.0454 | 0.0681 | 0.0539 | 0.0609 | 0.0607 | 0.0658 | 0.0532 | 0.0558 | n.a./n.r. | 0.0441 | n.a./n.r. | 0.0558 | 0.0497 | 0.0607 | 0.053 | 0.0521 | 0.0834 | n.a./n.r. | 0.1428 | 0.0399 | 0.0444 | 0.055 | 0.0484 | 0.0518 | 0.0472 | 8.738 | 0.0725
trans-1,2-Dichloroethene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | 0.0642 | 0.0363 | 0.0383 | 0.0326 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.987 | 0.0426
Methyl tert-butyl ether (MTBE) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4894 | n.a./n.r.
1,1-Dichloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2283 | n.a./n.r.
2,2-Dichloropropane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.1046 | n.a./n.r.
cis-1,2-Dichloroethene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0737 | 0.0297 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9754 | 0.0252
2-Butanone (MEK) | 1.8 | 9 | 18 | 36 | 90 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.4553 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.493 | 0.6677 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 3.4471 | n.a./n.r. | 14.883 | n.a./n.r.
Methyl acrylate | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0169 | n.a./n.r.
Bromochloromethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | 0.0151 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0943 | n.a./n.r. | 0.2473 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0223 | 0.017 | 9.0098 | n.a./n.r.
Methacrylonitrile | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4066 | n.a./n.r.
Tetrahydrofuran | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.4881 | 0.708 | 0.2158 | n.a./n.r. | n.a./n.r. | 5.473 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.9182 | n.a./n.r. | 8.4324 | n.a./n.r.
Trichloromethane (chloroform) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 2.7621 | 2.915 | 3.1189 | 2.4956 | 11.002 | n.a./n.r. | 2.6326 | n.a./n.r. | 2.3753 | 2.3328 | 1.0608 | 1.4764 | 3.8385 | 0.9818 | 11.4679 | 12.4584 | 0.0489 | 14.7845 | 4.7989 | 6.3686 | 7.1218 | 5.0787 | 13.3321 | 11.1901 | 9.2941 | n.a./n.r.
1,1,1-Trichloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7855 | n.a./n.r.
Dibromofluoromethane [SS1] | 19.5883 | 20.998 | 19.7289 | 19.8924 | 19.6736 | 20.1187 | 20.6841 | 19.962 | 19.6248 | 19.918 | 19.5359 | 19.7953 | 19.8218 | 20.5569 | 19.5993 | 21.3698 | 19.1364 | 19.5111 | 19.1953 | 19.531 | 19.302 | 19.7911 | 19.8771 | 19.7597 | 19.3632 | 19.4321 | 19.4057 | 19.6704 | 20.0165 | 19.7846 | 19.5405 | 21.9484 | 20.1022 | 18.6709 | 19.0542 | 18.93 | 18.7575
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2039 | n.a./n.r.
Carbon tetrachloride | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0876 | n.a./n.r.
1,1-Dichloropropene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2925 | n.a./n.r.
Benzene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7609 | n.a./n.r.
1,2-Dichloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3443 | n.a./n.r.
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0659 | n.a./n.r. | n.a./n.r. | 0.0221 | 0.0138 | n.a./n.r. | 0.011 | n.a./n.r. | 0.0081 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1708 | 0.0561
1,2-Dichloropropane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7778 | n.a./n.r.
Dibromomethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | 0.03 | n.a./n.r. | 0.0231 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.035 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.4325 | n.a./n.r. | 0.1272 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0647 | 0.0692 | 10.2467 | n.a./n.r.
Methyl methacrylate | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1462 | n.a./n.r.
Bromodichloromethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0629 | 0.0882 | n.a./n.r. | 0.5195 | n.a./n.r. | 0.0365 | n.a./n.r. | 0.0504 | 0.0513 | n.a./n.r. | 3.2668 | 3.4325 | n.a./n.r. | 8.7178 | 7.8563 | n.a./n.r. | 10.4465 | 4.3906 | 4.9493 | 5.7559 | 4.9328 | 11.5968 | 9.4369 | 9.8232 | n.a./n.r.
2-Nitropropane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8589 | n.a./n.r.
cis-1,3-Dichloropropene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5455 | n.a./n.r.
4-Methyl-2-pentanone (MIBK) | 1.8 | 9 | 18 | 36 | 90 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 16.4226 | n.a./n.r.
Toluene-d8 [SS2] | 19.2705 | 19.7427 | 19.7342 | 20.5865 | 20.4388 | 20.2273 | 20.3821 | 21.1209 | 20.0705 | 19.4384 | 20.0486 | 19.4839 | 19.6385 | 19.8677 | 19.6209 | 19.7368 | 19.3196 | 19.3024 | 19.1888 | 19.5419 | 19.6662 | 19.7636 | 19.6451 | 19.4074 | 19.8716 | 19.6532 | 19.1366 | 19.6989 | 19.3528 | 19.629 | 18.9935 | 19.6194 | 19.6156 | 19.3169 | 18.93 | 19.4279 | 18.8881
Toluene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0889 | 0.0616 | 0.0496 | n.a./n.r. | 0.0496 | n.a./n.r. | 0.0472 | 0.0458 | 0.0426 | n.a./n.r. | 0.0526 | 0.0444 | n.a./n.r. | 0.0549 | 0.0779 | 0.0484 | 0.0317 | 0.4717 | n.a./n.r. | 0.0616 | 0.0371 | 0.2018 | 0.0379 | 0.04 | n.a./n.r. | 0.0484 | 0.045 | n.a./n.r. | 0.0362 | 9.9631 | 0.0842
trans-1,3-Dichloropropene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5004 | n.a./n.r.
Ethyl methacrylate | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6734 | n.a./n.r.
1,1,2-Trichloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6751 | n.a./n.r.
Tetrachloroethene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0756 | 0.0564 | 0.0268 | 0.0245 | 0.017 | 0.0155 | 0.0106 | 0.0108 | 0.0103 | 0.0058 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0051 | n.a./n.r. | 0.004 | 0.0048 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0058 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4701 | 0.0767
1,3-Dichloropropane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.774 | n.a./n.r.
2-Hexanone | 1.8 | 9 | 18 | 36 | 90 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 16.4479 | n.a./n.r.
Dibromochloromethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1283 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0505 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.0656 | 1.9265 | n.a./n.r. | 10.7832 | 9.183 | 0.0263 | 6.444 | 3.1781 | 3.4426 | 3.754 | 3.527 | 16.122 | 12.721 | 9.9615 | 0.0393
1,2-Dibromoethane (EDB) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0612 | n.a./n.r.
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7885 | n.a./n.r.
1,1,1,2-Tetrachloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8464 | n.a./n.r.
Ethylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1023 | n.a./n.r. | 0.0665 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0693 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0021 | 0.1436
m/p-Xylene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1239 | 0.0958 | 0.0853 | 0.0245 | 0.0312 | 0.0248 | 0.0235 | 0.0248 | 0.0264 | n.a./n.r. | 0.0261 | 0.0227 | n.a./n.r. | 0.0216 | 0.0231 | n.a./n.r. | n.a./n.r. | 0.1302 | n.a./n.r. | 0.0255 | 0.0178 | 0.0171 | 0.0281 | 0.0269 | 0.0248 | 0.033 | 0.0332 | n.a./n.r. | 0.0188 | 9.9844 | 0.1605
o-Xylene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1065 | 0.0866 | 0.0743 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1182 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0294 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1095 | 0.1242
Styrene | 1 | 5 | 10 | 20 | 50 | 100 | 0.0902 | 0.0729 | 0.0508 | n.a./n.r. | 0.0207 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0531 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0123 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8104 | 0.1239
Bromoform | 1 | 5 | 10 | 20 | 50 | 100 | 0.035 | n.a./n.r. | 0.01 | 1.337 | 0.0159 | 0.004 | 0.0058 | 0.0066 | n.a./n.r. | 0.0079 | n.a./n.r. | 0.0045 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.9329 | 0.2329 | 0.0116 | 13.9934 | 11.6125 | 0.0724 | 1.3171 | 0.4001 | 0.4023 | 0.5446 | 0.434 | 19.8299 | 16.4524 | 10.0292 | 0.0544
Isopropylbenzene (cumene) | 1 | 5 | 10 | 20 | 50 | 100 | 0.094 | 0.0512 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0864 | 0.1242
1-Bromo-4-fluorobenzene (BFB) [SS3] | 19.8758 | 19.9045 | 19.2062 | 20.1723 | 20.3296 | 20.5114 | 22.086 | 22.8214 | 21.947 | 20.9188 | 21.1341 | 20.6136 | 20.6414 | 21.0152 | 20.4774 | 20.8944 | 20.5081 | 19.4793 | 20.4954 | 20.6704 | 19.8757 | 20.1184 | 20.7526 | 20.3921 | 19.9728 | 20.2293 | 20.6864 | 20.2007 | 19.9316 | 19.9414 | 20.1394 | 20.6006 | 19.5963 | 19.6219 | 19.0277 | 19.5942 | 19.6818
trans-1,4-Dichloro-2-butene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.3716 | 0.0417
Bromobenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1208 | 0.055 | 0.0399 | 0.0288 | 0.0226 | 0.0185 | n.a./n.r. | 0.0169 | 0.0157 | n.a./n.r. | 0.0115 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.01 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1524 | 0.1036
1,1,2,2-Tetrachloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6704 | n.a./n.r.
1,2,3-Trichloropropane (TCP) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8393 | 0.0375
n-Propylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1453 | 0.0876 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1901 | 0.1729
2-Chlorotoluene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1295 | 0.0698 | 0.0318 | n.a./n.r. | 0.0204 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0134 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1587 | 0.1402
1,3,5-Trimethylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | 0.0708 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.445 | 0.1331
4-Chlorotoluene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1627 | n.a./n.r. | 0.0502 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0181 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.3423 | 0.1566
tert-Butylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1539 | 0.061 | 0.031 | 0.0215 | 0.0147 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9295 | 0.1057
Pentachloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3652 | n.a./n.r.
1,2,4-Trimethylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1464 | 0.0909 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0371 | n.a./n.r. | 0.0358 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1369 | n.a./n.r. | 0.0224 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0239 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0236 | 10.3603 | 0.1475
1-Methylpropylbenzene (sec-butylbenzene) | 1 | 5 | 10 | 20 | 50 | 100 | 0.1515 | 0.087 | n.a./n.r. | n.a./n.r. | 0.0249 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1195 | 0.166
1,3-Dichlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.2359 | 0.1233 | 0.0602 | 0.0504 | 0.0385 | 0.0283 | 0.026 | 0.027 | 0.023 | 0.0232 | 0.0151 | n.a./n.r. | 0.014 | 0.012 | n.a./n.r. | 0.0125 | n.a./n.r. | n.a./n.r. | 0.0096 | 0.0094 | n.a./n.r. | 0.0077 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4587 | 0.165
4-Isopropyltoluene (p-cymene) | 1 | 5 | 10 | 20 | 50 | 100 | 0.154 | 0.0996 | 0.0483 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4234 | 0.1772
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4689 | n.a./n.r.
n-Butylbenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.245 | 0.1478 | 0.087 | 0.0617 | 0.0552 | 0.0375 | 0.0396 | 0.0338 | 0.0296 | 0.0237 | 0.0249 | 0.0171 | n.a./n.r. | 0.0188 | 0.0153 | 0.0136 | 0.0152 | n.a./n.r. | 0.0126 | 0.0151 | 0.0114 | 0.0145 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.2627 | 0.2004
1,2-Dichlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.1721 | 0.0922 | 0.0485 | 0.04 | 0.0327 | 0.0219 | 0.0244 | 0.0251 | 0.0213 | 0.0187 | 0.0167 | 0.0159 | 0.0146 | 0.0167 | n.a./n.r. | 0.0131 | 0.0108 | 0.0127 | 0.0113 | 0.0105 | n.a./n.r. | 0.0097 | 0.0101 | 0.0116 | 0.0124 | n.a./n.r. | n.a./n.r. | 0.0094 | 0.0105 | 10.4614 | 0.1485
Hexachloroethane | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.2328 | n.a./n.r.
1,2-Dibromo-3-chloropropane (DBCP) | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0129 | n.a./n.r.
Nitrobenzene | 1 | 5 | 10 | 20 | 50 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1192 | n.a./n.r.
1,2,4-Trichlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.4487 | 0.2122 | 0.1345 | 0.0803 | 0.0727 | 0.0481 | 0.0462 | 0.0501 | 0.0365 | 0.0319 | 0.0291 | 0.0294 | 0.0209 | 0.0247 | 0.0236 | 0.0186 | 0.018 | n.a./n.r. | 0.0159 | 0.0181 | 0.0134 | n.a./n.r. | 0.014 | 0.0148 | n.a./n.r. | 0.0119 | n.a./n.r. | 0.0094 | n.a./n.r. | 10.4262 | 0.233
Hexachloro-1,3-butadiene | 1 | 5 | 10 | 20 | 50 | 100 | 0.3036 | 0.1632 | 0.1049 | 0.0671 | 0.0594 | 0.0506 | 0.0405 | 0.0439 | 0.0334 | 0.0276 | 0.025 | 0.022 | 0.0206 | 0.0161 | 0.0214 | n.a./n.r. | n.a./n.r. | 0.0205 | 0.0165 | n.a./n.r. | 0.0088 | 0.0138 | n.a./n.r. | 0.0129 | 0.0106 | n.a./n.r. | 0.013 | n.a./n.r. | 0.0068 | 10.0413 | 0.1854
Naphthalene | 1 | 5 | 10 | 20 | 50 | 100 | 0.4567 | 0.2011 | 0.094 | 0.0641 | 0.0525 | 0.0441 | 0.0413 | 0.0446 | 0.0343 | 0.027 | 0.0221 | n.a./n.r. | 0.0162 | 0.0272 | 0.0233 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0121 | 0.012 | n.a./n.r. | 0.1491 | n.a./n.r. | n.a./n.r. | 0.0073 | n.a./n.r. | 0.0112 | n.a./n.r. | n.a./n.r. | 10.1502 | 0.1958
1,2,3-Trichlorobenzene | 1 | 5 | 10 | 20 | 50 | 100 | 0.3267 | 0.1353 | 0.0971 | 0.0625 | 0.049 | 0.0386 | 0.0406 | 0.0291 | 0.0285 | 0.0266 | 0.0264 | 0.0197 | 0.0162 | n.a./n.r. | n.a./n.r. | 0.0135 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0103 | 0.0102 | n.a./n.r. | n.a./n.r. | 0.0106 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2453 | 0.2294

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library | Area | Estimated | 20 ppb (TIC) | RT | Area
| No. | min | SI |  |  | SI |  |  | SI |  |  | counts*min | Concentration | Pentafluorobenzene [IS1] | 5.42 | 1092292
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC |  | (ppb) | 1,4-Difluorobenzene [IS2] | 6.17 | 1411137
10-28-JES-01 | 1 | 1.53 | 733 | 2-Butene | mainlib | 712 | 1-Propene, 2-methyl- | mainlib | 699 | 1-Butene | mainlib | 12818.192 | 0.2347 | Chlorobenzene-d5 [IS3] | 8.91 | 1744613
10-28-JES-01 | 5 | 3.26 | 898 | Cyclopentene | mainlib | 893 | 2,3-Diazabicyclo[2.2.1]-hept-2-ene | mainlib | 889 | Bicyclo[2.1.0]pentane | mainlib | 141351.219 | 2.5882 | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 1927012
10-28-JES-01 | 6 | 3.4 | 951 | Cyclopentane | mainlib | 924 | 1-Pentene | mainlib | 880 | Cyclopropane, ethyl- | mainlib | 10799930.2 | 197.748
10-28-JES-01 | 7 | 3.55 | 759 | Acetic acid, cyano-, 1,1-dimethylethyl ester | mainlib | 756 | 2-Propanol, 2-methyl- | mainlib | 730 | N,N-Dimethylformamide-di-t-butylacetal | mainlib | 54377.907 | 0.99567
10-28-JES-01 | 12 | 7.39 | 701 | 3-Hexanol, 5-methyl- | mainlib | 672 | 2-Butanol, 2,3-dimethyl- | mainlib | 656 | 3-Hexanol, 4-methyl- | mainlib | 15590.636 | 0.22097
10-28-JES-01 | 17 | 8.48 | 886 | Cyclopentanone | mainlib | 803 | 2-Ethylacrolein | mainlib | 782 | 2-Butenal, 2-methyl-, (E)- | mainlib | 130195.746 | 1.4925