## IS-RSD
[8 empty] | Peak Name | Peak Area
[9 empty] | counts*min | 1 | 2 | 3 | 4 | 5 | 6
Total Ion Count | [7 empty] | First Injection | Startup | H2O-R1 | H2O-R2 | H2O-R3 | H2O-R4
RSD | RSD (1-3) | RSD (2-4) | RSD (3-5) | RSD (4-6) | RSD (5-7) | RSD (6-8) | RSD (7-9) | TIC | TIC | TIC | TIC | TIC | TIC
3.3737 | 4.7668 | 3.3737 | 2.0694 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Pentafluorobenzene [IS1] |  |  | 1018568.737 | 1089628.013 | 1058199.582
2.2745 | 2.6669 | 2.2745 | 0.25596 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1,4-Difluorobenzene [IS2] |  |  | 1462448.572 | 1518665.636 | 1524172.829
4.0995 | 4.3966 | 4.0995 | 1.151 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Chlorobenzene-d5 [IS3] |  |  | 1731301.178 | 1842402.668 | 1872638.795
4.7418 | 4.088 | 4.7418 | 2.598 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1,4-Dichlorobenzene-d4 [IS4] |  |  | 1752017.886 | 1856323.635 | 1925804.652

## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2024/240201.seq/308.smp/MSDevice.xraw
Channel | TIC
Injection Information: |  |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Data Vault | ChromeleonLocal |  |  | 95 | 5699941 | 50-200% of mass 174 | 95/174 | 196.1938 | 50 | 200 | TRUE
Injection | H2O-R7 |  |  | 96 | 374163.625 | 5 to 9% of m/z 95 | 96/95 | 6.5643 | 5 | 9 | TRUE
Injection Number | 12 |  |  | 173 | #N/A | <2% of m/z 174 | 173/174 | #N/A | 0 | 2 | #N/A
Position | 1 |  |  | 174 | 2905260.25 | 50-200% of mass 95 | 174/95 | 50.97 | 50 | 200 | TRUE
Comment |  |  |  | 175 | 183819.5312 | 5 to 9% of m/z 174 | 175/174 | 6.3271 | 5 | 9 | TRUE
Processing Method | VOC-11 |  |  | 176 | 2801184 | 95 to 105% of m/z 174 | 176/174 | 96.4177 | 95 | 105 | TRUE
Instrument Method | VOC-27 |  |  | 177 | 175037.25 | 5 to 10% of m/z 176 | 177/176 | 6.2487 | 5 | 10 | TRUE
Type | Unknown
Status | Running
Injection Date | 2024-02-01
Injection Time | 14:38
Injection Volume (µl) | 20 | *N/A for P&T
Dilution Factor | 1899-12-31
Weight | 24:00
Raw Data Information:
Data Points | 230
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9808
Retention Time (min) | 9.8412
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3288
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0211
Base Peak Intensity | 5699941
Total Ion Count | 20221459.6024
Raw Data: |  | Rounded:
Mass (m/z) | Value (counts) | Mass (m/z) | Value (counts)
36 | 26643.4004 | 36 | 26643.4004
37 | 146982.3281 | 37 | 146982.3281
38 | 138128.8281 | 38 | 138128.8281
39 | 62325.0273 | 39 | 62325.0273
40 | 34448.4141 | 40 | 34448.4141
41.1 | 1608.142 | 41 | 1608.142
42.1 | 1561.6711 | 42 | 1561.6711
42.8 | 2088.5867 | 43 | 2088.5867
44 | 102932.8281 | 44 | 102932.8281
45 | 30693.8145 | 45 | 30693.8145
46 | 1477.0785 | 46 | 1477.0785
47 | 78740.3672 | 47 | 78740.3672
47.9 | 18700.0977 | 48 | 18700.0977
49 | 145945.125 | 49 | 145945.125
50 | 718372.5625 | 50 | 718372.5625
51 | 212203.1094 | 51 | 212203.1094
52 | 8447.3984 | 52 | 8447.3984
53.1 | 280.7314 | 53 | 280.7314
54 | 225.3855 | 54 | 225.3855
55 | 5648.2769 | 55 | 5648.2769
56 | 32964.6797 | 56 | 32964.6797
57 | 86648.125 | 57 | 86648.125
58 | 3230.9766 | 58 | 3230.9766
59.1 | 247.2806 | 59 | 247.2806
60 | 29679.3672 | 60 | 29679.3672
61 | 153421.2656 | 61 | 153421.2656
62 | 150792.1719 | 62 | 150792.1719
63 | 110773.9219 | 63 | 110773.9219
63.9 | 11349.7598 | 64 | 11349.7598
65 | 1214.6661 | 65 | 1214.6661
66 | 43.0963 | 66 | 43.0963
67 | 11682.1055 | 67 | 11682.1055
68 | 416728.4375 | 68 | 416728.4375
69 | 384432.875 | 69 | 384432.875
70.1 | 35735.125 | 70 | 35735.125
70.9 | 1128.6119 | 71 | 1128.6119
71.9 | 17323.502 | 72 | 17323.502
73 | 162133.4531 | 73 | 162133.4531
74 | 628213.4375 | 74 | 628213.4375
75 | 2366125 | 75 | 2366125
76.1 | 193777.7188 | 76 | 193777.7188
77 | 33717.8789 | 77 | 33717.8789
77.9 | 23484.0781 | 78 | 23484.0781
78.9 | 31748.3594 | 79 | 31748.3594
80 | 11859.5215 | 80 | 11859.5215
80.9 | 32982.8203 | 81 | 32982.8203
81.9 | 5578.0015 | 82 | 5578.0015
83.1 | 1321.3463 | 83 | 1321.3463
83.6 | 314.9723 | 84 | 314.9723
84.3 | 348.914 | 84 | 348.914
85.7 | 4440.6479 | 86 | 4440.6479
87 | 234809.3906 | 87 | 234809.3906
88 | 231341 | 88 | 231341
88.7 | 5488.5898 | 89 | 5488.5898
89.4 | 433.3797 | 89 | 433.3797
91 | 7671.9595 | 91 | 7671.9595
92 | 93964.5391 | 92 | 93964.5391
93 | 156629.7812 | 93 | 156629.7812
94 | 447261.1562 | 94 | 447261.1562
95 | 5699941 | 95 | 5699941
96 | 374163.625 | 96 | 374163.625
97.1 | 9986.3682 | 97 | 9986.3682
98.2 | 693.7884 | 98 | 693.7884
99 | 186.675 | 99 | 186.675
99.7 | 0.13028 | 100 | 0.13028
100.2 | 74.9177 | 100 | 74.9177
100.7 | 5e-06 | 101 | 5e-06
102.1 | 0.019375 | 102 | 0.019375
102.8 | 1100.371 | 103 | 1100.371
103.9 | 8824.9824 | 104 | 8824.9824
104.8 | 4404.8921 | 105 | 4404.8921
105.9 | 8337.2402 | 106 | 8337.2402
107 | 1950.3115 | 107 | 1950.3115
107.6 | 0.60541 | 108 | 0.60541
108.2 | 127.9693 | 108 | 127.9693
109.8 | 173.6938 | 110 | 173.6938
111 | 2112.4749 | 111 | 2112.4749
111.7 | 276.7997 | 112 | 276.7997
113.1 | 444.1409 | 113 | 444.1409
114 | 159.1235 | 114 | 159.1235
115 | 2553.7771 | 115 | 2553.7771
115.9 | 8414.0449 | 116 | 8414.0449
116.9 | 14734.1436 | 117 | 14734.1436
117.9 | 7656.0542 | 118 | 7656.0542
118.9 | 9960.9609 | 119 | 9960.9609
119.9 | 178.3394 | 120 | 178.3394
121.2 | 60.1379 | 121 | 60.1379
122 | 626.2733 | 122 | 626.2733
122.9 | 1086.9561 | 123 | 1086.9561
123.9 | 2060.1572 | 124 | 2060.1572
124.9 | 1960.9475 | 125 | 1960.9475
125.9 | 904.0575 | 126 | 904.0575
126.7 | 954.6217 | 127 | 954.6217
127.9 | 7753.4971 | 128 | 7753.4971
128.9 | 3426.4409 | 129 | 3426.4409
129.9 | 12365.4883 | 130 | 12365.4883
130.9 | 3912.072 | 131 | 3912.072
131.9 | 1846.5525 | 132 | 1846.5525
133 | 398.88 | 133 | 398.88
134.1 | 1080.8745 | 134 | 1080.8745
134.9 | 3222.3518 | 135 | 3222.3518
135.9 | 846.4373 | 136 | 846.4373
136.9 | 2686.2075 | 137 | 2686.2075
137.6 | 124.6506 | 138 | 124.6506
138.8 | 587.01 | 139 | 587.01
140 | 948.1526 | 140 | 948.1526
140.9 | 16146.7295 | 141 | 16146.7295
141.9 | 1226.7781 | 142 | 1226.7781
142.9 | 20908.0137 | 143 | 20908.0137
143.7 | 848.9271 | 144 | 848.9271
144.2 | 428.0378 | 144 | 428.0378
144.8 | 1524.6968 | 145 | 1524.6968
145.9 | 4066.3616 | 146 | 4066.3616
146.9 | 1490.8384 | 147 | 1490.8384
147.9 | 6012.0435 | 148 | 6012.0435
148.9 | 2558.5364 | 149 | 2558.5364
150 | 3053.7896 | 150 | 3053.7896
151.1 | 451.9929 | 151 | 451.9929
152 | 844.1075 | 152 | 844.1075
153 | 3697.6208 | 153 | 3697.6208
154 | 2221.1829 | 154 | 2221.1829
154.9 | 6999.4619 | 155 | 6999.4619
156.1 | 1291.6914 | 156 | 1291.6914
157 | 5927.3799 | 157 | 5927.3799
157.8 | 879.1988 | 158 | 879.1988
158.9 | 2362.3013 | 159 | 2362.3013
159.4 | 109.9633 | 159 | 109.9633
160.9 | 2490.384 | 161 | 2490.384
162.3 | 75.8921 | 162 | 75.8921
163.2 | 0.000293 | 163 | 0.000293
164.3 | 0 | 164 | 0
165.5 | 0 | 166 | 0
166.1 | 0 | 166 | 0
166.6 | 0 | 167 | 0
167.5 | 1.8e-05 | 168 | 1.8e-05
168.1 | 0 | 168 | 0
168.9 | 0 | 169 | 0
169.7 | 65.2388 | 170 | 65.2388
170.5 | 153.7077 | 171 | 153.7077
171.3 | 40.6828 | 171 | 40.6828
172.1 | 1323.5099 | 172 | 1323.5099
173.9 | 2905260.25 | 174 | 2905260.25
175 | 183819.5312 | 175 | 183819.5312
175.9 | 2801184 | 176 | 2801184
177 | 175037.25 | 177 | 175037.25
178 | 5895.1919 | 178 | 5895.1919
179.5 | 0 | 180 | 0
180 | 0.11994 | 180 | 0.11994
180.9 | 79.1076 | 181 | 79.1076
181.4 | 0.00393 | 181 | 0.00393
182 | 0 | 182 | 0
184.4 | 0 | 184 | 0
185.8 | 0.004604 | 186 | 0.004604
186.8 | 43.2823 | 187 | 43.2823
187.9 | 0.11197 | 188 | 0.11197
188.8 | 4.0734 | 189 | 4.0734
189.9 | 0 | 190 | 0
190.8 | 82.7365 | 191 | 82.7365
192.3 | 101.971 | 192 | 101.971
193 | 2e-05 | 193 | 2e-05
193.6 | 0 | 194 | 0
194.2 | 33.843 | 194 | 33.843
194.7 | 0 | 195 | 0
195.8 | 0 | 196 | 0
196.4 | 0.000186 | 196 | 0.000186
197.2 | 0.000108 | 197 | 0.000108
198.1 | 33.4836 | 198 | 33.4836
199 | 0 | 199 | 0
200.3 | 0 | 200 | 0
201.2 | 0 | 201 | 0
202.4 | 0.000167 | 202 | 0.000167
203 | 1e-06 | 203 | 1e-06
205.4 | 0 | 205 | 0
206 | 0 | 206 | 0
207.2 | 35.5442 | 207 | 35.5442
207.9 | 4.2e-05 | 208 | 4.2e-05
209.1 | 0 | 209 | 0
210.6 | 0 | 211 | 0
211.1 | 0 | 211 | 0
212.3 | 0 | 212 | 0
212.9 | 0 | 213 | 0
214 | 0.011025 | 214 | 0.011025
215.1 | 0 | 215 | 0
216.3 | 42.0791 | 216 | 42.0791
217.6 | 120.5884 | 218 | 120.5884
218.5 | 1.4756 | 219 | 1.4756
219.6 | 0 | 220 | 0
220.2 | 0 | 220 | 0
221.2 | 0 | 221 | 0
221.8 | 0 | 222 | 0
224.3 | 38.2146 | 224 | 38.2146
225.2 | 0 | 225 | 0
226.2 | 0 | 226 | 0
227.1 | 2.3793 | 227 | 2.3793
227.6 | 0 | 228 | 0
228.4 | 0 | 228 | 0
229.1 | 3e-05 | 229 | 3e-05
230.3 | 0 | 230 | 0
230.9 | 0 | 231 | 0
231.9 | 75.58 | 232 | 75.58
232.4 | 59.6409 | 232 | 59.6409
232.9 | 0 | 233 | 0
233.4 | 0.003275 | 233 | 0.003275
234.3 | 0 | 234 | 0
235 | 87.469 | 235 | 87.469
235.6 | 0 | 236 | 0
236.4 | 0 | 236 | 0
237.5 | 0 | 238 | 0
238.7 | 0.085348 | 239 | 0.085348
239.2 | 0 | 239 | 0
240.8 | 108.2421 | 241 | 108.2421
242.1 | 0 | 242 | 0
243.1 | 1.5e-05 | 243 | 1.5e-05
244.8 | 0 | 245 | 0
245.3 | 0.005579 | 245 | 0.005579
246.4 | 241.1767 | 246 | 241.1767
247 | 0 | 247 | 0
247.9 | 0 | 248 | 0
248.4 | 9e-06 | 248 | 9e-06
248.9 | 0.15144 | 249 | 0.15144
250.2 | 0.081248 | 250 | 0.081248
251.7 | 0 | 252 | 0
253.2 | 0 | 253 | 0
253.8 | 0 | 254 | 0
255.6 | 0 | 256 | 0
256.5 | 69.3905 | 257 | 69.3905
257.1 | 1.2e-05 | 257 | 1.2e-05
257.9 | 0 | 258 | 0
258.8 | 0.15275 | 259 | 0.15275
259.8 | 0 | 260 | 0
252.7 | 0.1609 | 253 | 0.1609
253.7 | 0.29099 | 254 | 0.29099
254.3 | 36.9879 | 254 | 36.9879
255.5 | 0.11669 | 256 | 0.11669
256.5 | 0 | 257 | 0
257.7 | 0.010091 | 258 | 0.010091
258.9 | 7e-06 | 259 | 7e-06
259.4 | 0.10145 | 259 | 0.10145
258 | 332.3039 | 258 | 332.3039
259.4 | 74.2735
260 | 0
259.9 | 2.1e-05
259.7 | 0
259.3 | 0
259.9 | 0
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[12 empty] | TV | [6 empty] | Measured | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
Name | RT | Window | Quant | Conf1 | Conf2 | Window |  | Channel | Eval. Type | IS | Cal.Type | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | RSD<20% | R^2>0.99 | Total Analytes | 81
Chloromethane (methyl chloride) | 1.47 | 0.050 AN | 50 | 52 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.09 | 1.777 | 4.8927 | 10.0865 | 25.0063 | 99.9982 | 109 | 88.85 | 97.854 | 100.865 | 100.0252 | 99.9982 | Chloromethane (methyl chloride) | 1.48 | Quad | 6 | 0.7167 | 0.99998 | TRUE | TRUE | RSD Failed | 0
Chloroethene (vinyl chloride) | 1.56 | 0.050 AN | 62 | 64 | 61 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0972 | 1.9047 | 5.3569 | 9.6189 | 25.1031 | 99.9946 | 109.72 | 95.235 | 107.138 | 96.189 | 100.4124 | 99.9946 | Chloroethene (vinyl chloride) | 1.57 | Quad | 6 | 1.3444 | 0.99993 | TRUE | TRUE | R^2 Failed | 0
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.4778 | 2.5011 | 5.8977 | 9.8118 | 24.7156 | 99.789 | 147.78 | 125.055 | 117.954 | 98.118 | 98.8624 | 99.789 | Bromomethane (methyl bromide) | 1.85 | Quad | 6 | 3.8351 | 0.99918 | TRUE | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.25 | 2.2373 | 6.1491 | 10.8411 | 29.3479 | 98.7642 | 125 | 111.865 | 122.982 | 108.411 | 117.3916 | 98.7642 | Chloroethane (ethyl chloride) | 1.96 | Lin | 6 | 8.5806 | 0.99682 | TRUE | TRUE
Trichlorofluoromethane | 2.2 | 0.050 AN | 101 | 103 | 105 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1952 | 2.1693 | 6.311 | 11.1396 | 29.5625 | 98.6745 | 119.52 | 108.465 | 126.22 | 111.396 | 118.25 | 98.6745 | Trichlorofluoromethane | 2.21 | Lin | 6 | 9.1187 | 0.99638 | TRUE | TRUE
Diethyl ether | 2.52 | 0.080 AN | 59 | 74 | 45 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9574 | 1.8905 | 5.1593 | 10.6321 | 28.6884 | 99.0093 | 95.74 | 94.525 | 103.186 | 106.321 | 114.7536 | 99.0093 | Diethyl ether | 2.52 | Lin | 6 | 7.1073 | 0.99791 | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2105 | 2.1569 | 5.9988 | 10.7714 | 28.8428 | 98.907 | 121.05 | 107.845 | 119.976 | 107.714 | 115.3712 | 98.907 | 1,1-Dichloroethene | 2.75 | Lin | 6 | 7.6169 | 0.99753 | TRUE | TRUE
Acetone | 2.87 | 0.088 AN | 43 | 58 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 2.7536 | 3.5933 | 9.0214 | 20.7331 | 49.8983 | 200.0018 | 137.68 | 89.8325 | 90.214 | 103.6655 | 99.7966 | 100.0009 | Acetone | 2.86 | Quad, WithOffset | 6 | 2.0452 | 0.9999 | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, 1/A | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1 | 2 | 5 | 10 | 25 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Iodomethane | 2.91 | P-P | 6 | 0 | 1 | TRUE | TRUE
Carbon disulfide | 2.97 | 0.050 AN | 76 | 78 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1983 | 2.1886 | 5.7351 | 10.77 | 28.5334 | 98.9971 | 119.83 | 109.43 | 114.702 | 107.7 | 114.1336 | 98.9971 | Carbon disulfide | 2.97 | Lin | 6 | 6.9782 | 0.99794 | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9277 | 1.7027 | 4.7992 | 9.1602 | 25.472 | 99.982 | 92.77 | 85.135 | 95.984 | 91.602 | 101.888 | 99.982 | 3-Chloropropene (allyl chloride) | 3.22 | Quad | 6 | 2.0638 | 0.99988 | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1393 | 1.8714 | 4.9915 | 9.29 | 25.38 | 99.9723 | 113.93 | 93.57 | 99.83 | 92.9 | 101.52 | 99.9723 | Methylene chloride (DCM) | 3.38 | Quad | 6 | 2.0085 | 0.99985 | TRUE | TRUE
trans-1,2-Dichloroethene | 3.68 | 0.100 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.163 | 2.1406 | 5.8077 | 10.3944 | 28.3523 | 99.0776 | 116.3 | 107.03 | 116.154 | 103.944 | 113.4092 | 99.0776 | trans-1,2-Dichloroethene | 3.7 | Lin | 6 | 6.5699 | 0.99819 | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9304 | 1.6897 | 4.9124 | 10.1352 | 26.9218 | 99.5173 | 93.04 | 84.485 | 98.248 | 101.352 | 107.6872 | 99.5173 | Methyl tert-butyl ether (MTBE) | 3.71 | Lin | 6 | 3.7488 | 0.99944 | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0846 | 1.9259 | 5.2279 | 10.0366 | 26.8947 | 99.5119 | 108.46 | 96.295 | 104.558 | 100.366 | 107.5788 | 99.5119 | 1,1-Dichloroethane | 4.2 | Lin | 6 | 3.6598 | 0.99946 | TRUE | TRUE
cis-1,2-Dichloroethene | 4.822 | 0.060 AN | 77 | 41 | 79 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0979 | 2.0123 | 5.466 | 10.215 | 27.382 | 99.3585 | 109.79 | 100.615 | 109.32 | 102.15 | 109.528 | 99.3585 | 2,2-Dichloropropane | 4.83 | Quad, WithOffset | 6 | 3.1822 | 0.99978 | TRUE | TRUE
2,2-Dichloropropane | 4.83 | 0.060 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.4758 | 2.1644 | 4.9988 | 8.8976 | 25.4808 | 99.9819 | 147.58 | 108.22 | 99.976 | 88.976 | 101.9232 | 99.9819 | cis-1,2-Dichloroethene | 4.84 | Lin | 6 | 4.6493 | 0.99912 | TRUE | TRUE
2-Butanone (MEK) | 4.86 | 0.060 AN | 43 | 72 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 1.078 | 2.6145 | 8.7686 | 19.1099 | 50.2353 | 200.1287 | 53.9 | 65.3625 | 87.686 | 95.5495 | 100.4706 | 100.0644 | 2-Butanone (MEK) | 4.86 | Lin | 6 | 2.1598 | 0.99983 | TRUE | TRUE
Methyl acrylate | 4.953 | 0.060 AN | 55 | 85 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.7347 | 2.2663 | 4.4548 | 9.0756 | 25.4804 | 99.9846 | 173.47 | 113.315 | 89.096 | 90.756 | 101.9216 | 99.9846 | Methyl acrylate | 4.95 | Quad, WithOffset | 6 | 3.1736 | 0.99981 | TRUE | TRUE
Bromochloromethane | 5.073 | 0.060 AN | 49 | 130 | 128 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0441 | 1.8956 | 5.2261 | 10.2989 | 27.6517 | 99.2975 | 104.41 | 94.78 | 104.522 | 102.989 | 110.6068 | 99.2975 | Methacrylonitrile | 5.08 | Lin | 6 | 3.4569 | 0.99954 | TRUE | TRUE
Methacrylonitrile | 5.076 | 0.060 AN | 67 | 52 | 40 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | 1.5089 | 4.6944 | 9.7829 | 26.6326 | 99.6432 | #VALUE! | 75.445 | 93.888 | 97.829 | 106.5304 | 99.6432 | Bromochloromethane | 5.09 | Lin | 6 | 5.1131 | 0.99894 | TRUE | TRUE
Tetrahydrofuran | 5.1 | 0.060 AN | 42 | 72 | 71 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.6933 | 1.4517 | 4.3273 | 10.2069 | 26.7251 | 99.5957 | 69.33 | 72.585 | 86.546 | 102.069 | 106.9004 | 99.5957 | Tetrahydrofuran | 5.1 | Lin, WithOffset | 6 | 4.2387 | 0.99945 | TRUE | TRUE
Trichloromethane (chloroform) | 5.211 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1006 | 2.0489 | 5.6112 | 10.6313 | 28.7874 | 98.9575 | 110.06 | 102.445 | 112.224 | 106.313 | 115.1496 | 98.9575 | Trichloromethane (chloroform) | 5.22 | Lin | 6 | 7.3438 | 0.99774 | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7901 | 1.5102 | 4.9441 | 8.8966 | 25.2422 | 100.0645 | 79.01 | 75.51 | 98.882 | 88.966 | 100.9688 | 100.0645 | 1,1,1-Trichloroethane | 5.35 | Lin | 6 | 2.3749 | 0.99979 | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 19.3319 | 19.5744 | 19.1811 | 20.6986 | 20.8382 | 20.3759 | 96.6595 | 97.872 | 95.9055 | 103.493 | 104.191 | 101.8795 | Dibromofluoromethane [SS1] | 5.37 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Pentafluorobenzene [IS1] | 5.44 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.486 | 0.040 AN | 56 | 41 | 43 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8339 | 1.5062 | 4.7513 | 8.3487 | 24.3231 | 100.3583 | 83.39 | 75.31 | 95.026 | 83.487 | 97.2924 | 100.3583 | 1-Chlorobutane (butyl chloride) | 5.49 | Lin | 6 | 3.6568 | 0.99951 | TRUE | TRUE
Carbon tetrachloride | 5.501 | 0.040 AN | 119 | 121 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.3176 | 2.0925 | 5.1081 | 9.1523 | 25.3395 | 99.9888 | 131.76 | 104.625 | 102.162 | 91.523 | 101.358 | 99.9888 | Carbon tetrachloride | 5.5 | Quad, WithOffset | 6 | 2.2555 | 0.9999 | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9319 | 1.7649 | 5.3331 | 9.4147 | 26.8008 | 99.5971 | 93.19 | 88.245 | 106.662 | 94.147 | 107.2032 | 99.5971 | 1,1-Dichloropropene | 5.52 | Lin | 6 | 3.6927 | 0.99946 | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9828 | 1.9138 | 5.5042 | 9.938 | 27.7419 | 99.2974 | 98.28 | 95.69 | 110.084 | 99.38 | 110.9676 | 99.2974 | Benzene | 5.71 | Lin | 6 | 5.3102 | 0.99886 | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9435 | 1.6737 | 5.0788 | 9.3526 | 25.3527 | 99.9766 | 94.35 | 83.685 | 101.576 | 93.526 | 101.4108 | 99.9766 | 1,2-Dichloroethane | 5.78 | Quad | 6 | 1.9386 | 0.99987 | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.181 | 0.080 AN | 114 | 88 | 63 | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Difluorobenzene [IS2] | 6.18 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9968 | 1.9406 | 5.5997 | 9.6927 | 26.273 | 99.6837 | 99.68 | 97.03 | 111.994 | 96.927 | 105.092 | 99.6837 | Trichloroethene | 6.4 | Lin | 6 | 2.7465 | 0.9997 | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8148 | 1.553 | 4.718 | 8.8684 | 25.3583 | 100.0485 | 81.48 | 77.65 | 94.36 | 88.684 | 101.4332 | 100.0485 | 1,2-Dichloropropane | 6.65 | Lin | 6 | 2.493 | 0.99977 | TRUE | TRUE
Dibromomethane | 6.729 | 0.060 AN | 174 | 93 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8569 | 1.5989 | 5.0465 | 8.9605 | 24.7221 | 100.1806 | 85.69 | 79.945 | 100.93 | 89.605 | 98.8884 | 100.1806 | Dibromomethane | 6.73 | Lin | 6 | 2.225 | 0.99981 | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.6063 | 2.2205 | 4.875 | 8.8398 | 25.4668 | 99.9883 | 160.63 | 111.025 | 97.5 | 88.398 | 101.8672 | 99.9883 | Methyl methacrylate | 6.75 | Quad, WithOffset | 6 | 2.8084 | 0.99986 | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7574 | 1.4122 | 4.5838 | 8.3891 | 24.293 | 100.3728 | 75.74 | 70.61 | 91.676 | 83.891 | 97.172 | 100.3728 | Bromodichloromethane | 6.93 | Lin | 6 | 3.747 | 0.99949 | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | n.a./n.r. | 4.6447 | 10.0761 | 25.1542 | 99.9931 | #VALUE! | #VALUE! | 92.894 | 100.761 | 100.6168 | 99.9931 | 2-Nitropropane | 7.16 | Quad | 6 | 2.0183 | 0.99989 | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7712 | 1.4001 | 4.6019 | 8.769 | 25.673 | 99.9803 | 77.12 | 70.005 | 92.038 | 87.69 | 102.692 | 99.9803 | cis-1,3-Dichloropropene | 7.36 | Quad | 6 | 2.7907 | 0.9998 | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.532 | 0.049 AN | 43 | 58 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 1.3748 | 2.6738 | 8.5469 | 17.4684 | 51.5785 | 199.9494 | 68.74 | 66.845 | 85.469 | 87.342 | 103.157 | 99.9747 | 4-Methyl-2-pentanone (MIBK) | 7.53 | Quad | 6 | 3.3026 | 0.99972 | TRUE | TRUE
Toluene-d8 [SS2] | 7.62 | 0.058 AN | 98 | 100 | 70 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 19.2488 | 19.2131 | 20.0488 | 20.9532 | 20.5483 | 19.9879 | 96.244 | 96.0655 | 100.244 | 104.766 | 102.7415 | 99.9395 | Toluene-d8 [SS2] | 7.61 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Toluene | 7.684 | 0.047 AN | 91 | 92 | 65 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.953 | 1.7491 | 5.2413 | 9.2707 | 26.9511 | 99.5786 | 95.3 | 87.455 | 104.826 | 92.707 | 107.8044 | 99.5786 | Toluene | 7.68 | Lin | 6 | 4.0208 | 0.99937 | TRUE | TRUE
trans-1,3-Dichloropropene | 7.931 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.5962 | 2.2797 | 4.9733 | 8.617 | 25.542 | 99.9859 | 159.62 | 113.985 | 99.466 | 86.17 | 102.168 | 99.9859 | trans-1,3-Dichloropropene | 7.93 | Quad, WithOffset | 6 | 3.3129 | 0.9998 | TRUE | TRUE
Ethyl methacrylate | 8 | 0.060 AN | 69 | 41 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.8664 | 2.3279 | 4.6568 | 8.5987 | 25.5601 | 99.9875 | 186.64 | 116.395 | 93.136 | 85.987 | 102.2404 | 99.9875 | Ethyl methacrylate | 8 | Quad, WithOffset | 6 | 3.2526 | 0.99983 | TRUE | TRUE
1,1,2-Trichloroethane | 8.11 | 0.060 AN | 97 | 83 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7764 | 1.4581 | 4.3481 | 8.2372 | 23.9463 | 100.4854 | 77.64 | 72.905 | 86.962 | 82.372 | 95.7852 | 100.4854 | 1,1,2-Trichloroethane | 8.11 | Lin | 6 | 4.4034 | 0.99931 | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1156 | 2.0687 | 5.9257 | 9.7463 | 25.5769 | 99.8323 | 111.56 | 103.435 | 118.514 | 97.463 | 102.3076 | 99.8323 | Tetrachloroethene | 8.16 | Lin | 6 | 2.1209 | 0.99982 | TRUE | TRUE
1,3-Dichloropropane | 8.25 | 0.060 AN | 76 | 41 | 78 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7524 | 1.4029 | 4.4087 | 8.5132 | 24.9723 | 100.1996 | 75.24 | 70.145 | 88.174 | 85.132 | 99.8892 | 100.1996 | 1,3-Dichloropropane | 8.25 | Lin | 6 | 3.3241 | 0.9996 | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 2.9927 | 4.1878 | 9.8371 | 18.2724 | 50.7266 | 199.9796 | 149.635 | 104.695 | 98.371 | 91.362 | 101.4532 | 99.9898 | 2-Hexanone | 8.32 | Quad, WithOffset | 6 | 2.258 | 0.9999 | TRUE | TRUE
Dibromochloromethane | 8.435 | 0.050 AN | 129 | 127 | 131 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8143 | 1.497 | 4.5185 | 8.4976 | 25.8085 | 99.9755 | 81.43 | 74.85 | 90.37 | 84.976 | 103.234 | 99.9755 | Dibromochloromethane | 8.43 | Quad | 6 | 3.3377 | 0.99971 | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.52 | 0.049 AN | 107 | 109 | 93 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8045 | 1.4869 | 4.7842 | 8.8627 | 25.6208 | 99.9787 | 80.45 | 74.345 | 95.684 | 88.627 | 102.4832 | 99.9787 | 1,2-Dibromoethane (EDB) | 8.52 | Quad | 6 | 2.7208 | 0.9998 | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.92 | 0.040 AF | 117 | 82 | 52 | 0.1 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Chlorobenzene-d5 [IS3] | 8.92 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Chlorobenzene | 8.94 | 0.040 AN | 112 | 77 | 114 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9999 | 1.8044 | 5.2499 | 9.2138 | 26.5642 | 99.679 | 99.99 | 90.22 | 104.998 | 92.138 | 106.2568 | 99.679 | Chlorobenzene | 8.94 | Lin | 6 | 3.3803 | 0.99955 | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.02 | 0.040 AN | 131 | 133 | 117 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7965 | 1.512 | 4.4954 | 8.2454 | 25.9452 | 99.9694 | 79.65 | 75.6 | 89.908 | 82.454 | 103.7808 | 99.9694 | 1,1,1,2-Tetrachloroethane | 9.02 | Quad | 6 | 3.9756 | 0.99958 | TRUE | TRUE
Ethylbenzene | 9.03 | 0.040 AN | 91 | 106 | 51 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.092 | 2.0414 | 5.8943 | 10.0928 | 29.0169 | 98.94 | 109.2 | 102.07 | 117.886 | 100.928 | 116.0676 | 98.94 | Ethylbenzene | 9.03 | Lin | 6 | 7.7569 | 0.99748 | TRUE | TRUE
m/p-Xylene | 9.138 | 0.060 AN | 91 | 106 | 105 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 2.3665 | 4.1325 | 12.8023 | 21.9032 | 61.7504 | 196.7257 | 118.325 | 103.3125 | 128.023 | 109.516 | 123.5008 | 98.3628 | m/p-Xylene | 9.14 | Lin | 6 | 11.3406 | 0.99431 | TRUE | TRUE
o-Xylene | 9.435 | 0.060 AN | 91 | 106 | 105 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0786 | 1.932 | 5.9313 | 10.1595 | 28.0385 | 99.1784 | 107.86 | 96.6 | 118.626 | 101.595 | 112.154 | 99.1784 | o-Xylene | 9.43 | Lin | 6 | 6.0297 | 0.9985 | TRUE | TRUE
Styrene | 9.451 | 0.060 AN | 104 | 78 | 103 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0826 | 2.0345 | 6.2215 | 10.7436 | 30.365 | 98.5218 | 108.26 | 101.725 | 124.43 | 107.436 | 121.46 | 98.5218 | Styrene | 9.45 | Lin | 6 | 10.3506 | 0.99534 | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7542 | 1.5522 | 4.1711 | 7.8705 | 26.0966 | 99.9704 | 75.42 | 77.61 | 83.422 | 78.705 | 104.3864 | 99.9704 | Bromoform | 9.58 | Quad | 6 | 4.3791 | 0.99954 | TRUE | TRUE
Isopropylbenzene (cumene) | 9.72 | 0.060 AN | 105 | 120 | 79 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0498 | 1.9075 | 5.0931 | 9.4063 | 28.2333 | 99.2478 | 104.98 | 95.375 | 101.862 | 94.063 | 112.9332 | 99.2478 | Isopropylbenzene (cumene) | 9.72 | Lin | 6 | 6.2487 | 0.99843 | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.843 | 0.060 AN | 95 | 174 | 176 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 21.06 | 20.3055 | 18.5432 | 20.2308 | 20.434 | 19.4265 | 105.3 | 101.5275 | 92.716 | 101.154 | 102.17 | 97.1325 | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Bromobenzene | 9.943 | 0.040 AN | 77 | 156 | 158 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0277 | 1.8389 | 4.6613 | 9.0702 | 26.6527 | 99.6997 | 102.77 | 91.945 | 93.226 | 90.702 | 106.6108 | 99.6997 | Bromobenzene | 9.94 | Lin | 6 | 3.6725 | 0.99948 | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.954 | 0.040 AN | 83 | 85 | 95 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.7997 | 2.3674 | 4.699 | 8.4069 | 25.7565 | 99.9698 | 179.97 | 118.37 | 93.98 | 84.069 | 103.026 | 99.9698 | 1,1,2,2-Tetrachloroethane | 9.95 | Quad, WithOffset | 6 | 5.0178 | 0.99947 | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.988 | 0.040 AN | 77 | 110 | 61 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.3865 | 1.9717 | 4.4186 | 8.6961 | 26.8278 | 99.6992 | 138.65 | 98.585 | 88.372 | 86.961 | 107.3112 | 99.6992 | 1,2,3-Trichloropropane (TCP) | 9.99 | Lin, WithOffset | 6 | 5.099 | 0.99924 | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.988 | 0.040 AN | 75 | 53 | 89 | 0.05 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.5507 | 2.1527 | 4.6291 | 8.7841 | 26.0024 | 99.881 | 155.07 | 107.635 | 92.582 | 87.841 | 104.0096 | 99.881 | trans-1,4-Dichloro-2-butene | 9.99 | Lin, WithOffset | 6 | 3.7375 | 0.9996 | TRUE | TRUE
n-Propylbenzene | 10.023 | 0.026 AN | 91 | 120 | 65 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0899 | 1.9556 | 5.2915 | 10.2193 | 29.3603 | 98.8734 | 108.99 | 97.78 | 105.83 | 102.193 | 117.4412 | 98.8734 | n-Propylbenzene | 10.02 | Lin | 6 | 8.2612 | 0.99716 | TRUE | TRUE
2-Chlorotoluene | 10.08 | 0.031 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1367 | 2.0224 | 5.2655 | 10.0036 | 29.3666 | 98.8929 | 113.67 | 101.12 | 105.31 | 100.036 | 117.4664 | 98.8929 | 2-Chlorotoluene | 10.08 | Lin | 6 | 8.2585 | 0.99716 | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.159 | 0.040 AN | 105 | 120 | 119 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8838 | 1.7311 | 4.7561 | 9.3689 | 28.4167 | 99.2277 | 88.38 | 86.555 | 95.122 | 93.689 | 113.6668 | 99.2277 | 1,3,5-Trimethylbenzene | 10.16 | Lin | 6 | 6.6526 | 0.99824 | TRUE | TRUE
4-Chlorotoluene | 10.176 | 0.040 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1376 | 2.0184 | 5.436 | 10.363 | 31.2349 | 98.3814 | 113.76 | 100.92 | 108.72 | 103.63 | 124.9396 | 98.3814 | 4-Chlorotoluene | 10.17 | Lin | 6 | 11.6816 | 0.99412 | TRUE | TRUE
tert-Butylbenzene | 10.38 | 0.040 AN | 119 | 91 | 134 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9072 | 1.7381 | 4.6962 | 9.0267 | 26.5066 | 99.742 | 90.72 | 86.905 | 93.924 | 90.267 | 106.0264 | 99.742 | tert-Butylbenzene | 10.38 | Lin | 6 | 3.4964 | 0.99953 | TRUE | TRUE
Pentachloroethane | 10.398 | 0.030 AN | 167 | 130 | 132 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 2.026 | n.a./n.r. | 4.586 | 8.3446 | 25.6711 | n.a./n.r. | 202.6 | #VALUE! | 91.72 | 83.446 | 102.6844 | #VALUE! | Pentachloroethane | 10.4 | Quad, WithOffset | 6 | 3.952 | 0.99974 | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.422 | 0.030 AN | 105 | 120 | 77 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9133 | 1.795 | 4.8329 | 9.7897 | 28.7726 | 99.0912 | 91.33 | 89.75 | 96.658 | 97.897 | 115.0904 | 99.0912 | 1,2,4-Trimethylbenzene | 10.42 | Lin | 6 | 7.2003 | 0.99791 | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.535 | 0.037 AN | 105 | 134 | 91 | 0.08 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9812 | 1.8595 | 5.1512 | 9.9055 | 28.8379 | 99.0454 | 98.12 | 92.975 | 103.024 | 99.055 | 115.3516 | 99.0454 | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | Lin | 6 | 7.2915 | 0.99783 | TRUE | TRUE
1,3-Dichlorobenzene | 10.61 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9856 | 1.7979 | 4.6393 | 8.9065 | 25.7522 | 99.9435 | 98.56 | 89.895 | 92.786 | 89.065 | 103.0088 | 99.9435 | 1,3-Dichlorobenzene | 10.61 | Lin | 6 | 2.6287 | 0.99974 | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9873 | 1.9464 | 5.1976 | 10.2782 | 29.905 | 98.7372 | 98.73 | 97.32 | 103.952 | 102.782 | 119.62 | 98.7372 | 4-Isopropyltoluene (p-cymene) | 10.64 | Lin | 6 | 9.2636 | 0.99641 | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 0.036 AN | 152 | 115 |  | 0.07 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1,4-Dichlorobenzene | 10.687 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1056 | 2.0068 | 5.2658 | 10.0131 | 28.3836 | 99.1383 | 110.56 | 100.34 | 105.316 | 100.131 | 113.5344 | 99.1383 | 1,4-Dichlorobenzene | 10.68 | Lin | 6 | 6.4424 | 0.9983 | TRUE | TRUE
n-Butylbenzene | 10.923 | 0.040 AN | 91 | 92 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9477 | 1.7742 | 5.0798 | 9.8905 | 29.4819 | 98.8915 | 94.77 | 88.71 | 101.596 | 98.905 | 117.9276 | 98.8915 | n-Butylbenzene | 10.92 | Lin | 6 | 8.4959 | 0.99704 | TRUE | TRUE
1,2-Dichlorobenzene | 10.93 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0132 | 1.916 | 5.0517 | 9.6224 | 28.6898 | 99.1143 | 101.32 | 95.8 | 101.034 | 96.224 | 114.7592 | 99.1143 | 1,2-Dichlorobenzene | 10.93 | Lin | 6 | 7.0394 | 0.99799 | TRUE | TRUE
Hexachloroethane | 11.11 | 0.040 AN | 117 | 119 | 201 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8399 | 1.5419 | 4.286 | 8.4836 | 25.8627 | 99.9735 | 83.99 | 77.095 | 85.72 | 84.836 | 103.4508 | 99.9735 | Hexachloroethane | 11.11 | Quad | 6 | 3.5218 | 0.99968 | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 0.040 AN | 157 | 155 | 75 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.7417 | 2.3593 | 4.6949 | 8.6328 | 25.5797 | 99.9848 | 174.17 | 117.965 | 93.898 | 86.328 | 102.3188 | 99.9848 | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | Quad, WithOffset | 6 | 3.495 | 0.99978 | TRUE | TRUE
Nitrobenzene | 11.58 | 0.040 AN | 77 | 51 | 123 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | 2.2862 | 4.4202 | 8.7467 | 25.624 | 99.9801 | #VALUE! | 114.31 | 88.404 | 87.467 | 102.496 | 99.9801 | Nitrobenzene | 11.59 | Quad, WithOffset | 6 | 4.0491 | 0.99969 | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.98 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7733 | 1.5076 | 4.2242 | 8.0622 | 24.0736 | 100.4763 | 77.33 | 75.38 | 84.484 | 80.622 | 96.2944 | 100.4763 | 1,2,4-Trichlorobenzene | 11.98 | Lin | 6 | 4.6192 | 0.99924 | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.07 | 0.040 AN | 225 | 227 | 223 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8083 | 1.4845 | 4.2599 | 7.8118 | 23.2068 | 100.7164 | 80.83 | 74.225 | 85.198 | 78.118 | 92.8272 | 100.7164 | Hexachloro-1,3-butadiene | 12.07 | Lin | 6 | 5.9387 | 0.99876 | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7893 | 1.4692 | 4.2602 | 8.1867 | 25.6774 | 100.0617 | 78.93 | 73.46 | 85.204 | 81.867 | 102.7096 | 100.0617 | Naphthalene | 12.15 | Lin | 6 | 4.1085 | 0.99938 | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.29 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8213 | 1.5449 | 4.1905 | 8.0172 | 24.69 | 100.3271 | 82.13 | 77.245 | 83.81 | 80.172 | 98.76 | 100.3271 | 1,2,3-Trichlorobenzene | 12.29 | Lin | 6 | 4.3107 | 0.99933 | TRUE | TRUE

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 2
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 2
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 2
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | 20
Pentafluorobenzene [IS1] | 20
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | 20
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 5
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 2
Toluene-d8 [SS2] | 20
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 2
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | 20
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 2
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | 20
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | 20
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 2
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.47 | 382 | 0.04 | 0.086 | Not confirmed | 50 | 52 | 32.87 | 27.41 | Confirmed | 49 | 9.4 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | 1.57 | 73 | 0.01 | 0.013 | Not confirmed | 62 | 64 | 31.72 | n.a. | Not confirmed | 61 | 7.92 | 0.87 | Confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.86 | 1953 | 0.19 | 0.404 | Confirmed | 94 | 96 | 93.44 | 96.02 | Confirmed | 93 | 20.09 | 11.09 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 30.86 | n.a. | Not confirmed | 49 | 21.2 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | 2.2 | 35 | 0 | 0.008 | Not confirmed | 101 | 103 | 63.84 | n.a. | Not confirmed | 105 | 9.93 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 76.45 | n.a. | Not confirmed | 45 | 67.36 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | 2.75 | 54 | 0.01 | 0.011 | Not confirmed | 61 | 96 | 68.4 | 70.28 | Confirmed | 98 | 42.11 | 86.79 | Not confirmed
2 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 39.42 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Iodomethane | 2.91 | 528 | 0.05 | 0.372 | Confirmed | 142 | 127 | 31.33 | 15.86 | Confirmed | 141 | 12.32 | 7.75 | Confirmed
1 | TRUE | FALSE | FALSE | Carbon disulfide | 2.97 | 1304 | 0.13 | 0.122 | Confirmed | 76 | 78 | 8.49 | 6.65 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 52.37 | n.a. | Not confirmed | 76 | 35.15 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Methylene chloride (DCM) | 3.39 | 420 | 0.04 | 0.07 | Confirmed | 49 | 84 | 87.09 | 60.7 | Confirmed | 86 | 56.06 | 52.33 | Confirmed
1 | TRUE | FALSE | FALSE | trans-1,2-Dichloroethene | 3.69 | 193 | 0.02 | 0.038 | Confirmed | 61 | 96 | 65.93 | 53 | Confirmed | 98 | 42.43 | 28.28 | Confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 41.7 | n.a. | Not confirmed | 57 | 23.28 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | 4.21 | 25 | 0 | 0.004 | Not confirmed | 63 | 65 | 32.18 | n.a. | Not confirmed | 83 | 11.21 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 157 | n.a. | Not confirmed | 79 | 30.84 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | cis-1,2-Dichloroethene | 4.84 | 116 | 0.01 | 0.021 | Confirmed | 61 | 96 | 71.05 | 43.55 | Confirmed | 98 | 45.12 | 53.24 | Confirmed
2 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 26.63 | n.a. | Not confirmed | 57 | 6.42 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 13.43 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 27.47 | n.a. | Not confirmed | 40 | 33.89 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | Bromochloromethane | 5.08 | 36 | 0 | 0.011 | Not confirmed | 49 | 130 | 67.47 | 127.75 | Not confirmed | 128 | 52.42 | 114.29 | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 40.16 | n.a. | Not confirmed | 71 | 41.81 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | 5.21 | 39 | 0 | 0.007 | Not confirmed | 83 | 85 | 64.62 | n.a. | Not confirmed | 47 | 18.21 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | 5.34 | 12 | 0 | 0.005 | Not confirmed | 97 | 99 | 62.69 | n.a. | Not confirmed | 61 | 78.38 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.37 | 51671 | 5.04 | 21.088 | Confirmed | 113 | 111 | 104.25 | 100.66 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.44 | 90998 | 8.88 | 20 | Confirmed | 168 | 99 | 67.18 | 64.58 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 56 | 41 | 63.4 | n.a. | Not confirmed | 43 | 24.09 | n.a. | Not confirmed
1 | FALSE | FALSE | TRUE | Carbon tetrachloride | 5.5 | 12 | 0 | 0.603 | Confirmed | 119 | 121 | 31.97 | 9.7 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | 1,1-Dichloropropene | 5.52 | 114 | 0.01 | 0.026 | Confirmed | 75 | 77 | 30.69 | 43.33 | Confirmed | 110 | 35.51 | 11.54 | Confirmed
1 | TRUE | TRUE | FALSE | Benzene | 5.72 | 209 | 0.02 | 0.014 | Not confirmed | 78 | 77 | 24.84 | n.a. | Not confirmed | 52 | 15.16 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | 5.79 | 76 | 0.01 | 0.019 | Not confirmed | 62 | 64 | 32.11 | n.a. | Not confirmed | 49 | 31.58 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.18 | 171169 | 16.7 | 20 | Confirmed | 114 | 88 | 20.74 | 21.09 | Confirmed | 63 | 21.84 | 21.44 | Confirmed
1 | TRUE | TRUE | FALSE | Trichloroethene | 6.39 | 96 | 0.01 | 0.031 | Not confirmed | 130 | 132 | 95.7 | 68.57 | Confirmed | 95 | 109.17 | 66.06 | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 69.35 | n.a. | Not confirmed | 41 | 53.36 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.73 | 62 | 0.01 | 0.038 | Not confirmed | 174 | 93 | 129.88 | 112.37 | Confirmed | 95 | 107.3 | 69.39 | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 92.09 | n.a. | Not confirmed | 39 | 42.41 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | 6.91 | 31 | 0 | 0.009 | Not confirmed | 83 | 85 | 62.69 | n.a. | Not confirmed | 47 | 15.96 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 94.62 | n.a. | Not confirmed | 39 | 24.38 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | 7.37 | 66 | 0.01 | 0.025 | Not confirmed | 75 | 39 | 39.52 | n.a. | Not confirmed | 77 | 31.53 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 39.83 | n.a. | Not confirmed | 41 | 22.3 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.61 | 247988 | 24.19 | 19.949 | Confirmed | 98 | 100 | 60.5 | 59.17 | Confirmed | 70 | 12.15 | 12.12 | Confirmed
1 | TRUE | FALSE | FALSE | Toluene | 7.68 | 549 | 0.05 | 0.034 | Confirmed | 91 | 92 | 54.03 | 64.48 | Confirmed | 65 | 12.34 | 17.77 | Confirmed
1 | FALSE | TRUE | FALSE | trans-1,3-Dichloropropene | 7.93 | 61 | 0.01 | 1.034 | Not confirmed | 75 | 39 | 40.52 | n.a. | Not confirmed | 77 | 30.97 | n.a. | Not confirmed
1 | FALSE | TRUE | FALSE | Ethyl methacrylate | 8 | 35 | 0 | 1.385 | Not confirmed | 69 | 41 | 58.74 | n.a. | Not confirmed | 99 | 20.87 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | 8.1 | 23 | 0 | 0.008 | Not confirmed | 97 | 83 | 91.41 | n.a. | Not confirmed | 99 | 61.61 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.16 | 133 | 0.01 | 0.038 | Confirmed | 166 | 164 | 79.25 | 76.55 | Confirmed | 129 | 88.26 | 84.09 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3-Dichloropropane | 8.25 | 85 | 0.01 | 0.016 | Not confirmed | 76 | 41 | 60.52 | n.a. | Not confirmed | 78 | 32.15 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 57.4 | n.a. | Not confirmed | 57 | 19.64 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | 8.44 | 12 | 0 | 0.007 | Not confirmed | 129 | 127 | 76.5 | 142.66 | Not confirmed | 131 | 22.9 | 54.69 | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 92.55 | n.a. | Not confirmed | 93 | 4.26 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.92 | 156538 | 15.27 | 20 | Confirmed | 117 | 82 | 67.96 | 67.09 | Confirmed | 52 | 16.73 | 15.98 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.94 | 373 | 0.04 | 0.042 | Not confirmed | 112 | 77 | 74.99 | 212.56 | Not confirmed | 114 | 31.2 | 29.61 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | 9.02 | 9 | 0 | 0.006 | Not confirmed | 131 | 133 | 93.93 | n.a. | Not confirmed | 117 | 86.73 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Ethylbenzene | 9.03 | 480 | 0.05 | 0.033 | Confirmed | 91 | 106 | 32.43 | 31.85 | Confirmed | 51 | 9.37 | 13.89 | Confirmed
1 | TRUE | FALSE | FALSE | m/p-Xylene | 9.13 | 1042 | 0.1 | 0.088 | Confirmed | 91 | 106 | 46.77 | 40.23 | Confirmed | 105 | 20.4 | 21.04 | Confirmed
1 | TRUE | FALSE | FALSE | o-Xylene | 9.43 | 497 | 0.05 | 0.038 | Confirmed | 91 | 106 | 45.31 | 34.63 | Confirmed | 105 | 23.13 | 16.76 | Confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.45 | 357 | 0.03 | 0.04 | Confirmed | 104 | 78 | 61.9 | 68.42 | Confirmed | 103 | 56.81 | 57.58 | Confirmed
1 | TRUE | FALSE | FALSE | Bromoform | 9.58 | 15 | 0 | 0.019 | Confirmed | 173 | 171 | 50.55 | 33.34 | Confirmed | 175 | 46.49 | 40.97 | Confirmed
1 | TRUE | TRUE | FALSE | Isopropylbenzene (cumene) | 9.71 | 429 | 0.04 | 0.029 | Not confirmed | 105 | 120 | 27.2 | 22.68 | Confirmed | 79 | 17.55 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 91059 | 8.88 | 20.53 | Confirmed | 95 | 174 | 52.45 | 56.54 | Confirmed | 176 | 50.78 | 53.58 | Confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.94 | 350 | 0.03 | 0.05 | Confirmed | 77 | 156 | 44.57 | 55.83 | Confirmed | 158 | 42.71 | 41.37 | Confirmed
1 | FALSE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | 9.95 | 84 | 0.01 | 1.125 | Not confirmed | 83 | 85 | 63.54 | 64.42 | Confirmed | 95 | 13.15 | 35.49 | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 110 | 76.82 | n.a. | Not confirmed | 61 | 67.36 | n.a. | Not confirmed
1 | FALSE | TRUE | FALSE | trans-1,4-Dichloro-2-butene | 9.99 | 108 | 0.01 | 0.839 | Not confirmed | 75 | 53 | 22.86 | n.a. | Not confirmed | 89 | 4.9 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | n-Propylbenzene | 10.02 | 883 | 0.09 | 0.051 | Confirmed | 91 | 120 | 22.49 | 20.16 | Confirmed | 65 | 10.91 | 9.67 | Confirmed
1 | TRUE | FALSE | FALSE | 2-Chlorotoluene | 10.08 | 555 | 0.05 | 0.053 | Confirmed | 91 | 126 | 31.37 | 26.43 | Confirmed | 89 | 17.52 | 15.59 | Confirmed
1 | TRUE | FALSE | FALSE | 1,3,5-Trimethylbenzene | 10.15 | 405 | 0.04 | 0.031 | Confirmed | 105 | 120 | 43.91 | 43.94 | Confirmed | 119 | 10.33 | 14.11 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Chlorotoluene | 10.17 | 759 | 0.07 | 0.065 | Confirmed | 91 | 126 | 27.59 | 27.81 | Confirmed | 89 | 11.19 | 11.15 | Confirmed
1 | TRUE | FALSE | FALSE | tert-Butylbenzene | 10.38 | 321 | 0.03 | 0.027 | Confirmed | 119 | 91 | 76.37 | 90.35 | Confirmed | 134 | 23.56 | 21.32 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 167 | 130 | 145.01 | n.a. | Not confirmed | 132 | 144.82 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,4-Trimethylbenzene | 10.42 | 540 | 0.05 | 0.043 | Not confirmed | 105 | 120 | 41.67 | 46.94 | Confirmed | 77 | 12.21 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 701 | 0.07 | 0.043 | Not confirmed | 105 | 134 | 18.5 | 17.05 | Confirmed | 91 | 15.72 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.61 | 519 | 0.05 | 0.082 | Confirmed | 146 | 148 | 62.47 | 66.18 | Confirmed | 111 | 50.97 | 46.33 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.64 | 531 | 0.05 | 0.044 | Confirmed | 119 | 91 | 31.48 | 28.23 | Confirmed | 134 | 27.22 | 26.22 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 72030 | 7.03 | 20 | Confirmed | 152 | 115 | 75.75 | 75.14 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.68 | 646 | 0.06 | 0.113 | Not confirmed | 146 | 148 | 63.45 | 77.65 | Confirmed | 111 | 55.36 | 181.44 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.92 | 914 | 0.09 | 0.069 | Confirmed | 91 | 92 | 52.22 | 52.15 | Confirmed | 134 | 23.38 | 19.33 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.93 | 461 | 0.04 | 0.078 | Confirmed | 146 | 148 | 62.83 | 64.04 | Confirmed | 111 | 51.55 | 42.02 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 96.74 | n.a. | Not confirmed | 201 | 58.13 | n.a. | Not confirmed
1 | FALSE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 28 | 0 | 1.228 | Not confirmed | 157 | 155 | 76.5 | n.a. | Not confirmed | 75 | 130.93 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 49.24 | n.a. | Not confirmed | 123 | 41.22 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.98 | 556 | 0.05 | 0.148 | Confirmed | 180 | 182 | 94.66 | 90.84 | Confirmed | 145 | 41.13 | 39.57 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.07 | 109 | 0.01 | 0.076 | Confirmed | 225 | 227 | 64.13 | 52.27 | Confirmed | 223 | 63.42 | 62.39 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.15 | 1911 | 0.19 | 0.149 | Confirmed | 128 | 127 | 12.85 | 14.27 | Confirmed | 129 | 10.36 | 10.12 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.29 | 524 | 0.05 | 0.145 | Confirmed | 180 | 182 | 94.11 | 91.49 | Confirmed | 145 | 43.31 | 40.89 | Confirmed

## CCV
[11 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? |  |  | 70-130 | 80-120 |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.43 | 96447 | 5.44 | 133044 | TRUE | TRUE | % of TV | True Value | Pass? |  | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.18 | 182958 | 6.18 | 263138 | TRUE | TRUE | 77.8 | 10 | TRUE | FALSE | Chloromethane (methyl chloride) | 1.47 | 35766 | 0.35 | 7.78 | Confirmed | 50 | 52 | 32.87 | 33.87 | Confirmed | 49 | 9.4 | 9.3 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 171399 | 8.92 | 236091 | TRUE | TRUE | 82.71 | 10 | TRUE | TRUE | Chloroethene (vinyl chloride) | 1.57 | 49912 | 0.48 | 8.271 | Confirmed | 62 | 64 | 31.72 | 32.51 | Confirmed | 61 | 7.92 | 8.05 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 81811 | 10.67 | 100630 | TRUE | TRUE | 120.18 | 10 | TRUE | FALSE | Bromomethane (methyl bromide) | 1.85 | 57804 | 0.56 | 12.018 | Confirmed | 94 | 96 | 93.44 | 97.14 | Confirmed | 93 | 20.09 | 20.38 | Confirmed
[7 empty] | 104.13 | 10 | TRUE | TRUE | Chloroethane (ethyl chloride) | 1.96 | 35118 | 0.34 | 10.413 | Confirmed | 64 | 66 | 30.86 | 33.82 | Confirmed | 49 | 21.2 | 21.54 | Confirmed
[7 empty] | 96.63 | 10 | TRUE | TRUE | Trichlorofluoromethane | 2.2 | 41986 | 0.41 | 9.663 | Confirmed | 101 | 103 | 63.84 | 65.26 | Confirmed | 105 | 9.93 | 9.88 | Confirmed
Total Analytes | 81 | [5 empty] | 88.71 | 10 | TRUE | TRUE | Diethyl ether | 2.51 | 31453 | 0.3 | 8.871 | Confirmed | 59 | 74 | 76.45 | 78.45 | Confirmed | 45 | 67.36 | 72.27 | Confirmed
Failed | 1 | [5 empty] | 90.84 | 10 | TRUE | TRUE | 1,1-Dichloroethene | 2.75 | 47141 | 0.45 | 9.084 | Confirmed | 61 | 96 | 68.4 | 69.11 | Confirmed | 98 | 42.11 | 42.87 | Confirmed
| 7 | [5 empty] | 86.0056 | 18 | TRUE | TRUE | Acetone | 2.84 | 19095 | 0.18 | 15.481 | Confirmed | 43 | 58 | 39.42 | 35.78 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 63.52 | 10 | FALSE | FALSE | Iodomethane | 2.9 | 20296 | 0.2 | 6.352 | Confirmed | 142 | 127 | 31.33 | 30.46 | Confirmed | 141 | 12.32 | 13.64 | Confirmed
[7 empty] | 88.76 | 10 | TRUE | TRUE | Carbon disulfide | 2.97 | 100130 | 0.97 | 8.876 | Confirmed | 76 | 78 | 8.49 | 8.65 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 85.13 | 10 | TRUE | TRUE | 3-Chloropropene (allyl chloride) | 3.21 | 37043 | 0.36 | 8.513 | Confirmed | 41 | 39 | 52.37 | 54.77 | Confirmed | 76 | 35.15 | 36.19 | Confirmed
[7 empty] | 76.93 | 10 | TRUE | FALSE | Methylene chloride (DCM) | 3.37 | 48156 | 0.46 | 7.693 | Confirmed | 49 | 84 | 87.09 | 89.18 | Confirmed | 86 | 56.06 | 57.04 | Confirmed
[7 empty] | 87.95 | 10 | TRUE | TRUE | trans-1,2-Dichloroethene | 3.69 | 47327 | 0.46 | 8.795 | Confirmed | 61 | 96 | 65.93 | 67.9 | Confirmed | 98 | 42.43 | 43.23 | Confirmed
[7 empty] | 82.13 | 10 | TRUE | TRUE | Methyl tert-butyl ether (MTBE) | 3.7 | 58638 | 0.57 | 8.213 | Confirmed | 73 | 41 | 41.7 | 40.68 | Confirmed | 57 | 23.28 | 23.95 | Confirmed
[7 empty] | 90.78 | 10 | TRUE | TRUE | 1,1-Dichloroethane | 4.2 | 55806 | 0.54 | 9.078 | Confirmed | 63 | 65 | 32.18 | 31.87 | Confirmed | 83 | 11.21 | 11.59 | Confirmed
[7 empty] | 78.89 | 10 | TRUE | FALSE | 2,2-Dichloropropane | 4.82 | 16024 | 0.15 | 7.889 | Confirmed | 77 | 41 | 157 | 145.36 | Confirmed | 79 | 30.84 | 31.67 | Confirmed
[7 empty] | 86.96 | 10 | TRUE | TRUE | cis-1,2-Dichloroethene | 4.83 | 51414 | 0.5 | 8.696 | Confirmed | 61 | 96 | 71.05 | 73.46 | Confirmed | 98 | 45.12 | 46.06 | Confirmed
[7 empty] | 91.6667 | 18 | TRUE | TRUE | 2-Butanone (MEK) | 4.85 | 29135 | 0.28 | 16.5 | Confirmed | 43 | 72 | 26.63 | 25.39 | Confirmed | 57 | 6.42 | 6.61 | Confirmed
[7 empty] | 90.82 | 10 | TRUE | TRUE | Methyl acrylate | 4.94 | 16256 | 0.16 | 9.082 | Confirmed | 55 | 85 | 13.43 | 15.33 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 78.36 | 10 | TRUE | FALSE | Methacrylonitrile | 5.07 | 16801 | 0.16 | 7.836 | Confirmed | 67 | 52 | 27.47 | 30.18 | Confirmed | 40 | 33.89 | 34.41 | Confirmed
[7 empty] | 88.4 | 10 | TRUE | TRUE | Bromochloromethane | 5.08 | 30601 | 0.3 | 8.84 | Confirmed | 49 | 130 | 67.47 | 71.73 | Confirmed | 128 | 52.42 | 57.13 | Confirmed
[7 empty] | 78.53 | 10 | TRUE | FALSE | Tetrahydrofuran | 5.09 | 10447 | 0.1 | 7.853 | Confirmed | 42 | 72 | 40.16 | 39 | Confirmed | 71 | 41.81 | 44.01 | Confirmed
[7 empty] | 94.76 | 10 | TRUE | TRUE | Trichloromethane (chloroform) | 5.21 | 53994 | 0.52 | 9.476 | Confirmed | 83 | 85 | 64.62 | 64.81 | Confirmed | 47 | 18.21 | 18.56 | Confirmed
[7 empty] | 88.34 | 10 | TRUE | TRUE | 1,1,1-Trichloroethane | 5.34 | 24982 | 0.24 | 8.834 | Confirmed | 97 | 99 | 62.69 | 61.64 | Confirmed | 61 | 78.38 | 69.95 | Confirmed
[7 empty] | 98.565 | 20 | TRUE | TRUE | Dibromofluoromethane [SS1] | 5.37 | 51195 | 0.49 | 19.713 | Confirmed | 113 | 111 | 104.25 | 102.65 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | TRUE | Pentafluorobenzene [IS1] | 5.43 | 96447 | 0.93 | 20 | Confirmed | 168 | 99 | 67.18 | 63.2 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 81.43 | 10 | TRUE | TRUE | 1-Chlorobutane (butyl chloride) | 5.49 | 49911 | 0.48 | 8.143 | Confirmed | 56 | 41 | 63.4 | 62 | Confirmed | 43 | 24.09 | 24.6 | Confirmed
[7 empty] | 84.01 | 10 | TRUE | TRUE | Carbon tetrachloride | 5.5 | 18065 | 0.17 | 8.401 | Confirmed | 119 | 121 | 31.97 | 32.43 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 86.7 | 10 | TRUE | TRUE | 1,1-Dichloropropene | 5.52 | 40814 | 0.39 | 8.67 | Confirmed | 75 | 77 | 30.69 | 30.54 | Confirmed | 110 | 35.51 | 35.96 | Confirmed
[7 empty] | 91.97 | 10 | TRUE | TRUE | Benzene | 5.71 | 145955 | 1.41 | 9.197 | Confirmed | 78 | 77 | 24.84 | 25.29 | Confirmed | 52 | 15.16 | 14.72 | Confirmed
[7 empty] | 87.87 | 10 | TRUE | TRUE | 1,2-Dichloroethane | 5.78 | 37256 | 0.36 | 8.787 | Confirmed | 62 | 64 | 32.11 | 32.15 | Confirmed | 49 | 31.58 | 30.48 | Confirmed
[7 empty] | 100 | 20 | TRUE | TRUE | 1,4-Difluorobenzene [IS2] | 6.18 | 182958 | 1.77 | 20 | Confirmed | 114 | 88 | 20.74 | 20.78 | Confirmed | 63 | 21.84 | 21.3 | Confirmed
[7 empty] | 93.09 | 10 | TRUE | TRUE | Trichloroethene | 6.39 | 31377 | 0.3 | 9.309 | Confirmed | 130 | 132 | 95.7 | 95.61 | Confirmed | 95 | 109.17 | 107.33 | Confirmed
[7 empty] | 86.84 | 10 | TRUE | TRUE | 1,2-Dichloropropane | 6.65 | 30593 | 0.3 | 8.684 | Confirmed | 63 | 62 | 69.35 | 68.04 | Confirmed | 41 | 53.36 | 50.41 | Confirmed
[7 empty] | 98.76 | 10 | TRUE | TRUE | Dibromomethane | 6.73 | 17845 | 0.17 | 9.876 | Confirmed | 174 | 93 | 129.88 | 113.53 | Confirmed | 95 | 107.3 | 97.53 | Confirmed
[7 empty] | 100.77 | 10 | TRUE | TRUE | Methyl methacrylate | 6.74 | 13277 | 0.13 | 10.077 | Confirmed | 41 | 69 | 92.09 | 101.04 | Confirmed | 39 | 42.41 | 46.3 | Confirmed
[7 empty] | 79.67 | 10 | TRUE | FALSE | Bromodichloromethane | 6.92 | 29145 | 0.28 | 7.967 | Confirmed | 83 | 85 | 62.69 | 64.89 | Confirmed | 47 | 15.96 | 15.78 | Confirmed
[7 empty] | 83.13 | 10 | TRUE | TRUE | 2-Nitropropane | 7.15 | 2478 | 0.02 | 8.313 | Confirmed | 43 | 41 | 94.62 | 96.37 | Confirmed | 39 | 24.38 | 26.38 | Confirmed
[7 empty] | 91.49 | 10 | TRUE | TRUE | cis-1,3-Dichloropropene | 7.36 | 27802 | 0.27 | 9.149 | Confirmed | 75 | 39 | 39.52 | 41.49 | Confirmed | 77 | 31.53 | 30.65 | Confirmed
[7 empty] | 97.6278 | 18 | TRUE | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.52 | 56455 | 0.54 | 17.573 | Confirmed | 43 | 58 | 39.83 | 41.49 | Confirmed | 41 | 22.3 | 22.01 | Confirmed
[7 empty] | 99.53 | 20 | TRUE | TRUE | Toluene-d8 [SS2] | 7.61 | 270938 | 2.61 | 19.906 | Confirmed | 98 | 100 | 60.5 | 58.81 | Confirmed | 70 | 12.15 | 11.92 | Confirmed
[7 empty] | 86.96 | 10 | TRUE | TRUE | Toluene | 7.68 | 155054 | 1.5 | 8.696 | Confirmed | 91 | 92 | 54.03 | 55.03 | Confirmed | 65 | 12.34 | 12.55 | Confirmed
[7 empty] | 93.02 | 10 | TRUE | TRUE | trans-1,3-Dichloropropene | 7.93 | 18815 | 0.18 | 9.302 | Confirmed | 75 | 39 | 40.52 | 40.89 | Confirmed | 77 | 30.97 | 31.3 | Confirmed
[7 empty] | 104.39 | 10 | TRUE | TRUE | Ethyl methacrylate | 7.99 | 21201 | 0.2 | 10.439 | Confirmed | 69 | 41 | 58.74 | 59.97 | Confirmed | 99 | 20.87 | 21.7 | Confirmed
[7 empty] | 80.72 | 10 | TRUE | TRUE | 1,1,2-Trichloroethane | 8.1 | 27398 | 0.26 | 8.072 | Confirmed | 97 | 83 | 91.41 | 94.63 | Confirmed | 99 | 61.61 | 62.7 | Confirmed
[7 empty] | 97.85 | 10 | TRUE | TRUE | Tetrachloroethene | 8.16 | 37349 | 0.36 | 9.785 | Confirmed | 166 | 164 | 79.25 | 80.99 | Confirmed | 129 | 88.26 | 85.56 | Confirmed
[7 empty] | 84.53 | 10 | TRUE | TRUE | 1,3-Dichloropropane | 8.24 | 50615 | 0.49 | 8.453 | Confirmed | 76 | 41 | 60.52 | 60.22 | Confirmed | 78 | 32.15 | 30.89 | Confirmed
[7 empty] | 104.0889 | 18 | TRUE | TRUE | 2-Hexanone | 8.31 | 35260 | 0.34 | 18.736 | Confirmed | 43 | 58 | 57.4 | 57.26 | Confirmed | 57 | 19.64 | 19.74 | Confirmed
[7 empty] | 92.76 | 10 | TRUE | TRUE | Dibromochloromethane | 8.43 | 17772 | 0.17 | 9.276 | Confirmed | 129 | 127 | 76.5 | 76.95 | Confirmed | 131 | 22.9 | 24.39 | Confirmed
[7 empty] | 91.87 | 10 | TRUE | TRUE | 1,2-Dibromoethane (EDB) | 8.52 | 20991 | 0.2 | 9.187 | Confirmed | 107 | 109 | 92.55 | 95.26 | Confirmed | 93 | 4.26 | 4.48 | Confirmed
[7 empty] | 100 | 20 | TRUE | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 171399 | 1.65 | 20 | Confirmed | 117 | 82 | 67.96 | 68.51 | Confirmed | 52 | 16.73 | 16.2 | Confirmed
[7 empty] | 89.82 | 10 | TRUE | TRUE | Chlorobenzene | 8.94 | 86796 | 0.84 | 8.982 | Confirmed | 112 | 77 | 74.99 | 72.74 | Confirmed | 114 | 31.2 | 31.41 | Confirmed
[7 empty] | 89.92 | 10 | TRUE | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 16182 | 0.16 | 8.992 | Confirmed | 131 | 133 | 93.93 | 93.27 | Confirmed | 117 | 86.73 | 81.25 | Confirmed
[7 empty] | 95.8 | 10 | TRUE | TRUE | Ethylbenzene | 9.02 | 153374 | 1.48 | 9.58 | Confirmed | 91 | 106 | 32.43 | 32.77 | Confirmed | 51 | 9.37 | 9.26 | Confirmed
[7 empty] | 105.93 | 20 | TRUE | TRUE | m/p-Xylene | 9.13 | 273438 | 2.64 | 21.186 | Confirmed | 91 | 106 | 46.77 | 46.22 | Confirmed | 105 | 20.4 | 20.86 | Confirmed
[7 empty] | 96.58 | 10 | TRUE | TRUE | o-Xylene | 9.43 | 137350 | 1.33 | 9.658 | Confirmed | 91 | 106 | 45.31 | 45.3 | Confirmed | 105 | 23.13 | 23.93 | Confirmed
[7 empty] | 104.64 | 10 | TRUE | TRUE | Styrene | 9.45 | 101904 | 0.98 | 10.464 | Confirmed | 104 | 78 | 61.9 | 61.42 | Confirmed | 103 | 56.81 | 57.12 | Confirmed
[7 empty] | 95.97 | 10 | TRUE | TRUE | Bromoform | 9.58 | 8913 | 0.09 | 9.597 | Confirmed | 173 | 171 | 50.55 | 51.2 | Confirmed | 175 | 46.49 | 47.15 | Confirmed
[7 empty] | 90.91 | 10 | TRUE | TRUE | Isopropylbenzene (cumene) | 9.71 | 150574 | 1.45 | 9.091 | Confirmed | 105 | 120 | 27.2 | 28.02 | Confirmed | 79 | 17.55 | 17.25 | Confirmed
[7 empty] | 92.8 | 20 | TRUE | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 93500 | 0.9 | 18.56 | Confirmed | 95 | 174 | 52.45 | 59.09 | Confirmed | 176 | 50.78 | 55.85 | Confirmed
[7 empty] | 84 | 10 | TRUE | TRUE | Bromobenzene | 9.94 | 66146 | 0.64 | 8.4 | Confirmed | 77 | 156 | 44.57 | 48.78 | Confirmed | 158 | 42.71 | 47.15 | Confirmed
[7 empty] | 85.42 | 10 | TRUE | TRUE | 1,1,2,2-Tetrachloroethane | 9.95 | 34030 | 0.33 | 8.542 | Confirmed | 83 | 85 | 63.54 | 64.46 | Confirmed | 95 | 13.15 | 12.69 | Confirmed
[7 empty] | 88.06 | 10 | TRUE | TRUE | 1,2,3-Trichloropropane (TCP) | 9.98 | 11779 | 0.11 | 8.806 | Confirmed | 77 | 110 | 76.82 | 78.44 | Confirmed | 61 | 67.36 | 67.86 | Confirmed
[7 empty] | 87.92 | 10 | TRUE | TRUE | trans-1,4-Dichloro-2-butene | 9.98 | 35054 | 0.34 | 8.792 | Confirmed | 75 | 53 | 22.86 | 20.18 | Confirmed | 89 | 4.9 | 4.73 | Confirmed
[7 empty] | 94.84 | 10 | TRUE | TRUE | n-Propylbenzene | 10.02 | 186008 | 1.8 | 9.484 | Confirmed | 91 | 120 | 22.49 | 22.92 | Confirmed | 65 | 10.91 | 11.02 | Confirmed
[7 empty] | 95.11 | 10 | TRUE | TRUE | 2-Chlorotoluene | 10.08 | 112076 | 1.08 | 9.511 | Confirmed | 91 | 126 | 31.37 | 31.37 | Confirmed | 89 | 17.52 | 17.77 | Confirmed
[7 empty] | 92.06 | 10 | TRUE | TRUE | 1,3,5-Trimethylbenzene | 10.15 | 137047 | 1.32 | 9.206 | Confirmed | 105 | 120 | 43.91 | 45.15 | Confirmed | 119 | 10.33 | 10.76 | Confirmed
[7 empty] | 101.24 | 10 | TRUE | TRUE | 4-Chlorotoluene | 10.17 | 135200 | 1.3 | 10.124 | Confirmed | 91 | 126 | 27.59 | 28.15 | Confirmed | 89 | 11.19 | 11.27 | Confirmed
[7 empty] | 90.06 | 10 | TRUE | TRUE | tert-Butylbenzene | 10.37 | 119771 | 1.16 | 9.006 | Confirmed | 119 | 91 | 76.37 | 76.19 | Confirmed | 134 | 23.56 | 24.27 | Confirmed
[7 empty] | 100.72 | 10 | TRUE | TRUE | Pentachloroethane | 10.39 | 6333 | 0.06 | 10.072 | Confirmed | 167 | 130 | 145.01 | 127.22 | Confirmed | 132 | 144.82 | 124.22 | Confirmed
[7 empty] | 93.88 | 10 | TRUE | TRUE | 1,2,4-Trimethylbenzene | 10.42 | 132641 | 1.28 | 9.388 | Confirmed | 105 | 120 | 41.67 | 44.46 | Confirmed | 77 | 12.21 | 12.34 | Confirmed
[7 empty] | 94.03 | 10 | TRUE | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 175535 | 1.69 | 9.403 | Confirmed | 105 | 134 | 18.5 | 19.09 | Confirmed | 91 | 15.72 | 15.64 | Confirmed
[7 empty] | 90.47 | 10 | TRUE | TRUE | 1,3-Dichlorobenzene | 10.61 | 65271 | 0.63 | 9.047 | Confirmed | 146 | 148 | 62.47 | 63.18 | Confirmed | 111 | 50.97 | 48.94 | Confirmed
[7 empty] | 98.45 | 10 | TRUE | TRUE | 4-Isopropyltoluene (p-cymene) | 10.64 | 134448 | 1.3 | 9.845 | Confirmed | 119 | 91 | 31.48 | 31.14 | Confirmed | 134 | 27.22 | 27.24 | Confirmed
[7 empty] | 100 | 20 | TRUE | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 81811 | 0.79 | 20 | Confirmed | 152 | 115 | 75.75 | 59.16 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 98.41 | 10 | TRUE | TRUE | 1,4-Dichlorobenzene | 10.68 | 63968 | 0.62 | 9.841 | Confirmed | 146 | 148 | 63.45 | 65.72 | Confirmed | 111 | 55.36 | 50.27 | Confirmed
[7 empty] | 89.2 | 10 | TRUE | TRUE | n-Butylbenzene | 10.91 | 134075 | 1.29 | 8.92 | Confirmed | 91 | 92 | 52.22 | 51.86 | Confirmed | 134 | 23.38 | 23.82 | Confirmed
[7 empty] | 99.59 | 10 | TRUE | TRUE | 1,2-Dichlorobenzene | 10.92 | 67086 | 0.65 | 9.959 | Confirmed | 146 | 148 | 62.83 | 63.53 | Confirmed | 111 | 51.55 | 49.43 | Confirmed
[7 empty] | 87.01 | 10 | TRUE | TRUE | Hexachloroethane | 11.11 | 9992 | 0.1 | 8.701 | Confirmed | 117 | 119 | 96.74 | 99.79 | Confirmed | 201 | 58.13 | 64.77 | Confirmed
[7 empty] | 100.13 | 10 | TRUE | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 3970 | 0.04 | 10.013 | Confirmed | 157 | 155 | 76.5 | 76.62 | Confirmed | 75 | 130.93 | 112.92 | Confirmed
[7 empty] | 85.68 | 10 | TRUE | TRUE | Nitrobenzene | 11.58 | 687 | 0.01 | 8.568 | Confirmed | 77 | 51 | 49.24 | 47.8 | Confirmed | 123 | 41.22 | 39.78 | Confirmed
[7 empty] | 82.05 | 10 | TRUE | TRUE | 1,2,4-Trichlorobenzene | 11.98 | 34974 | 0.34 | 8.205 | Confirmed | 180 | 182 | 94.66 | 95.98 | Confirmed | 145 | 41.13 | 39.75 | Confirmed
[7 empty] | 85.92 | 10 | TRUE | TRUE | Hexachloro-1,3-butadiene | 12.07 | 14102 | 0.14 | 8.592 | Confirmed | 225 | 227 | 64.13 | 64.15 | Confirmed | 223 | 63.42 | 61.68 | Confirmed
[7 empty] | 83.3 | 10 | TRUE | TRUE | Naphthalene | 12.15 | 121324 | 1.17 | 8.33 | Confirmed | 128 | 127 | 12.85 | 13.03 | Confirmed | 129 | 10.36 | 10.41 | Confirmed
[7 empty] | 87.41 | 10 | TRUE | TRUE | 1,2,3-Trichlorobenzene | 12.29 | 35791 | 0.35 | 8.741 | Confirmed | 180 | 182 | 94.11 | 93.54 | Confirmed | 145 | 43.31 | 40.39 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2024
240201
Peak Name | Amount
ppb
First Injection | CAL1 | CAL2 | CAL3 | CAL4 | CAL5 | CAL6 | H2O-V2 | H2O-V3 | 01.26.KDP.01 | 01.26.KDP.02 | 01.26.KDP.03 | 01.26.KDP.04 | 01.26.KDP.05 | 01.26.KDP.06 | CCV
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | 1.09 | 1.777 | 4.8927 | 10.0865 | 25.0063 | 99.9982 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7797
Chloroethene (vinyl chloride) | 1.0972 | 1.9047 | 5.3569 | 9.6189 | 25.1031 | 99.9946 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2714
Bromomethane (methyl bromide) | 1.4778 | 2.5011 | 5.8977 | 9.8118 | 24.7156 | 99.789 | 0.5556 | 0.4038 | 0.2629 | 0.1895 | n.a./n.r. | 0.1759 | 0.1535 | 0.13 | 12.0184
Chloroethane (ethyl chloride) | 1.25 | 2.2373 | 6.1491 | 10.8411 | 29.3479 | 98.7642 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4128
Trichlorofluoromethane | 1.1952 | 2.1693 | 6.311 | 11.1396 | 29.5625 | 98.6745 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6629
Diethyl ether | 0.9574 | 1.8905 | 5.1593 | 10.6321 | 28.6884 | 99.0093 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8714
1,1-Dichloroethene | 1.2105 | 2.1569 | 5.9988 | 10.7714 | 28.8428 | 98.907 | 0.0512 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0837
Acetone | 2.7536 | 3.5933 | 9.0214 | 20.7331 | 49.8983 | 200.0018 | n.a./n.r. | n.a./n.r. | 11.5521 | 13.5666 | 15.4965 | 12.2748 | 10.3418 | 10.6894 | 15.4811
Iodomethane | 1 | 2 | 5 | 10 | 25 | 100 | 0.6965 | 0.372 | 0.178 | 0.1997 | 0.1545 | 0.0856 | n.a./n.r. | 0.1266 | 6.3515
Carbon disulfide | 1.1983 | 2.1886 | 5.7351 | 10.77 | 28.5334 | 98.9971 | 0.365 | 0.1225 | 0.0787 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8764
3-Chloropropene (allyl chloride) | 0.9277 | 1.7027 | 4.7992 | 9.1602 | 25.472 | 99.982 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.513
Methylene chloride (DCM) | 1.1393 | 1.8714 | 4.9915 | 9.29 | 25.38 | 99.9723 | 0.0949 | 0.0699 | n.a./n.r. | 0.0573 | 0.0592 | n.a./n.r. | 0.0478 | 0.2542 | 7.6935
trans-1,2-Dichloroethene | 1.163 | 2.1406 | 5.8077 | 10.3944 | 28.3523 | 99.0776 | 0.105 | 0.0381 | 0.0227 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7947
Methyl tert-butyl ether (MTBE) | 0.9304 | 1.6897 | 4.9124 | 10.1352 | 26.9218 | 99.5173 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2133
1,1-Dichloroethane | 1.0846 | 1.9259 | 5.2279 | 10.0366 | 26.8947 | 99.5119 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0782
cis-1,2-Dichloroethene | 1.0979 | 2.0123 | 5.466 | 10.215 | 27.382 | 99.3585 | 0.0648 | 0.0208 | 0.0086 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6962
2,2-Dichloropropane | 1.4758 | 2.1644 | 4.9988 | 8.8976 | 25.4808 | 99.9819 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8888
2-Butanone (MEK) | 1.078 | 2.6145 | 8.7686 | 19.1099 | 50.2353 | 200.1287 | n.a./n.r. | n.a./n.r. | 0.5078 | 0.8744 | 0.2824 | 0.4934 | n.a./n.r. | n.a./n.r. | 16.4997
Methyl acrylate | 1.7347 | 2.2663 | 4.4548 | 9.0756 | 25.4804 | 99.9846 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0817
Bromochloromethane | 1.0441 | 1.8956 | 5.2261 | 10.2989 | 27.6517 | 99.2975 | 0.0566 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0114 | 0.0061 | n.a./n.r. | n.a./n.r. | 8.84
Methacrylonitrile | n.a./n.r. | 1.5089 | 4.6944 | 9.7829 | 26.6326 | 99.6432 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8357
Tetrahydrofuran | 0.6933 | 1.4517 | 4.3273 | 10.2069 | 26.7251 | 99.5957 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 3.124 | 7.8527
Trichloromethane (chloroform) | 1.1006 | 2.0489 | 5.6112 | 10.6313 | 28.7874 | 98.9575 | n.a./n.r. | n.a./n.r. | 2.5423 | 2.3864 | 32.0745 | 32.8186 | 2.2812 | n.a./n.r. | 9.4757
1,1,1-Trichloroethane | 0.7901 | 1.5102 | 4.9441 | 8.8966 | 25.2422 | 100.0645 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8344
Dibromofluoromethane [SS1] | 19.3319 | 19.5744 | 19.1811 | 20.6986 | 20.8382 | 20.3759 | 21.1351 | 21.0882 | 20.8314 | 20.5853 | 20.2263 | 19.7468 | 19.521 | 20.0313 | 19.7133
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | 0.8339 | 1.5062 | 4.7513 | 8.3487 | 24.3231 | 100.3583 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1431
Carbon tetrachloride | 1.3176 | 2.0925 | 5.1081 | 9.1523 | 25.3395 | 99.9888 | n.a./n.r. | 0.6032 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4013
1,1-Dichloropropene | 0.9319 | 1.7649 | 5.3331 | 9.4147 | 26.8008 | 99.5971 | n.a./n.r. | 0.0258 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6703
Benzene | 0.9828 | 1.9138 | 5.5042 | 9.938 | 27.7419 | 99.2974 | 0.0364 | n.a./n.r. | 0.0295 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0228 | 9.1973
1,2-Dichloroethane | 0.9435 | 1.6737 | 5.0788 | 9.3526 | 25.3527 | 99.9766 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7874
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | 0.9968 | 1.9406 | 5.5997 | 9.6927 | 26.273 | 99.6837 | 0.078 | n.a./n.r. | 0.017 | 0.0109 | n.a./n.r. | 0.0144 | n.a./n.r. | n.a./n.r. | 9.3088
1,2-Dichloropropane | 0.8148 | 1.553 | 4.718 | 8.8684 | 25.3583 | 100.0485 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6844
Dibromomethane | 0.8569 | 1.5989 | 5.0465 | 8.9605 | 24.7221 | 100.1806 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8761
Methyl methacrylate | 1.6063 | 2.2205 | 4.875 | 8.8398 | 25.4668 | 99.9883 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.077
Bromodichloromethane | 0.7574 | 1.4122 | 4.5838 | 8.3891 | 24.293 | 100.3728 | 0.0193 | n.a./n.r. | 0.02 | n.a./n.r. | 2.1875 | 0.8416 | n.a./n.r. | n.a./n.r. | 7.9674
2-Nitropropane | n.a./n.r. | n.a./n.r. | 4.6447 | 10.0761 | 25.1542 | 99.9931 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3126
cis-1,3-Dichloropropene | 0.7712 | 1.4001 | 4.6019 | 8.769 | 25.673 | 99.9803 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1489
4-Methyl-2-pentanone (MIBK) | 1.3748 | 2.6738 | 8.5469 | 17.4684 | 51.5785 | 199.9494 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 17.5729
Toluene-d8 [SS2] | 19.2488 | 19.2131 | 20.0488 | 20.9532 | 20.5483 | 19.9879 | 20.3187 | 19.9491 | 19.9204 | 19.7865 | 19.5301 | 19.9088 | 19.5704 | 19.2124 | 19.9055
Toluene | 0.953 | 1.7491 | 5.2413 | 9.2707 | 26.9511 | 99.5786 | 0.0716 | 0.0337 | 0.088 | 0.0899 | 0.043 | 0.0468 | 0.102 | 0.4524 | 8.6963
trans-1,3-Dichloropropene | 1.5962 | 2.2797 | 4.9733 | 8.617 | 25.542 | 99.9859 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3015
Ethyl methacrylate | 1.8664 | 2.3279 | 4.6568 | 8.5987 | 25.5601 | 99.9875 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4385
1,1,2-Trichloroethane | 0.7764 | 1.4581 | 4.3481 | 8.2372 | 23.9463 | 100.4854 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0719
Tetrachloroethene | 1.1156 | 2.0687 | 5.9257 | 9.7463 | 25.5769 | 99.8323 | 0.1124 | 0.0381 | 0.0204 | 0.0132 | 0.0116 | 0.0084 | n.a./n.r. | n.a./n.r. | 9.7845
1,3-Dichloropropane | 0.7524 | 1.4029 | 4.4087 | 8.5132 | 24.9723 | 100.1996 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4531
2-Hexanone | 2.9927 | 4.1878 | 9.8371 | 18.2724 | 50.7266 | 199.9796 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 18.7358
Dibromochloromethane | 0.8143 | 1.497 | 4.5185 | 8.4976 | 25.8085 | 99.9755 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.2284 | 0.0617 | n.a./n.r. | n.a./n.r. | 9.2759
1,2-Dibromoethane (EDB) | 0.8045 | 1.4869 | 4.7842 | 8.8627 | 25.6208 | 99.9787 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.187
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | 0.9999 | 1.8044 | 5.2499 | 9.2138 | 26.5642 | 99.679 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9817
1,1,1,2-Tetrachloroethane | 0.7965 | 1.512 | 4.4954 | 8.2454 | 25.9452 | 99.9694 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9923
Ethylbenzene | 1.092 | 2.0414 | 5.8943 | 10.0928 | 29.0169 | 98.94 | 0.0878 | 0.0328 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0822 | 9.5802
m/p-Xylene | 2.3665 | 4.1325 | 12.8023 | 21.9032 | 61.7504 | 196.7257 | 0.2162 | 0.0884 | 0.0931 | 0.0819 | 0.0613 | 0.045 | 0.0664 | 0.0593 | 21.1858
o-Xylene | 1.0786 | 1.932 | 5.9313 | 10.1595 | 28.0385 | 99.1784 | 0.0968 | 0.0383 | 0.0276 | 0.0303 | 0.0187 | 0.0115 | 0.0212 | n.a./n.r. | 9.6583
Styrene | 1.0826 | 2.0345 | 6.2215 | 10.7436 | 30.365 | 98.5218 | 0.1001 | 0.0402 | 0.0271 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.4435 | 10.4644
Bromoform | 0.7542 | 1.5522 | 4.1711 | 7.8705 | 26.0966 | 99.9704 | 0.0723 | 0.0189 | n.a./n.r. | n.a./n.r. | 0.0114 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.597
Isopropylbenzene (cumene) | 1.0498 | 1.9075 | 5.0931 | 9.4063 | 28.2333 | 99.2478 | 0.0978 | n.a./n.r. | 0.0205 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0913
1-Bromo-4-fluorobenzene (BFB) [SS3] | 21.06 | 20.3055 | 18.5432 | 20.2308 | 20.434 | 19.4265 | 21.3572 | 20.5301 | 19.5126 | 19.1524 | 18.9408 | 19.2035 | 19.5924 | 18.9526 | 18.5601
Bromobenzene | 1.0277 | 1.8389 | 4.6613 | 9.0702 | 26.6527 | 99.6997 | 0.1798 | 0.0504 | 0.0317 | 0.0231 | 0.0126 | 0.0171 | n.a./n.r. | n.a./n.r. | 8.3998
1,1,2,2-Tetrachloroethane | 1.7997 | 2.3674 | 4.699 | 8.4069 | 25.7565 | 99.9698 | 1.1906 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5418
1,2,3-Trichloropropane (TCP) | 1.3865 | 1.9717 | 4.4186 | 8.6961 | 26.8278 | 99.6992 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8063
trans-1,4-Dichloro-2-butene | 1.5507 | 2.1527 | 4.6291 | 8.7841 | 26.0024 | 99.881 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7917
n-Propylbenzene | 1.0899 | 1.9556 | 5.2915 | 10.2193 | 29.3603 | 98.8734 | 0.1473 | 0.0511 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4835
2-Chlorotoluene | 1.1367 | 2.0224 | 5.2655 | 10.0036 | 29.3666 | 98.8929 | 0.1536 | 0.0535 | 0.025 | 0.0177 | 0.0137 | 0.0085 | n.a./n.r. | n.a./n.r. | 9.5112
1,3,5-Trimethylbenzene | 0.8838 | 1.7311 | 4.7561 | 9.3689 | 28.4167 | 99.2277 | 0.1076 | 0.0309 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2058
4-Chlorotoluene | 1.1376 | 2.0184 | 5.436 | 10.363 | 31.2349 | 98.3814 | 0.1965 | 0.0646 | 0.0363 | 0.0254 | 0.0175 | 0.0152 | n.a./n.r. | 0.0078 | 10.1244
tert-Butylbenzene | 0.9072 | 1.7381 | 4.6962 | 9.0267 | 26.5066 | 99.742 | 0.1286 | 0.0274 | 0.0189 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0064
Pentachloroethane | 2.026 | n.a./n.r. | 4.586 | 8.3446 | 25.6711 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0717
1,2,4-Trimethylbenzene | 0.9133 | 1.795 | 4.8329 | 9.7897 | 28.7726 | 99.0912 | 0.1359 | n.a./n.r. | 0.0446 | n.a./n.r. | 0.0256 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3876
1-Methylpropylbenzene (sec-butylbenzene) | 0.9812 | 1.8595 | 5.1512 | 9.9055 | 28.8379 | 99.0454 | 0.1448 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.403
1,3-Dichlorobenzene | 0.9856 | 1.7979 | 4.6393 | 8.9065 | 25.7522 | 99.9435 | 0.2623 | 0.0817 | 0.0422 | 0.0277 | 0.0228 | 0.0217 | 0.0165 | 0.0124 | 9.0469
4-Isopropyltoluene (p-cymene) | 0.9873 | 1.9464 | 5.1976 | 10.2782 | 29.905 | 98.7372 | 0.151 | 0.0442 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1456 | 9.8455
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | 1.1056 | 2.0068 | 5.2658 | 10.0131 | 28.3836 | 99.1383 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8413
n-Butylbenzene | 0.9477 | 1.7742 | 5.0798 | 9.8905 | 29.4819 | 98.8915 | 0.2254 | 0.0691 | 0.041 | 0.0316 | 0.0205 | 0.0171 | 0.0138 | 0.0117 | 8.9204
1,2-Dichlorobenzene | 1.0132 | 1.916 | 5.0517 | 9.6224 | 28.6898 | 99.1143 | 0.2758 | 0.0777 | 0.0416 | n.a./n.r. | 0.0202 | 0.0198 | 0.0154 | 0.0162 | 9.9595
Hexachloroethane | 0.8399 | 1.5419 | 4.286 | 8.4836 | 25.8627 | 99.9735 | 0.1048 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7005
1,2-Dibromo-3-chloropropane (DBCP) | 1.7417 | 2.3593 | 4.6949 | 8.6328 | 25.5797 | 99.9848 | 1.4054 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0132
Nitrobenzene | n.a./n.r. | 2.2862 | 4.4202 | 8.7467 | 25.624 | 99.9801 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5677
1,2,4-Trichlorobenzene | 0.7733 | 1.5076 | 4.2242 | 8.0622 | 24.0736 | 100.4763 | 0.5493 | 0.1483 | 0.0841 | 0.0537 | 0.0399 | 0.0377 | 0.0258 | 0.0232 | 8.2053
Hexachloro-1,3-butadiene | 0.8083 | 1.4845 | 4.2599 | 7.8118 | 23.2068 | 100.7164 | 0.3133 | 0.0755 | 0.0465 | n.a./n.r. | n.a./n.r. | 0.0183 | 0.0154 | 0.0091 | 8.5923
Naphthalene | 0.7893 | 1.4692 | 4.2602 | 8.1867 | 25.6774 | 100.0617 | 0.7388 | 0.149 | 0.0822 | 0.0612 | 0.0352 | 0.0313 | 0.0306 | 0.0179 | 8.3298
1,2,3-Trichlorobenzene | 0.8213 | 1.5449 | 4.1905 | 8.0172 | 24.69 | 100.3271 | 0.7205 | 0.1452 | 0.0884 | 0.0604 | 0.0431 | 0.0345 | 0.0289 | 0.018 | 8.741

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library | Area | Estimated | 20 ppb (TIC) | RT | Area
| No. | min | SI |  |  | SI |  |  | SI |  |  | counts*min | Concentration | Pentafluorobenzene [IS1] | 5.43 | 254740
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC |  | (ppb) | 1,4-Difluorobenzene [IS2] | 6.18 | 373775
| 1 | 1.43 | 950 | Dimethyl ether | mainlib | 893 | Ethanol | mainlib | 847 | Hydrazine, methyl- | mainlib | 816734.829 | 64.123 | Chlorobenzene-d5 [IS3] | 7.61 | 560429
| 7 | 3.42 | 955 | Cyclopentane | mainlib | 929 | 1-Pentene | mainlib | 887 | Cyclopropane, ethyl- | mainlib | 5901252.863 | 463.3158 | 1,4-Dichlorobenzene-d4 [IS4] | 8.91 | 445540