## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2023/231025.seq/261.smp/MSDevice.xraw
Channel | TIC
Injection Information: |  |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Data Vault | ChromeleonLocal |  |  | 95 | 5939169 | 50-200% of mass 174 | 95/174 | 130.2231 | 50 | 200 | TRUE
Injection | H2O-R2 |  |  | 96 | 376614.2812 | 5 to 9% of m/z 95 | 96/95 | 6.3412 | 5 | 9 | TRUE
Injection Number | 3 |  |  | 173 | 9449.5205 | <2% of m/z 174 | 173/174 | 0.20719 | 0 | 2 | TRUE
Position | 0 |  |  | 174 | 4560764 | 50-200% of mass 95 | 174/95 | 76.7913 | 50 | 200 | TRUE
Comment |  |  |  | 175 | 313309.5938 | 5 to 9% of m/z 174 | 175/174 | 6.8697 | 5 | 9 | TRUE
Processing Method | VOC-9 |  |  | 176 | 4390088 | 95 to 105% of m/z 174 | 176/174 | 96.2577 | 95 | 105 | TRUE
Instrument Method | VOC-25 |  |  | 177 | 276595.125 | 5 to 10% of m/z 176 | 177/176 | 6.3004 | 5 | 10 | TRUE
Type | Unknown
Status | Finished
Injection Date | 2023-10-25
Injection Time | 17:28
Injection Volume (µl) | 20 | *N/A for P&T
Dilution Factor | 1
Weight | 1
Raw Data Information:
Data Points | 228
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9808
Retention Time (min) | 9.8361
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3286
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0229
Base Peak Intensity | 5939169
Total Ion Count | 24000093.9424
Raw Data: |  | Rounded:
Mass (m/z) | Value (counts) | Mass (m/z) | Value (counts)
35.3 | 332.6864 | 35 | 332.6864
36 | 24757.084 | 36 | 24757.084
37.1 | 155860.125 | 37 | 155860.125
38.1 | 128094.0156 | 38 | 128094.0156
39.1 | 56459.5547 | 39 | 56459.5547
40 | 36293.7656 | 40 | 36293.7656
41.1 | 2080.5518 | 41 | 2080.5518
42.1 | 1317.1638 | 42 | 1317.1638
43.1 | 1842.5793 | 43 | 1842.5793
44 | 60348.4883 | 44 | 60348.4883
45 | 33843.9102 | 45 | 33843.9102
45.7 | 309.1955 | 46 | 309.1955
46.2 | 2399.855 | 46 | 2399.855
47.1 | 71415.4844 | 47 | 71415.4844
48 | 22480.5684 | 48 | 22480.5684
49 | 143119.3281 | 49 | 143119.3281
50 | 734700.625 | 50 | 734700.625
51 | 212590.7188 | 51 | 212590.7188
52 | 8049.0166 | 52 | 8049.0166
53.2 | 706.9816 | 53 | 706.9816
55 | 6418.0186 | 55 | 6418.0186
56 | 40997.9922 | 56 | 40997.9922
57 | 86098.9062 | 57 | 86098.9062
58 | 2241.5808 | 58 | 2241.5808
59.1 | 439.9609 | 59 | 439.9609
60 | 27069.2852 | 60 | 27069.2852
61 | 149058.4219 | 61 | 149058.4219
62 | 156038.7656 | 62 | 156038.7656
63 | 109150.7812 | 63 | 109150.7812
64 | 11081.8164 | 64 | 11081.8164
65 | 2286.6321 | 65 | 2286.6321
66 | 667.4095 | 66 | 667.4095
67 | 8873.1709 | 67 | 8873.1709
68 | 419395.0312 | 68 | 419395.0312
69 | 375235.3438 | 69 | 375235.3438
70 | 26515.3477 | 70 | 26515.3477
71 | 1070.4137 | 71 | 1070.4137
72 | 16468.9277 | 72 | 16468.9277
73 | 166751.9531 | 73 | 166751.9531
74 | 617786.1875 | 74 | 617786.1875
75 | 2365574 | 75 | 2365574
76.1 | 192127.0312 | 76 | 192127.0312
77.1 | 29431.0137 | 77 | 29431.0137
78 | 23290.7422 | 78 | 23290.7422
78.9 | 36474.9648 | 79 | 36474.9648
80 | 15858.1465 | 80 | 15858.1465
80.9 | 38276.4883 | 81 | 38276.4883
82 | 5668.6704 | 82 | 5668.6704
83 | 1065.3978 | 83 | 1065.3978
83.6 | 207.848 | 84 | 207.848
84.3 | 0.019524 | 84 | 0.019524
84.8 | 4.8134 | 85 | 4.8134
86 | 5656.9575 | 86 | 5656.9575
87 | 237724.7188 | 87 | 237724.7188
88 | 236753.5781 | 88 | 236753.5781
89.2 | 164.4024 | 89 | 164.4024
90.9 | 10764.2109 | 91 | 10764.2109
92 | 100426.6953 | 92 | 100426.6953
93 | 170148.7969 | 93 | 170148.7969
94 | 468646.5938 | 94 | 468646.5938
95 | 5939169 | 95 | 5939169
96 | 376614.2812 | 96 | 376614.2812
97.1 | 9044.3047 | 97 | 9044.3047
98.1 | 266.9487 | 98 | 266.9487
99.3 | 302.7762 | 99 | 302.7762
100.3 | 471.1996 | 100 | 471.1996
101.1 | 0 | 101 | 0
101.8 | 2.7327 | 102 | 2.7327
103.1 | 1156.1034 | 103 | 1156.1034
103.9 | 9706.7451 | 104 | 9706.7451
104.9 | 3883.6624 | 105 | 3883.6624
105.9 | 11241.4932 | 106 | 11241.4932
107 | 1676.6174 | 107 | 1676.6174
108 | 270.5235 | 108 | 270.5235
108.9 | 118.8784 | 109 | 118.8784
109.9 | 732.5402 | 110 | 732.5402
110.9 | 1535.5646 | 111 | 1535.5646
111.9 | 516.0637 | 112 | 516.0637
112.6 | 441.4137 | 113 | 441.4137
113.2 | 205.5961 | 113 | 205.5961
113.9 | 60.0355 | 114 | 60.0355
115 | 2624.4536 | 115 | 2624.4536
115.9 | 9747.4404 | 116 | 9747.4404
117 | 19052.4844 | 117 | 19052.4844
117.9 | 7253.2134 | 118 | 7253.2134
118.9 | 14123.1738 | 119 | 14123.1738
120.1 | 540.6302 | 120 | 540.6302
121 | 130.5527 | 121 | 130.5527
121.9 | 307.5013 | 122 | 307.5013
122.7 | 343.631 | 123 | 343.631
123.9 | 1940.2236 | 124 | 1940.2236
124.8 | 720.3721 | 125 | 720.3721
126.1 | 569.9866 | 126 | 569.9866
127 | 471.9933 | 127 | 471.9933
127.9 | 11861.4697 | 128 | 11861.4697
129 | 4832.418 | 129 | 4832.418
129.9 | 13606.4365 | 130 | 13606.4365
130.9 | 4233.9961 | 131 | 4233.9961
132 | 556.5607 | 132 | 556.5607
133.1 | 287.5861 | 133 | 287.5861
134 | 902.0736 | 134 | 902.0736
134.9 | 4377.332 | 135 | 4377.332
135.9 | 740.1766 | 136 | 740.1766
137 | 4599.3047 | 137 | 4599.3047
138.5 | 276.7761 | 139 | 276.7761
139 | 48.6136 | 139 | 48.6136
139.9 | 1244.3536 | 140 | 1244.3536
140.9 | 27216.0977 | 141 | 27216.0977
141.8 | 2154.2202 | 142 | 2154.2202
142.9 | 24953.9941 | 143 | 24953.9941
143.9 | 1835.6947 | 144 | 1835.6947
144.9 | 1736.4485 | 145 | 1736.4485
145.9 | 4947.8843 | 146 | 4947.8843
146.9 | 2143.6719 | 147 | 2143.6719
147.9 | 9556.9414 | 148 | 9556.9414
149 | 2280.1138 | 149 | 2280.1138
150 | 2095.3384 | 150 | 2095.3384
151.1 | 243.4434 | 151 | 243.4434
151.9 | 2011.7875 | 152 | 2011.7875
153 | 3077.3926 | 153 | 3077.3926
154 | 2822.9719 | 154 | 2822.9719
154.9 | 8440.1377 | 155 | 8440.1377
155.9 | 1611.5864 | 156 | 1611.5864
156.9 | 4814.5391 | 157 | 4814.5391
158 | 1107.9017 | 158 | 1107.9017
158.9 | 2845.0454 | 159 | 2845.0454
160.1 | 398.846 | 160 | 398.846
160.9 | 3311.27 | 161 | 3311.27
161.9 | 400.5527 | 162 | 400.5527
163 | 126.6625 | 163 | 126.6625
163.7 | 0 | 164 | 0
165 | 3.1488 | 165 | 3.1488
165.6 | 50.1018 | 166 | 50.1018
166.5 | 105.0583 | 167 | 105.0583
167.1 | 0 | 167 | 0
167.9 | 318.9973 | 168 | 318.9973
168.7 | 124.7018 | 169 | 124.7018
169.4 | 217.4709 | 169 | 217.4709
170.1 | 369.6842 | 170 | 369.6842
171.1 | 278.1899 | 171 | 278.1899
172 | 5105.8384 | 172 | 5105.8384
172.7 | 9449.5205 | 173 | 9449.5205
173.9 | 4560764 | 174 | 4560764
175 | 313309.5938 | 175 | 313309.5938
175.9 | 4390088 | 176 | 4390088
177 | 276595.125 | 177 | 276595.125
178 | 7594.2773 | 178 | 7594.2773
179 | 4.5886 | 179 | 4.5886
180.3 | 0 | 180 | 0
181.3 | 0 | 181 | 0
181.9 | 0.006831 | 182 | 0.006831
182.5 | 0 | 183 | 0
183.1 | 0.012026 | 183 | 0.012026
183.9 | 0 | 184 | 0
184.4 | 0 | 184 | 0
185 | 0 | 185 | 0
185.9 | 0 | 186 | 0
186.8 | 0 | 187 | 0
187.9 | 0 | 188 | 0
188.7 | 4.4567 | 189 | 4.4567
189.3 | 181.3055 | 189 | 181.3055
190.2 | 266.0359 | 190 | 266.0359
190.8 | 375.0984 | 191 | 375.0984
191.4 | 75.3758 | 191 | 75.3758
192.1 | 0.009054 | 192 | 0.009054
193 | 182.1265 | 193 | 182.1265
193.6 | 0.003275 | 194 | 0.003275
194.3 | 60.6861 | 194 | 60.6861
195.3 | 0.019254 | 195 | 0.019254
196.6 | 0 | 197 | 0
197.6 | 0.000523 | 198 | 0.000523
199.5 | 0 | 200 | 0
200.9 | 0 | 201 | 0
202.9 | 0 | 203 | 0
203.9 | 0.17869 | 204 | 0.17869
204.5 | 0 | 205 | 0
205.2 | 0.028013 | 205 | 0.028013
205.8 | 0.12929 | 206 | 0.12929
206.8 | 1557.844 | 207 | 1557.844
208.1 | 585.7047 | 208 | 585.7047
209.1 | 189.6801 | 209 | 189.6801
209.8 | 3e-06 | 210 | 3e-06
210.4 | 0 | 210 | 0
211.1 | 0 | 211 | 0
215.3 | 0 | 215 | 0
216.5 | 0 | 217 | 0
217.1 | 0 | 217 | 0
218.9 | 4.5218 | 219 | 4.5218
219.5 | 0 | 220 | 0
220.1 | 0 | 220 | 0
221.2 | 0 | 221 | 0
222.1 | 1.8947 | 222 | 1.8947
222.6 | 0 | 223 | 0
223.2 | 0.00655 | 223 | 0.00655
224 | 0 | 224 | 0
224.6 | 0 | 225 | 0
226.2 | 0 | 226 | 0
227.5 | 0 | 228 | 0
229 | 0.11002 | 229 | 0.11002
231.2 | 133.3086 | 231 | 133.3086
232.2 | 0.34547 | 232 | 0.34547
233.2 | 7.2799 | 233 | 7.2799
233.8 | 0 | 234 | 0
234.3 | 0 | 234 | 0
235 | 0.038508 | 235 | 0.038508
235.5 | 0 | 236 | 0
237 | 0 | 237 | 0
238.6 | 0 | 239 | 0
240.6 | 1e-06 | 241 | 1e-06
241.3 | 0.024755 | 241 | 0.024755
241.9 | 0 | 242 | 0
244.3 | 0 | 244 | 0
244.9 | 0 | 245 | 0
245.6 | 0 | 246 | 0
246.9 | 0.012615 | 247 | 0.012615
247.4 | 0 | 247 | 0
248.9 | 0.011476 | 249 | 0.011476
249.4 | 0 | 249 | 0
250.3 | 106.447 | 250 | 106.447
251.6 | 0.066507 | 252 | 0.066507
252.7 | 0 | 253 | 0
253.5 | 0 | 254 | 0
254.2 | 0 | 254 | 0
255.3 | 0 | 255 | 0
256.6 | 0 | 257 | 0
258.1 | 0 | 258 | 0
259.1 | 0 | 259 | 0
260 | 128.3679 | 260 | 128.3679
250.3 | 4.1e-05 | 250 | 4.1e-05
251.2 | 315.1255 | 251 | 315.1255
252.1 | 0 | 252 | 0
253.5 | 3.3149 | 254 | 3.3149
254.5 | 0 | 255 | 0
255.1 | 0 | 255 | 0
256.1 | 0 | 256 | 0
256.7 | 0 | 257 | 0
257.7 | 0 | 258 | 0
258.8 | 2.7447 | 259 | 2.7447
259.3 | 0 | 259 | 0
260 | 121.3069
258.7 | 4e-06
259.9 | 229.5571
257.3 | 0.004411
258.3 | 181.4352
259 | 0.043369
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[8 empty] | TV | [7 empty] | Actual | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
Name | RT | Window | Quant | Conf1 | Conf2 | IS |  | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | MRL | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | RSD<20% | R^2>0.99
Chloromethane (methyl chloride) | 1.464 | 0.050 AN | 50 | 52 | 49 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Chloroethene (vinyl chloride) | 1.553 | 0.050 AN | 62 | 64 | 61 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Trichlorofluoromethane | 2.182 | 0.050 AN | 101 | 103 | 105 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Diethyl ether | 2.505 | 0.080 AN | 59 | 74 | 45 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Acetone | 2.852 | 0.088 AN | 43 | 58 |  | Internal Pentafluorobenzene [IS1] | P-P |  | 1.8 |  | 9 | 36 | 180 | ppb | 1.8 | [18 empty] | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Carbon disulfide | 2.961 | 0.050 AN | 76 | 78 |  | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
trans-1,2-Dichloroethene | 3.678 | 0.100 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
2,2-Dichloropropane | 4.819 | 0.060 AN | 77 | 41 | 79 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
cis-1,2-Dichloroethene | 4.823 | 0.060 AN | 61 | 96 | 98 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
2-Butanone (MEK) | 4.851 | 0.060 AN | 43 | 72 | 57 | Internal Pentafluorobenzene [IS1] | P-P |  | 1.8 |  | 9 | 36 | 180 | ppb | 1.8 | [18 empty] | TRUE | TRUE
Methyl acrylate | 4.939 | 0.060 AN | 55 | 85 |  | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Bromochloromethane | 5.073 | 0.060 AN | 49 | 130 | 128 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Methacrylonitrile | 5.076 | 0.060 AN | 67 | 52 | 40 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Tetrahydrofuran | 5.09 | 0.060 AN | 42 | 72 | 71 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Trichloromethane (chloroform) | 5.198 | 0.060 AN | 83 | 85 | 47 | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | Internal Pentafluorobenzene [IS1] | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | ISTD Internal | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.476 | 0.040 AN | 56 | 41 | 43 | Internal 1,4-Difluorobenzene [IS2] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Carbon tetrachloride | 5.493 | 0.040 AN | 119 | 121 |  | Internal Pentafluorobenzene [IS1] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | Internal 1,4-Difluorobenzene [IS2] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.167 | 0.080 AN | 114 | 88 | 63 | ISTD Internal | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Dibromomethane | 6.715 | 0.060 AN | 174 | 93 | 95 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.518 | 0.049 AN | 43 | 58 | 41 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1.8 |  | 9 | 36 | 180 | ppb | 1.8 | [18 empty] | TRUE | TRUE
Toluene-d8 [SS2] | 7.607 | 0.058 AN | 98 | 100 | 70 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
Toluene | 7.67 | 0.047 AN | 91 | 92 | 65 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
trans-1,3-Dichloropropene | 7.92 | 0.060 AN | 75 | 39 | 77 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Ethyl methacrylate | 7.99 | 0.060 AN | 69 | 41 | 99 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1,2-Trichloroethane | 8.098 | 0.060 AN | 97 | 83 | 99 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,3-Dichloropropane | 8.24 | 0.060 AN | 76 | 41 | 78 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1.8 |  | 9 | 36 | 180 | ppb | 1.8 | [18 empty] | TRUE | TRUE
Dibromochloromethane | 8.418 | 0.050 AN | 129 | 127 | 131 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.517 | 0.049 AN | 107 | 109 | 93 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.915 | 0.040 AF | 117 | 82 | 52 | ISTD Internal | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
Chlorobenzene | 8.937 | 0.040 AN | 112 | 77 | 114 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.01 | 0.040 AN | 131 | 133 | 117 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Ethylbenzene | 9.018 | 0.040 AN | 91 | 106 | 51 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
m/p-Xylene | 9.125 | 0.060 AN | 91 | 106 | 105 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
o-Xylene | 9.423 | 0.060 AN | 91 | 106 | 105 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Styrene | 9.439 | 0.060 AN | 104 | 78 | 103 | Internal Chlorobenzene-d5 [IS3] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Isopropylbenzene (cumene) | 9.726 | 0.060 AN | 105 | 120 | 79 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.835 | 0.060 AN | 95 | 174 | 176 | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
Bromobenzene | 9.936 | 0.040 AN | 77 | 156 | 158 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.944 | 0.040 AN | 83 | 85 | 95 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.969 | 0.040 AN | 75 | 53 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.975 | 0.040 AN | 77 | 110 | 61 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
n-Propylbenzene | 10.016 | 0.026 AN | 91 | 120 | 65 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
2-Chlorotoluene | 10.071 | 0.031 AN | 91 | 126 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.154 | 0.040 AN | 105 | 120 | 119 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
4-Chlorotoluene | 10.168 | 0.040 AN | 91 | 126 | 89 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
tert-Butylbenzene | 10.357 | 0.040 AN | 119 | 91 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Pentachloroethane | 10.381 | 0.030 AN | 167 | 130 | 132 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.411 | 0.030 AN | 105 | 120 | 77 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.521 | 0.037 AN | 105 | 134 | 91 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,3-Dichlorobenzene | 10.603 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.656 | 0.036 AN | 152 | 150 |  | ISTD Internal | Lin, Avg |  | 20 |  | 20 | 20 | 20 | ppb | 20 | [18 empty] | TRUE | TRUE
1,4-Dichlorobenzene | 10.679 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
n-Butylbenzene | 10.9 | 0.040 AN | 91 | 92 | 134 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2-Dichlorobenzene | 10.914 | 0.040 AN | 146 | 148 | 111 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Hexachloroethane | 11.093 | 0.040 AN | 117 | 119 | 201 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.447 | 0.040 AN | 157 | 155 | 75 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Nitrobenzene | 11.566 | 0.040 AN | 77 | 51 | 123 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.966 | 0.040 AN | 180 | 182 | 145 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.066 | 0.040 AN | 225 | 227 | 223 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.273 | 0.040 AN | 180 | 182 | 145 | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P |  | 1 |  | 5 | 20 | 100 | ppb | 1 | [18 empty] | TRUE | TRUE

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 1.8
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 1.8
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 1
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | n.a.
Pentafluorobenzene [IS1] | n.a.
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | n.a.
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 1
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 1.8
Toluene-d8 [SS2] | n.a.
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 1.8
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | n.a.
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 1
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | n.a.
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | n.a.
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 1
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.45 | 212 | 0.02 | 0.109 | Not confirmed | 50 | 52 | 32.52 | 43.18 | Confirmed | 49 | 9.04 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | 1.55 | 53 | 0 | 0.022 | Not confirmed | 62 | 64 | 31.54 | n.a. | Not confirmed | 61 | 7.07 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.84 | 1245 | 0.09 | 0.271 | Confirmed | 94 | 96 | 90.83 | 97.14 | Confirmed | 93 | 19.04 | 17.62 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 31.29 | n.a. | Not confirmed | 49 | 22.1 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | 2.19 | 21 | 0 | 0.009 | Not confirmed | 101 | 103 | 65.64 | n.a. | Not confirmed | 105 | 10.65 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 74.44 | n.a. | Not confirmed | 45 | 70.22 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | 2.73 | 65 | 0 | 0.021 | Not confirmed | 61 | 96 | 70.15 | n.a. | Not confirmed | 98 | 44.75 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 37.92 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Iodomethane | 2.89 | 733 | 0.05 | 0.387 | Confirmed | 142 | 127 | 30.19 | 23.48 | Confirmed | 141 | 12.62 | 8.1 | Confirmed
1 | TRUE | TRUE | FALSE | Carbon disulfide | 2.95 | 1096 | 0.08 | 0.141 | Not confirmed | 76 | 78 | 8.71 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 52.32 | n.a. | Not confirmed | 76 | 39.26 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Methylene chloride (DCM) | 3.35 | 926 | 0.07 | 0.192 | Confirmed | 49 | 84 | 87.46 | 84.75 | Confirmed | 86 | 54.29 | 53.43 | Confirmed
1 | TRUE | FALSE | FALSE | trans-1,2-Dichloroethene | 3.67 | 161 | 0.01 | 0.044 | Confirmed | 61 | 96 | 71.11 | 62.37 | Confirmed | 98 | 44.52 | 34.58 | Confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 25.35 | n.a. | Not confirmed | 57 | 22.76 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 65 | 31.28 | n.a. | Not confirmed | 83 | 11.68 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | cis-1,2-Dichloroethene | 4.81 | 102 | 0.01 | 0.022 | Confirmed | 61 | 96 | 76.84 | 61.73 | Confirmed | 98 | 48.66 | 65 | Confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 77.59 | n.a. | Not confirmed | 79 | 32.27 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 27.41 | n.a. | Not confirmed | 57 | 7.73 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 15.58 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Bromochloromethane | 5.06 | 40 | 0 | 0.012 | Confirmed | 49 | 130 | 82.24 | 106.19 | Confirmed | 128 | 63.47 | 87.49 | Confirmed
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 29.87 | n.a. | Not confirmed | 40 | 32.42 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 44.72 | n.a. | Not confirmed | 71 | 45.06 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | 5.19 | 50 | 0 | 0.01 | Not confirmed | 83 | 85 | 64.48 | n.a. | Not confirmed | 47 | 16.96 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 99 | 63.03 | n.a. | Not confirmed | 61 | 42.06 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.35 | 66980 | 4.83 | 20.327 | Confirmed | 113 | 111 | 102.66 | 104.97 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.42 | 132028 | 9.53 | 20 | Confirmed | 168 | 99 | 50.27 | 49.85 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | 5.47 | 89 | 0.01 | 0.023 | Not confirmed | 56 | 41 | 54.04 | n.a. | Not confirmed | 43 | 25.36 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon tetrachloride | n.a. | n.a. | n.a. | n.a. | Not confirmed | 119 | 121 | 32.02 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,1-Dichloropropene | 5.5 | 120 | 0.01 | 0.042 | Not confirmed | 75 | 77 | 30.9 | n.a. | Not confirmed | 110 | 40.86 | 21.32 | Confirmed
1 | TRUE | TRUE | FALSE | Benzene | 5.7 | 223 | 0.02 | 0.018 | Not confirmed | 78 | 77 | 23.83 | n.a. | Not confirmed | 52 | 14.26 | 19.01 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | 5.76 | 57 | 0 | 0.013 | Not confirmed | 62 | 64 | 32.7 | n.a. | Not confirmed | 49 | 28.82 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.16 | 217164 | 15.68 | 20 | Confirmed | 114 | 88 | 18.11 | 18.3 | Confirmed | 63 | 18.38 | 18.32 | Confirmed
1 | TRUE | FALSE | FALSE | Trichloroethene | 6.38 | 190 | 0.01 | 0.056 | Confirmed | 130 | 132 | 96.9 | 95.47 | Confirmed | 95 | 97.17 | 88.08 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 69.06 | n.a. | Not confirmed | 41 | 39.67 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.71 | 63 | 0 | 0.024 | Not confirmed | 174 | 93 | 93.18 | 106.26 | Confirmed | 95 | 79.32 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 94.07 | n.a. | Not confirmed | 39 | 42.42 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | 6.92 | 52 | 0 | 0.014 | Not confirmed | 83 | 85 | 62.94 | n.a. | Not confirmed | 47 | 14.8 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 80.97 | n.a. | Not confirmed | 39 | 25.57 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | 7.35 | 140 | 0.01 | 0.035 | Not confirmed | 75 | 39 | 38.19 | n.a. | Not confirmed | 77 | 31.65 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 41.82 | n.a. | Not confirmed | 41 | 22.91 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.6 | 288990 | 20.86 | 19.725 | Confirmed | 98 | 100 | 64.11 | 64.19 | Confirmed | 70 | 11.17 | 11.32 | Confirmed
1 | TRUE | FALSE | FALSE | Toluene | 7.67 | 1228 | 0.09 | 0.094 | Confirmed | 91 | 92 | 57.72 | 54.74 | Confirmed | 65 | 11.76 | 11.4 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,3-Dichloropropene | 7.92 | 213 | 0.02 | 0.072 | Not confirmed | 75 | 39 | 38.97 | 47.92 | Confirmed | 77 | 31.74 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 69 | 41 | 57.84 | n.a. | Not confirmed | 99 | 23.85 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 83 | 88.16 | n.a. | Not confirmed | 99 | 62.45 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.15 | 217 | 0.02 | 0.042 | Confirmed | 166 | 164 | 78.07 | 81.16 | Confirmed | 129 | 72.13 | 71.57 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3-Dichloropropane | 8.24 | 105 | 0.01 | 0.017 | Not confirmed | 76 | 41 | 58.21 | n.a. | Not confirmed | 78 | 32.03 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 58.21 | n.a. | Not confirmed | 57 | 20.79 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | 8.42 | 24 | 0 | 0.008 | Not confirmed | 129 | 127 | 77.67 | 107.62 | Confirmed | 131 | 23.52 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 95.55 | n.a. | Not confirmed | 93 | 4.12 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.91 | 214507 | 15.48 | 20 | Confirmed | 117 | 82 | 58.25 | 58.66 | Confirmed | 52 | 14.18 | 13.45 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.93 | 823 | 0.06 | 0.091 | Not confirmed | 112 | 77 | 65.05 | 202.28 | Not confirmed | 114 | 31.26 | 31.93 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 131 | 133 | 95.33 | n.a. | Not confirmed | 117 | 80.9 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Ethylbenzene | 9.02 | 1688 | 0.12 | 0.126 | Confirmed | 91 | 106 | 36.31 | 34.33 | Confirmed | 51 | 8.66 | 8.67 | Confirmed
1 | TRUE | FALSE | FALSE | m/p-Xylene | 9.12 | 3354 | 0.24 | 0.142 | Confirmed | 91 | 106 | 52.86 | 51.48 | Confirmed | 105 | 21.3 | 19.94 | Confirmed
1 | TRUE | FALSE | FALSE | o-Xylene | 9.42 | 1216 | 0.09 | 0.095 | Confirmed | 91 | 106 | 50.73 | 49.56 | Confirmed | 105 | 24.79 | 30.86 | Confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.44 | 1404 | 0.1 | 0.135 | Confirmed | 104 | 78 | 54.72 | 50.95 | Confirmed | 103 | 53.56 | 50.46 | Confirmed
1 | TRUE | FALSE | FALSE | Bromoform | 9.57 | 55 | 0 | 0.026 | Confirmed | 173 | 171 | 50.58 | 48.54 | Confirmed | 175 | 48.87 | 69.44 | Confirmed
1 | TRUE | FALSE | FALSE | Isopropylbenzene (cumene) | 9.7 | 868 | 0.06 | 0.071 | Confirmed | 105 | 120 | 30.39 | 24.4 | Confirmed | 79 | 15.94 | 14.63 | Confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.83 | 118824 | 8.58 | 20.83 | Confirmed | 95 | 174 | 76.28 | 76.06 | Confirmed | 176 | 74.38 | 75.35 | Confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.93 | 734 | 0.05 | 0.108 | Confirmed | 77 | 156 | 61.67 | 65.91 | Confirmed | 158 | 59.93 | 63.19 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | 9.94 | 66 | 0 | 0.015 | Not confirmed | 83 | 85 | 63.62 | n.a. | Not confirmed | 95 | 13.5 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | trans-1,4-Dichloro-2-butene | 9.97 | 205 | 0.01 | 0.036 | Confirmed | 75 | 53 | 18.09 | 41.99 | Confirmed | 89 | 10.48 | 23.17 | Confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | 9.98 | 57 | 0 | 0.029 | Not confirmed | 77 | 110 | 77.48 | 40.98 | Not confirmed | 61 | 57.54 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | n-Propylbenzene | 10.01 | 2400 | 0.17 | 0.162 | Confirmed | 91 | 120 | 26.61 | 25.93 | Confirmed | 65 | 10.23 | 9.16 | Confirmed
1 | TRUE | FALSE | FALSE | 2-Chlorotoluene | 10.07 | 1097 | 0.08 | 0.109 | Confirmed | 91 | 126 | 36.74 | 29.37 | Confirmed | 89 | 17.3 | 12.75 | Confirmed
1 | TRUE | FALSE | FALSE | 1,3,5-Trimethylbenzene | 10.14 | 908 | 0.07 | 0.081 | Confirmed | 105 | 120 | 50.94 | 52.99 | Confirmed | 119 | 11.44 | 15.22 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Chlorotoluene | 10.16 | 2043 | 0.15 | 0.174 | Confirmed | 91 | 126 | 32.86 | 34.73 | Confirmed | 89 | 10.97 | 9.9 | Confirmed
1 | TRUE | FALSE | FALSE | tert-Butylbenzene | 10.36 | 508 | 0.04 | 0.052 | Confirmed | 119 | 91 | 65.97 | 65.39 | Confirmed | 134 | 24.1 | 23.38 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | 10.38 | 14 | 0 | 0.022 | Not confirmed | 167 | 130 | 53.27 | n.a. | Not confirmed | 132 | 56.54 | 116.37 | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trimethylbenzene | 10.41 | 1346 | 0.1 | 0.117 | Confirmed | 105 | 120 | 49.16 | 46.71 | Confirmed | 77 | 11.38 | 11.35 | Confirmed
1 | TRUE | FALSE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.52 | 1639 | 0.12 | 0.12 | Confirmed | 105 | 134 | 21.9 | 21.31 | Confirmed | 91 | 15.15 | 15.06 | Confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.6 | 1599 | 0.12 | 0.211 | Confirmed | 146 | 148 | 64.16 | 63.85 | Confirmed | 111 | 42.33 | 40.23 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.63 | 1635 | 0.12 | 0.145 | Confirmed | 119 | 91 | 27.95 | 25.33 | Confirmed | 134 | 30.53 | 31.66 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 124579 | 8.99 | 20 | Confirmed | 152 | 115 | 60.23 | 60.03 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.67 | 1857 | 0.13 | 0.227 | Not confirmed | 146 | 148 | 64.59 | 72.19 | Confirmed | 111 | 45.9 | 102.24 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.91 | 3211 | 0.23 | 0.304 | Confirmed | 91 | 92 | 52.39 | 51.55 | Confirmed | 134 | 29.3 | 28.18 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.91 | 1105 | 0.08 | 0.137 | Confirmed | 146 | 148 | 64.17 | 65.45 | Confirmed | 111 | 42.6 | 43.55 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 97.97 | n.a. | Not confirmed | 201 | 88.19 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 157 | 155 | 76.86 | n.a. | Not confirmed | 75 | 89.98 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 45.1 | n.a. | Not confirmed | 123 | 44.86 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.97 | 1557 | 0.11 | 0.298 | Confirmed | 180 | 182 | 95.1 | 92.58 | Confirmed | 145 | 32.08 | 30.48 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.06 | 481 | 0.03 | 0.224 | Confirmed | 225 | 227 | 62.87 | 70.78 | Confirmed | 223 | 62.66 | 62.08 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.14 | 2609 | 0.19 | 0.159 | Confirmed | 128 | 127 | 12.77 | 11.44 | Confirmed | 129 | 10.48 | 11.06 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.28 | 1232 | 0.09 | 0.224 | Confirmed | 180 | 182 | 95.85 | 97.86 | Confirmed | 145 | 33.64 | 32.2 | Confirmed

## CCV1
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? | <70% highlighted |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 134508 | 5.42 | 161264 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.16 | 225262 | 6.16 | 244452 | TRUE | TRUE | 84.53 | 10 | TRUE | Chloromethane (methyl chloride) | 1.46 | 19123 | 0.16 | 8.453 | Confirmed | 50 | 52 | 31.71 | 32.32 | Confirmed | 49 | 9.01 | 9.1 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 211361 | 8.91 | 230817 | TRUE | TRUE | 88.9 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.55 | 34740 | 0.29 | 8.89 | Confirmed | 62 | 64 | 31.1 | 31.44 | Confirmed | 61 | 7.37 | 7.09 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 130447 | 10.66 | 138670 | TRUE | TRUE | 109.24 | 10 | TRUE | Bromomethane (methyl bromide) | 1.83 | 46376 | 0.39 | 10.924 | Confirmed | 94 | 96 | 92.8 | 93.98 | Confirmed | 93 | 18.28 | 19.45 | Confirmed
[7 empty] | 85.27 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.94 | 26013 | 0.22 | 8.527 | Confirmed | 64 | 66 | 31.17 | 30.67 | Confirmed | 49 | 22.55 | 22.31 | Confirmed
[7 empty] | 83.05 | 10 | TRUE | Trichlorofluoromethane | 2.18 | 37608 | 0.32 | 8.305 | Confirmed | 101 | 103 | 63.97 | 66.76 | Confirmed | 105 | 9.43 | 10.02 | Confirmed
Total Analytes | 81 | [5 empty] | 88.19 | 10 | TRUE | Diethyl ether | 2.49 | 32678 | 0.28 | 8.819 | Confirmed | 59 | 74 | 73.64 | 76.79 | Confirmed | 45 | 71.41 | 70.84 | Confirmed
Failed | 3 | [5 empty] | 83.08 | 10 | TRUE | 1,1-Dichloroethene | 2.73 | 41113 | 0.35 | 8.308 | Confirmed | 61 | 96 | 76.66 | 72.05 | Confirmed | 98 | 49.44 | 46.89 | Confirmed
Allowance | 16.2 | [5 empty] | 82.4389 | 18 | TRUE | Acetone | 2.81 | 18241 | 0.15 | 14.839 | Confirmed | 43 | 58 | 33.08 | 41.76 | Confirmed | n.a. | n.a. | n.a. | n.a.
Pass? | TRUE | [5 empty] | 66.17 | 10 | FALSE | Iodomethane | 2.88 | 26683 | 0.23 | 6.617 | Confirmed | 142 | 127 | 29.47 | 30.71 | Confirmed | 141 | 12.1 | 12.52 | Confirmed
[7 empty] | 83.47 | 10 | TRUE | Carbon disulfide | 2.95 | 100995 | 0.85 | 8.347 | Confirmed | 76 | 78 | 8.53 | 8.65 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 82.97 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.19 | 45954 | 0.39 | 8.297 | Confirmed | 41 | 39 | 50.39 | 50.64 | Confirmed | 76 | 34.85 | 39.79 | Confirmed
[7 empty] | 85.11 | 10 | TRUE | Methylene chloride (DCM) | 3.35 | 46766 | 0.39 | 8.511 | Confirmed | 49 | 84 | 85.32 | 90.9 | Confirmed | 86 | 53.17 | 57.62 | Confirmed
[7 empty] | 84.52 | 10 | TRUE | trans-1,2-Dichloroethene | 3.67 | 44769 | 0.38 | 8.452 | Confirmed | 61 | 96 | 74.1 | 72.39 | Confirmed | 98 | 46.54 | 46.37 | Confirmed
[7 empty] | 89.53 | 10 | TRUE | Methyl tert-butyl ether (MTBE) | 3.68 | 90938 | 0.77 | 8.953 | Confirmed | 73 | 41 | 26.71 | 25.46 | Confirmed | 57 | 22.34 | 22.07 | Confirmed
[7 empty] | 87.07 | 10 | TRUE | 1,1-Dichloroethane | 4.18 | 60377 | 0.51 | 8.707 | Confirmed | 63 | 65 | 31.43 | 31.71 | Confirmed | 83 | 11.93 | 12.23 | Confirmed
[7 empty] | 76.61 | 10 | FALSE | 2,2-Dichloropropane | 4.81 | 25327 | 0.21 | 7.661 | Confirmed | 77 | 41 | 76.36 | 73.71 | Confirmed | 79 | 31.16 | 31.7 | Confirmed
[7 empty] | 86 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 50796 | 0.43 | 8.6 | Confirmed | 61 | 96 | 79.58 | 78.09 | Confirmed | 98 | 50.7 | 49.3 | Confirmed
[7 empty] | 87.4833 | 18 | TRUE | 2-Butanone (MEK) | 4.83 | 30705 | 0.26 | 15.747 | Confirmed | 43 | 72 | 26.8 | 29.77 | Confirmed | 57 | 8.71 | 7.99 | Confirmed
[7 empty] | 91.06 | 10 | TRUE | Methyl acrylate | 4.93 | 32574 | 0.28 | 9.106 | Confirmed | 55 | 85 | 16.43 | 17.66 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 87.1 | 10 | TRUE | Methacrylonitrile | 5.05 | 18861 | 0.16 | 8.71 | Confirmed | 67 | 52 | 33.66 | 30.35 | Confirmed | 40 | 41.59 | 36.35 | Confirmed
[7 empty] | 87.97 | 10 | TRUE | Bromochloromethane | 5.06 | 32814 | 0.28 | 8.797 | Confirmed | 49 | 130 | 87.46 | 84.24 | Confirmed | 128 | 66.82 | 63.44 | Confirmed
[7 empty] | 85.45 | 10 | TRUE | Tetrahydrofuran | 5.08 | 13079 | 0.11 | 8.545 | Confirmed | 42 | 72 | 39.18 | 41.53 | Confirmed | 71 | 39.07 | 43.62 | Confirmed
[7 empty] | 87.03 | 10 | TRUE | Trichloromethane (chloroform) | 5.19 | 58089 | 0.49 | 8.703 | Confirmed | 83 | 85 | 65.26 | 65.28 | Confirmed | 47 | 18.4 | 17.42 | Confirmed
[7 empty] | 85.32 | 10 | TRUE | 1,1,1-Trichloroethane | 5.33 | 36572 | 0.31 | 8.532 | Confirmed | 97 | 99 | 63.31 | 63.67 | Confirmed | 61 | 39.67 | 42.21 | Confirmed
[7 empty] | 101.285 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.35 | 68003 | 0.57 | 20.257 | Confirmed | 113 | 111 | 101.16 | 105.3 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 134508 | 1.14 | 20 | Confirmed | 168 | 99 | 44.54 | 49.68 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 83.25 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 53619 | 0.45 | 8.325 | Confirmed | 56 | 41 | 57.62 | 55.49 | Confirmed | 43 | 26.17 | 26.01 | Confirmed
[7 empty] | 85.82 | 10 | TRUE | Carbon tetrachloride | 5.48 | 27194 | 0.23 | 8.582 | Confirmed | 119 | 121 | 31.46 | 32.44 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 84.17 | 10 | TRUE | 1,1-Dichloropropene | 5.5 | 38822 | 0.33 | 8.417 | Confirmed | 75 | 77 | 31.06 | 31.7 | Confirmed | 110 | 44.53 | 40.85 | Confirmed
[7 empty] | 83.56 | 10 | TRUE | Benzene | 5.7 | 145039 | 1.22 | 8.356 | Confirmed | 78 | 77 | 23.38 | 24.54 | Confirmed | 52 | 14.64 | 14.33 | Confirmed
[7 empty] | 88.1 | 10 | TRUE | 1,2-Dichloroethane | 5.77 | 41631 | 0.35 | 8.81 | Confirmed | 62 | 64 | 31.86 | 32.21 | Confirmed | 49 | 31.03 | 29.38 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.16 | 225262 | 1.9 | 20 | Confirmed | 114 | 88 | 17.05 | 17.96 | Confirmed | 63 | 16.81 | 18.18 | Confirmed
[7 empty] | 82.25 | 10 | TRUE | Trichloroethene | 6.38 | 36804 | 0.31 | 8.225 | Confirmed | 130 | 132 | 97.52 | 97.05 | Confirmed | 95 | 89.89 | 96.47 | Confirmed
[7 empty] | 87.33 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 37259 | 0.31 | 8.733 | Confirmed | 63 | 62 | 69.37 | 69.47 | Confirmed | 41 | 43.58 | 39.86 | Confirmed
[7 empty] | 93.09 | 10 | TRUE | Dibromomethane | 6.72 | 27221 | 0.23 | 9.309 | Confirmed | 174 | 93 | 81.5 | 90.4 | Confirmed | 95 | 68.91 | 75.78 | Confirmed
[7 empty] | 92.79 | 10 | TRUE | Methyl methacrylate | 6.73 | 26519 | 0.22 | 9.279 | Confirmed | 41 | 69 | 84.46 | 93.38 | Confirmed | 39 | 46.2 | 42.97 | Confirmed
[7 empty] | 85.77 | 10 | TRUE | Bromodichloromethane | 6.91 | 39314 | 0.33 | 8.577 | Confirmed | 83 | 85 | 64.36 | 65.35 | Confirmed | 47 | 16.04 | 15.01 | Confirmed
[7 empty] | 88.35 | 10 | TRUE | 2-Nitropropane | 7.14 | 7509 | 0.06 | 8.835 | Confirmed | 43 | 41 | 81.84 | 78.95 | Confirmed | 39 | 25.82 | 26.23 | Confirmed
[7 empty] | 87.5 | 10 | TRUE | cis-1,3-Dichloropropene | 7.35 | 44265 | 0.37 | 8.75 | Confirmed | 75 | 39 | 44.86 | 38.37 | Confirmed | 77 | 32.49 | 32.1 | Confirmed
[7 empty] | 87.0111 | 18 | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.51 | 67799 | 0.57 | 15.662 | Confirmed | 43 | 58 | 40.35 | 42.94 | Confirmed | 41 | 23.25 | 23.09 | Confirmed
[7 empty] | 99.49 | 20 | TRUE | Toluene-d8 [SS2] | 7.6 | 287242 | 2.43 | 19.898 | Confirmed | 98 | 100 | 65.28 | 64.17 | Confirmed | 70 | 10.31 | 11.04 | Confirmed
[7 empty] | 83.99 | 10 | TRUE | Toluene | 7.67 | 150172 | 1.27 | 8.399 | Confirmed | 91 | 92 | 58.02 | 57.26 | Confirmed | 65 | 10.97 | 11.57 | Confirmed
[7 empty] | 89.76 | 10 | TRUE | trans-1,3-Dichloropropene | 7.92 | 34440 | 0.29 | 8.976 | Confirmed | 75 | 39 | 44.91 | 38.67 | Confirmed | 77 | 31.89 | 31.24 | Confirmed
[7 empty] | 92.35 | 10 | TRUE | Ethyl methacrylate | 7.99 | 45321 | 0.38 | 9.235 | Confirmed | 69 | 41 | 64.81 | 57.93 | Confirmed | 99 | 26.63 | 24.52 | Confirmed
[7 empty] | 91.53 | 10 | TRUE | 1,1,2-Trichloroethane | 8.09 | 35035 | 0.3 | 9.153 | Confirmed | 97 | 83 | 84.16 | 86.46 | Confirmed | 99 | 62.29 | 61.11 | Confirmed
[7 empty] | 78.97 | 10 | FALSE | Tetrachloroethene | 8.15 | 48161 | 0.41 | 7.897 | Confirmed | 166 | 164 | 77.08 | 79.12 | Confirmed | 129 | 67.48 | 72.06 | Confirmed
[7 empty] | 91.92 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 59165 | 0.5 | 9.192 | Confirmed | 76 | 41 | 64.98 | 57.89 | Confirmed | 78 | 32.57 | 31.66 | Confirmed
[7 empty] | 88.8556 | 18 | TRUE | 2-Hexanone | 8.31 | 47451 | 0.4 | 15.994 | Confirmed | 43 | 58 | 55.72 | 59.86 | Confirmed | 57 | 19.29 | 20.58 | Confirmed
[7 empty] | 90.5 | 10 | TRUE | Dibromochloromethane | 8.42 | 31026 | 0.26 | 9.05 | Confirmed | 129 | 127 | 76.76 | 78.14 | Confirmed | 131 | 23.36 | 23.41 | Confirmed
[7 empty] | 92.88 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.51 | 34198 | 0.29 | 9.288 | Confirmed | 107 | 109 | 94.98 | 96.19 | Confirmed | 93 | 3.9 | 4.45 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 211361 | 1.78 | 20 | Confirmed | 117 | 82 | 53.61 | 58 | Confirmed | 52 | 13.31 | 13.62 | Confirmed
[7 empty] | 84.47 | 10 | TRUE | Chlorobenzene | 8.93 | 99651 | 0.84 | 8.447 | Confirmed | 112 | 77 | 60.85 | 63.71 | Confirmed | 114 | 31.83 | 32.36 | Confirmed
[7 empty] | 85.64 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 27376 | 0.23 | 8.564 | Confirmed | 131 | 133 | 96.02 | 95.85 | Confirmed | 117 | 82.72 | 77.07 | Confirmed
[7 empty] | 85.07 | 10 | TRUE | Ethylbenzene | 9.02 | 159460 | 1.35 | 8.507 | Confirmed | 91 | 106 | 37.89 | 36.24 | Confirmed | 51 | 8.74 | 8.64 | Confirmed
[7 empty] | 88.13 | 10 | TRUE | m/p-Xylene | 9.12 | 286576 | 2.42 | 8.813 | Confirmed | 91 | 106 | 55.32 | 51.54 | Confirmed | 105 | 21.95 | 20.98 | Confirmed
[7 empty] | 86.44 | 10 | TRUE | o-Xylene | 9.42 | 148103 | 1.25 | 8.644 | Confirmed | 91 | 106 | 52.77 | 50.66 | Confirmed | 105 | 25.68 | 25.09 | Confirmed
[7 empty] | 87.73 | 10 | TRUE | Styrene | 9.44 | 123962 | 1.05 | 8.773 | Confirmed | 104 | 78 | 49.97 | 54.69 | Confirmed | 103 | 52.14 | 53.79 | Confirmed
[7 empty] | 86.51 | 10 | TRUE | Bromoform | 9.57 | 22329 | 0.19 | 8.651 | Confirmed | 173 | 171 | 50.06 | 51.71 | Confirmed | 175 | 47.63 | 48.3 | Confirmed
[7 empty] | 83.6 | 10 | TRUE | Isopropylbenzene (cumene) | 9.7 | 164677 | 1.39 | 8.36 | Confirmed | 105 | 120 | 31.21 | 30.22 | Confirmed | 79 | 14.85 | 15.92 | Confirmed
[7 empty] | 97.74 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.83 | 116760 | 0.99 | 19.548 | Confirmed | 95 | 174 | 88.43 | 79.25 | Confirmed | 176 | 85.06 | 75.83 | Confirmed
[7 empty] | 86.33 | 10 | TRUE | Bromobenzene | 9.93 | 74650 | 0.63 | 8.633 | Confirmed | 77 | 156 | 72.68 | 63.48 | Confirmed | 158 | 71.4 | 61.76 | Confirmed
[7 empty] | 90.01 | 10 | TRUE | 1,1,2,2-Tetrachloroethane | 9.94 | 46651 | 0.39 | 9.001 | Confirmed | 83 | 85 | 65.36 | 65.99 | Confirmed | 95 | 14.99 | 13.06 | Confirmed
[7 empty] | 88.3 | 10 | TRUE | 1,2,3-Trichloropropane (TCP) | 9.97 | 17812 | 0.15 | 8.83 | Confirmed | 77 | 110 | 82.83 | 78.3 | Confirmed | 61 | 55.97 | 58.37 | Confirmed
[7 empty] | 86.6 | 10 | TRUE | trans-1,4-Dichloro-2-butene | 9.98 | 53734 | 0.45 | 8.66 | Confirmed | 75 | 53 | 19.65 | 17.84 | Confirmed | 89 | 12 | 10.72 | Confirmed
[7 empty] | 82.34 | 10 | TRUE | n-Propylbenzene | 10.01 | 196913 | 1.66 | 8.234 | Confirmed | 91 | 120 | 28.56 | 26.47 | Confirmed | 65 | 9.75 | 10.19 | Confirmed
[7 empty] | 81.25 | 10 | TRUE | 2-Chlorotoluene | 10.07 | 120481 | 1.02 | 8.125 | Confirmed | 91 | 126 | 39.64 | 38.01 | Confirmed | 89 | 17.23 | 17.62 | Confirmed
[7 empty] | 81.11 | 10 | TRUE | 1,3,5-Trimethylbenzene | 10.14 | 151286 | 1.28 | 8.111 | Confirmed | 105 | 120 | 52.59 | 50.11 | Confirmed | 119 | 11.93 | 11.65 | Confirmed
[7 empty] | 81.21 | 10 | TRUE | 4-Chlorotoluene | 10.16 | 146471 | 1.24 | 8.121 | Confirmed | 91 | 126 | 35.46 | 33.22 | Confirmed | 89 | 10.88 | 11.31 | Confirmed
[7 empty] | 85.24 | 10 | TRUE | tert-Butylbenzene | 10.37 | 143697 | 1.21 | 8.524 | Confirmed | 119 | 91 | 63.34 | 63.24 | Confirmed | 134 | 25.06 | 23.88 | Confirmed
[7 empty] | 97.93 | 10 | TRUE | Pentachloroethane | 10.39 | 13675 | 0.12 | 9.793 | Confirmed | 167 | 130 | 53.8 | 55.45 | Confirmed | 132 | 62.91 | 53.75 | Confirmed
[7 empty] | 80.48 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.41 | 149262 | 1.26 | 8.048 | Confirmed | 105 | 120 | 50.81 | 48.67 | Confirmed | 77 | 10.34 | 11.4 | Confirmed
[7 empty] | 81.53 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.52 | 187465 | 1.58 | 8.153 | Confirmed | 105 | 134 | 22.93 | 21.87 | Confirmed | 91 | 14.28 | 14.99 | Confirmed
[7 empty] | 85.02 | 10 | TRUE | 1,3-Dichlorobenzene | 10.6 | 89163 | 0.75 | 8.502 | Confirmed | 146 | 148 | 63.81 | 62.38 | Confirmed | 111 | 39.53 | 42.69 | Confirmed
[7 empty] | 84 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.63 | 157900 | 1.33 | 8.4 | Confirmed | 119 | 91 | 25.02 | 27.29 | Confirmed | 134 | 31.11 | 29.42 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 130447 | 1.1 | 20 | Confirmed | 152 | 115 | 57.32 | 53.78 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 86.58 | 10 | TRUE | 1,4-Dichlorobenzene | 10.67 | 92716 | 0.78 | 8.658 | Confirmed | 146 | 148 | 63.43 | 63.68 | Confirmed | 111 | 44.06 | 43.3 | Confirmed
[7 empty] | 80.35 | 10 | TRUE | n-Butylbenzene | 10.91 | 141614 | 1.2 | 8.035 | Confirmed | 91 | 92 | 52.7 | 53.94 | Confirmed | 134 | 31.86 | 29.13 | Confirmed
[7 empty] | 86.93 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 95647 | 0.81 | 8.693 | Confirmed | 146 | 148 | 62.65 | 62.47 | Confirmed | 111 | 40.2 | 42.67 | Confirmed
[7 empty] | 83.56 | 10 | TRUE | Hexachloroethane | 11.1 | 17683 | 0.15 | 8.356 | Confirmed | 117 | 119 | 96.49 | 98.02 | Confirmed | 201 | 98.24 | 88.11 | Confirmed
[7 empty] | 88.89 | 10 | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.44 | 11372 | 0.1 | 8.889 | Confirmed | 157 | 155 | 76.06 | 77.01 | Confirmed | 75 | 77.06 | 90.06 | Confirmed
[7 empty] | 80.55 | 10 | TRUE | Nitrobenzene | 11.57 | 2891 | 0.02 | 8.055 | Confirmed | 77 | 51 | 49.41 | 46.53 | Confirmed | 123 | 54.32 | 47.05 | Confirmed
[7 empty] | 85.21 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.97 | 61108 | 0.52 | 8.521 | Confirmed | 180 | 182 | 96.18 | 94.83 | Confirmed | 145 | 30.56 | 31.27 | Confirmed
[7 empty] | 85.15 | 10 | TRUE | Hexachloro-1,3-butadiene | 12.06 | 26790 | 0.23 | 8.515 | Confirmed | 225 | 227 | 62.56 | 62.61 | Confirmed | 223 | 61.61 | 61.68 | Confirmed
[7 empty] | 84.14 | 10 | TRUE | Naphthalene | 12.14 | 172653 | 1.46 | 8.414 | Confirmed | 128 | 127 | 12.21 | 12.7 | Confirmed | 129 | 10.62 | 10.49 | Confirmed
[7 empty] | 85.64 | 10 | TRUE | 1,2,3-Trichlorobenzene | 12.28 | 61995 | 0.52 | 8.564 | Confirmed | 180 | 182 | 96.06 | 93.88 | Confirmed | 145 | 31.93 | 32.87 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2023
231025
Peak Name | Amount
ppb
First Injection | CAL2-1 | CAL4-5 | CAL5-20 | CAL6-100 | H2O-V3 | H2O-V4 | H2O-V5 | MB | 10-10-23-JI-01 | 10-10-23-JI-02 | 10-10-23-JI-03 | 10-10-23-JI-04 | 10-10-23-JI-05 | 10-12-23-JI-01 | 10-12-23-JI-02 | 10-12-23-JI-03 | 10-12-23-JI-04 | CCV | H2O-V6
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4526 | n.a./n.r.
Chloroethene (vinyl chloride) | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8898 | n.a./n.r.
Bromomethane (methyl bromide) | 1 | 5 | 20 | 100 | 0.5296 | 0.3859 | 0.2713 | 0.251 | n.a./n.r. | 0.2001 | 0.2439 | 0.2338 | 0.1938 | 0.2099 | 0.1706 | 0.184 | 0.1821 | 10.9238 | 0.2296
Chloroethane (ethyl chloride) | 0.659 | 4.7855 | 20.6621 | 99.8817 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5269 | n.a./n.r.
Trichlorofluoromethane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3049 | n.a./n.r.
Diethyl ether | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8189 | n.a./n.r.
1,1-Dichloroethene | 1 | 5 | 20 | 100 | 0.1477 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3082 | n.a./n.r.
Acetone | 1.8 | 9 | 36 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 25.4181 | 23.2717 | 36.8503 | 41.2477 | 52.8314 | 21.6683 | 28.206 | 25.5856 | 27.029 | 14.8389 | n.a./n.r.
Iodomethane | 1 | 5 | 20 | 100 | 0.7872 | 0.5842 | 0.3874 | 0.3592 | 0.3422 | 0.2794 | 0.3161 | 0.2131 | 0.2268 | 0.1807 | 0.2211 | 0.2173 | n.a./n.r. | 6.6171 | 0.2734
Carbon disulfide | 1 | 5 | 20 | 100 | 0.687 | 0.2844 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3473 | n.a./n.r.
3-Chloropropene (allyl chloride) | 1 | 5 | 20 | 100 | 0.0746 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2967 | n.a./n.r.
Methylene chloride (DCM) | 1 | 5 | 20 | 100 | 0.1046 | 0.0635 | 0.1919 | 0.1422 | n.a./n.r. | 0.0355 | 0.0526 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0309 | n.a./n.r. | n.a./n.r. | 8.5111 | 0.1745
trans-1,2-Dichloroethene | 1 | 5 | 20 | 100 | 0.3134 | 0.1013 | 0.0438 | n.a./n.r. | 0.0239 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4524 | n.a./n.r.
Methyl tert-butyl ether (MTBE) | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.953 | n.a./n.r.
1,1-Dichloroethane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7072 | n.a./n.r.
2,2-Dichloropropane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.6608 | n.a./n.r.
cis-1,2-Dichloroethene | 1 | 5 | 20 | 100 | 0.2016 | 0.0422 | 0.0221 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6001 | n.a./n.r.
2-Butanone (MEK) | 1.8 | 9 | 36 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.7474 | n.a./n.r.
Methyl acrylate | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1061 | n.a./n.r.
Bromochloromethane | 1 | 5 | 20 | 100 | 0.1104 | 0.0369 | 0.0125 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7974 | n.a./n.r.
Methacrylonitrile | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7101 | n.a./n.r.
Tetrahydrofuran | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5445 | n.a./n.r.
Trichloromethane (chloroform) | 1 | 5 | 20 | 100 | 0.0807 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 42.5147 | 62.9478 | 50.2325 | 39.1619 | 34.177 | 32.1047 | 31.3511 | 29.0068 | 31.2316 | 8.7035 | 0.0229
1,1,1-Trichloroethane | 1 | 5 | 20 | 100 | 0.0723 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5319 | n.a./n.r.
Dibromofluoromethane [SS1] | 19.7378 | 20.0422 | 20.16 | 20.06 | 20.7458 | 20.5973 | 20.3272 | 21.8528 | 20.3913 | 20.2792 | 20.8532 | 20.6102 | 20.7877 | 20.715 | 20.9483 | 20.4983 | 20.8187 | 20.257 | 19.8423
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | 1 | 5 | 20 | 100 | 0.3196 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3255 | n.a./n.r.
Carbon tetrachloride | 1 | 5 | 20 | 100 | 0.0488 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5818 | n.a./n.r.
1,1-Dichloropropene | 1 | 5 | 20 | 100 | 0.4633 | 0.1281 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4175 | n.a./n.r.
Benzene | 1 | 5 | 20 | 100 | 0.1349 | 0.0403 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3561 | n.a./n.r.
1,2-Dichloroethane | 1 | 5 | 20 | 100 | 0.1052 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8097 | n.a./n.r.
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | 1 | 5 | 20 | 100 | 0.3812 | 0.107 | 0.0559 | n.a./n.r. | 0.0211 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2255 | n.a./n.r.
1,2-Dichloropropane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7329 | n.a./n.r.
Dibromomethane | 1 | 5 | 20 | 100 | 0.2078 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0156 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3091 | n.a./n.r.
Methyl methacrylate | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2792 | n.a./n.r.
Bromodichloromethane | 1 | 5 | 20 | 100 | 0.1164 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.5209 | 18.78 | 18.2568 | 14.5607 | 12.5401 | 13.0497 | 13.0393 | 11.6895 | 12.1207 | 8.577 | n.a./n.r.
2-Nitropropane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8354 | n.a./n.r.
cis-1,3-Dichloropropene | 1 | 5 | 20 | 100 | 0.3311 | 0.0949 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7504 | n.a./n.r.
4-Methyl-2-pentanone (MIBK) | 1.8 | 9 | 36 | 180 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.6621 | n.a./n.r.
Toluene-d8 [SS2] | 19.3955 | 20.1511 | 20.1549 | 20.2986 | 19.881 | 19.6778 | 19.7251 | 18.7953 | 18.9459 | 19.0285 | 19.5003 | 19.0098 | 19.3569 | 19.4198 | 19.5219 | 19.4815 | 19.2194 | 19.8977 | 18.9663
Toluene | 1 | 5 | 20 | 100 | 0.459 | 0.1547 | 0.0936 | 0.0806 | 0.3995 | 0.2874 | 0.3734 | 0.2711 | 0.3042 | 0.2161 | 0.1505 | 0.1568 | 0.1893 | 8.3985 | 0.098
trans-1,3-Dichloropropene | 1 | 5 | 20 | 100 | 0.5139 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9756 | n.a./n.r.
Ethyl methacrylate | 1 | 5 | 20 | 100 | 0.1391 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2346 | n.a./n.r.
1,1,2-Trichloroethane | 1 | 5 | 20 | 100 | 0.1467 | 0.0332 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1528 | n.a./n.r.
Tetrachloroethene | 1 | 5 | 20 | 100 | 0.4598 | 0.0987 | 0.0422 | 0.0353 | 0.024 | 0.0221 | 0.0168 | n.a./n.r. | 0.0129 | 0.0119 | n.a./n.r. | 0.0065 | n.a./n.r. | 7.8975 | 0.063
1,3-Dichloropropane | 1 | 5 | 20 | 100 | 0.1881 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1921 | n.a./n.r.
2-Hexanone | 1.8 | 9 | 36 | 180 | 0.2413 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.9941 | n.a./n.r.
Dibromochloromethane | 1 | 5 | 20 | 100 | 0.1871 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.0789 | 7.0219 | 8.5465 | 5.9214 | 4.5691 | 5.33 | 4.8862 | 4.3668 | 4.9497 | 9.0504 | n.a./n.r.
1,2-Dibromoethane (EDB) | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.288 | n.a./n.r.
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | 1 | 5 | 20 | 100 | 0.7114 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4466 | n.a./n.r.
1,1,1,2-Tetrachloroethane | 1 | 5 | 20 | 100 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5636 | n.a./n.r.
Ethylbenzene | 1 | 5 | 20 | 100 | 0.8822 | 0.2359 | 0.1257 | 0.094 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5071 | 0.1403
m/p-Xylene | 1 | 5 | 20 | 100 | 0.8929 | 0.2484 | 0.1415 | 0.1133 | 0.0529 | 0.0586 | 0.0365 | 0.0295 | 0.0291 | 0.0217 | 0.0264 | 0.0246 | 0.0261 | 8.8133 | 0.1487
o-Xylene | 1 | 5 | 20 | 100 | 0.5762 | 0.1652 | 0.0953 | 0.0832 | 0.0283 | n.a./n.r. | 0.0246 | n.a./n.r. | n.a./n.r. | 0.0123 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6442 | 0.1155
Styrene | 1 | 5 | 20 | 100 | 0.9215 | 0.2543 | 0.1354 | 0.0922 | 0.0472 | 0.0383 | 0.0263 | 0.0226 | 0.0233 | n.a./n.r. | 0.017 | n.a./n.r. | n.a./n.r. | 8.7729 | 0.1285
Bromoform | 1 | 5 | 20 | 100 | 0.2504 | 0.0494 | 0.0258 | 0.0183 | 0.9128 | 0.5187 | 1.117 | 0.5264 | 0.3924 | 0.5438 | 0.3302 | 0.315 | 0.4273 | 8.6513 | 0.024
Isopropylbenzene (cumene) | 1 | 5 | 20 | 100 | 0.8901 | 0.192 | 0.0708 | 0.0408 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3602 | 0.1334
1-Bromo-4-fluorobenzene (BFB) [SS3] | 19.4911 | 19.7434 | 20.109 | 20.6565 | 19.971 | 20.7179 | 20.83 | 21.1664 | 19.8513 | 20.4703 | 20.1196 | 20.2598 | 20.6587 | 20.4922 | 21.2655 | 20.0697 | 20.6332 | 19.5475 | 21.2132
Bromobenzene | 1 | 5 | 20 | 100 | 0.8886 | 0.2621 | 0.108 | 0.079 | 0.0541 | 0.0366 | 0.0279 | 0.029 | 0.0163 | 0.0184 | 0.0168 | 0.0122 | n.a./n.r. | 8.6326 | 0.12
1,1,2,2-Tetrachloroethane | 1 | 5 | 20 | 100 | 0.2154 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0007 | n.a./n.r.
trans-1,4-Dichloro-2-butene | 1 | 5 | 20 | 100 | 0.3073 | 0.0732 | 0.0357 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6599 | 0.0445
1,2,3-Trichloropropane (TCP) | 1 | 5 | 20 | 100 | 0.3066 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8305 | n.a./n.r.
n-Propylbenzene | 1 | 5 | 20 | 100 | 1.3601 | 0.3949 | 0.162 | 0.0866 | 0.0625 | 0.0422 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2341 | 0.2037
2-Chlorotoluene | 1 | 5 | 20 | 100 | 0.9898 | 0.2497 | 0.1086 | 0.0648 | 0.0413 | 0.0317 | 0.0222 | n.a./n.r. | 0.0191 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1252 | 0.1423
1,3,5-Trimethylbenzene | 1 | 5 | 20 | 100 | 1.0169 | 0.233 | 0.0814 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.111 | 0.1469
4-Chlorotoluene | 1 | 5 | 20 | 100 | 1.226 | 0.3806 | 0.1737 | 0.1127 | 0.0781 | 0.0516 | 0.0384 | 0.0313 | n.a./n.r. | 0.0261 | n.a./n.r. | 0.0167 | n.a./n.r. | 8.1208 | 0.1714
Pentachloroethane | 1 | 5 | 20 | 100 | 0.465 | 0.067 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7932 | n.a./n.r.
tert-Butylbenzene | 1 | 5 | 20 | 100 | 0.771 | 0.1554 | 0.0521 | 0.0317 | 0.0197 | 0.0126 | 0.0102 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5241 | 0.1141
1,2,4-Trimethylbenzene | 1 | 5 | 20 | 100 | 1.1016 | 0.2743 | 0.1169 | 0.0795 | 0.0631 | 0.0515 | 0.0425 | 0.0401 | 0.0374 | 0.0288 | 0.0364 | 0.0293 | 0.0346 | 8.048 | 0.1631
1-Methylpropylbenzene (sec-butylbenzene) | 1 | 5 | 20 | 100 | 1.3013 | n.a./n.r. | 0.1202 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1526 | 0.1979
1,3-Dichlorobenzene | 1 | 5 | 20 | 100 | 1.3794 | 0.4306 | 0.2113 | 0.1437 | 0.091 | 0.0673 | 0.0489 | 0.0446 | 0.036 | 0.0332 | 0.0278 | 0.0187 | 0.0216 | 8.502 | 0.1817
4-Isopropyltoluene (p-cymene) | 1 | 5 | 20 | 100 | 1.4377 | 0.4053 | 0.1453 | 0.0735 | 0.0493 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3997 | 0.2025
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | 1 | 5 | 20 | 100 | 1.379 | 0.4608 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6579 | n.a./n.r.
n-Butylbenzene | 1 | 5 | 20 | 100 | 1.8714 | 0.6871 | 0.3037 | 0.1778 | 0.1239 | 0.0668 | 0.0623 | 0.0486 | 0.0368 | 0.0354 | 0.0324 | 0.0211 | 0.0194 | 8.0352 | 0.2642
1,2-Dichlorobenzene | 1 | 5 | 20 | 100 | 1.1225 | 0.3118 | 0.1365 | 0.0881 | 0.065 | 0.0424 | 0.0395 | 0.0289 | 0.0272 | 0.0256 | 0.0187 | 0.0178 | 0.0166 | 8.693 | 0.1631
Hexachloroethane | 1 | 5 | 20 | 100 | 0.4231 | 0.1069 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3563 | n.a./n.r.
1,2-Dibromo-3-chloropropane (DBCP) | 1 | 5 | 20 | 100 | 0.5127 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8885 | n.a./n.r.
Nitrobenzene | 1 | 5 | 20 | 100 | 1.1173 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0553 | n.a./n.r.
1,2,4-Trichlorobenzene | 1 | 5 | 20 | 100 | 1.9231 | 0.6427 | 0.2985 | 0.205 | 0.1383 | 0.0897 | 0.0704 | 0.0678 | 0.0447 | 0.0501 | 0.0353 | 0.0358 | 0.0293 | 8.5207 | 0.2532
Hexachloro-1,3-butadiene | 1 | 5 | 20 | 100 | 1.7046 | 0.563 | 0.2245 | 0.1482 | 0.1027 | 0.0671 | 0.0511 | 0.042 | 0.0263 | 0.0298 | 0.0263 | 0.0157 | 0.0203 | 8.5152 | 0.2159
Naphthalene | 1 | 5 | 20 | 100 | 1.4855 | 0.4012 | 0.1585 | 0.112 | 0.0751 | 0.0505 | 0.0418 | 0.0337 | 0.0318 | 0.0272 | 0.0218 | 0.0194 | 0.0199 | 8.4138 | 0.1938
1,2,3-Trichlorobenzene | 1 | 5 | 20 | 100 | 1.8407 | 0.5495 | 0.2238 | 0.1677 | 0.0973 | 0.0634 | 0.0523 | 0.0513 | 0.0326 | 0.0333 | 0.0235 | 0.0233 | 0.0192 | 8.5639 | 0.2415

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library | Area
| No. | min | SI |  |  | SI |  |  | SI |  |  | counts*min
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC