## IS-RSD
[8 empty] | Peak Name | Peak Area
[9 empty] | counts*min | 1 | 2 | 3 | 4 | 5 | 6
Total Ion Count | [7 empty] | First Injection | Startup | H2O-R1 | H2O-R2 | H2O-R3 | H2O-R4
RSD | RSD (1-3) | RSD (2-4) | RSD (3-5) | RSD (4-6) | RSD (5-7) | RSD (6-8) | RSD (7-9) | TIC | TIC | TIC | TIC | TIC | TIC
0.4431 | #DIV/0! | #DIV/0! | 0.62645 | 0.4431 | 0.3246 | #DIV/0! | #DIV/0! | Pentafluorobenzene [IS1] |  |  |  |  | 439321.644 | 443231.047 | 441201.047
1.2248 | #DIV/0! | #DIV/0! | 1.4942 | 1.2248 | 0.022124 | #DIV/0! | #DIV/0! | 1,4-Difluorobenzene [IS2] |  |  |  |  | 637772.245 | 651392.997 | 651596.833
2.0358 | #DIV/0! | #DIV/0! | 2.0023 | 2.0358 | 0.80403 | #DIV/0! | #DIV/0! | Chlorobenzene-d5 [IS3] |  |  |  |  | 775665.321 | 797945.198 | 807070.248
4.2845 | #DIV/0! | #DIV/0! | 2.0351 | 4.2845 | 3.8905 | #DIV/0! | #DIV/0! | 1,4-Dichlorobenzene-d4 [IS4] |  |  |  |  | 762843.135 | 785118.989 | 829538.603

## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2024/240111.seq/345.smp/MSDevice.xraw
Channel | TIC
Injection Information: |  |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Data Vault | ChromeleonLocal |  |  | 95 | 8944251 | 50-200% of mass 174 | 95/174 | 197.2869 | 50 | 200 | TRUE
Injection | H2O-V1 |  |  | 96 | 620187.1875 | 5 to 9% of m/z 95 | 96/95 | 6.9339 | 5 | 9 | TRUE
Injection Number | 9 |  |  | 173 | 12063.8975 | <2% of m/z 174 | 173/174 | 0.2661 | 0 | 2 | TRUE
Position | 2 |  |  | 174 | 4533625.5 | 50-200% of mass 95 | 174/95 | 50.6876 | 50 | 200 | TRUE
Comment |  |  |  | 175 | 362018.4375 | 5 to 9% of m/z 174 | 175/174 | 7.9852 | 5 | 9 | TRUE
Processing Method | VOC-11 |  |  | 176 | 4314496.5 | 95 to 105% of m/z 174 | 176/174 | 95.1666 | 95 | 105 | TRUE
Instrument Method | VOC-26 |  |  | 177 | 262960.25 | 5 to 10% of m/z 176 | 177/176 | 6.0948 | 5 | 10 | TRUE
Type | Unknown
Status | Finished
Injection Date | 2024-01-11
Injection Time | 15:55
Injection Volume (µl) | 20 | *N/A for P&T
Dilution Factor | 1899-12-31
Weight | 24:00
Raw Data Information:
Data Points | 236
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9808
Retention Time (min) | 9.8489
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3291
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0377
Base Peak Intensity | 8944251
Total Ion Count | 32848007.6761
Raw Data: |  | Rounded:
Mass (m/z) | Value (counts) | Mass (m/z) | Value (counts)
35.3 | 1976.7543 | 35 | 1976.7543
36 | 41904.2422 | 36 | 41904.2422
37 | 231001.5156 | 37 | 231001.5156
38 | 205799.2344 | 38 | 205799.2344
39.1 | 99145.6641 | 39 | 99145.6641
40 | 55264.4531 | 40 | 55264.4531
41.1 | 18054.5488 | 41 | 18054.5488
42.1 | 7308.3179 | 42 | 7308.3179
43.2 | 15827.3613 | 43 | 15827.3613
44 | 259014.5 | 44 | 259014.5
45 | 45113.8594 | 45 | 45113.8594
46.1 | 4880.856 | 46 | 4880.856
47 | 106593.4375 | 47 | 106593.4375
48 | 38824.8203 | 48 | 38824.8203
49 | 223434.7344 | 49 | 223434.7344
50 | 1208721.75 | 50 | 1208721.75
51 | 333946.1875 | 51 | 333946.1875
52.1 | 15427.0811 | 52 | 15427.0811
52.9 | 2436.7952 | 53 | 2436.7952
53.8 | 2903.8752 | 54 | 2903.8752
55.1 | 21324.1582 | 55 | 21324.1582
56 | 83388.0391 | 56 | 83388.0391
57 | 144478.3125 | 57 | 144478.3125
58.1 | 6503.8345 | 58 | 6503.8345
58.9 | 2189.2007 | 59 | 2189.2007
60 | 49442.8359 | 60 | 49442.8359
61 | 262859.375 | 61 | 262859.375
62 | 256458.3281 | 62 | 256458.3281
63 | 177411.4219 | 63 | 177411.4219
64 | 23002.7363 | 64 | 23002.7363
65 | 5021.1221 | 65 | 5021.1221
66.9 | 20640.6484 | 67 | 20640.6484
68 | 716727.25 | 68 | 716727.25
69 | 636170.875 | 69 | 636170.875
70 | 63305.9258 | 70 | 63305.9258
71.1 | 3351.4092 | 71 | 3351.4092
72 | 35255.5781 | 72 | 35255.5781
73 | 279379.5625 | 73 | 279379.5625
74 | 1093831.5 | 74 | 1093831.5
75 | 4026114.5 | 75 | 4026114.5
76 | 329235.1562 | 76 | 329235.1562
77.1 | 65124.8828 | 77 | 65124.8828
78.1 | 34830.5938 | 78 | 34830.5938
78.9 | 60821.0469 | 79 | 60821.0469
80 | 20489.8105 | 80 | 20489.8105
80.9 | 64559.7031 | 81 | 64559.7031
82.1 | 8864.0518 | 82 | 8864.0518
83 | 3409.634 | 83 | 3409.634
83.6 | 297.449 | 84 | 297.449
84.3 | 468.7755 | 84 | 468.7755
85.3 | 2161.8748 | 85 | 2161.8748
85.9 | 8351.6943 | 86 | 8351.6943
87 | 414475.3125 | 87 | 414475.3125
88 | 392537.8438 | 88 | 392537.8438
89.5 | 13.6405 | 90 | 13.6405
90.2 | 258.8912 | 90 | 258.8912
91 | 21036.9766 | 91 | 21036.9766
92 | 160441.2812 | 92 | 160441.2812
93 | 254709.9375 | 93 | 254709.9375
94 | 762718.875 | 94 | 762718.875
95 | 8944251 | 95 | 8944251
96 | 620187.1875 | 96 | 620187.1875
97.1 | 18773.8301 | 97 | 18773.8301
98.1 | 1681.9819 | 98 | 1681.9819
99.3 | 244.0296 | 99 | 244.0296
100.3 | 267.5058 | 100 | 267.5058
102.2 | 314.5019 | 102 | 314.5019
103 | 3293.8672 | 103 | 3293.8672
104 | 22239.8867 | 104 | 22239.8867
104.9 | 8644.1006 | 105 | 8644.1006
105.9 | 13442.4131 | 106 | 13442.4131
106.9 | 3393.136 | 107 | 3393.136
107.9 | 644.7021 | 108 | 644.7021
108.5 | 0.95416 | 109 | 0.95416
109.5 | 476.3502 | 110 | 476.3502
110.5 | 2514.261 | 111 | 2514.261
111.1 | 832.2116 | 111 | 832.2116
112.1 | 1479.7797 | 112 | 1479.7797
112.9 | 1021.7532 | 113 | 1021.7532
113.8 | 17.4319 | 114 | 17.4319
114.9 | 4353.7124 | 115 | 4353.7124
115.9 | 9800.6729 | 116 | 9800.6729
116.9 | 25627.4688 | 117 | 25627.4688
117.9 | 11634.2207 | 118 | 11634.2207
118.5 | 1001.6533 | 119 | 1001.6533
119 | 17949.8477 | 119 | 17949.8477
120.1 | 925.0338 | 120 | 925.0338
120.9 | 786.2062 | 121 | 786.2062
121.5 | 760.0021 | 122 | 760.0021
122.1 | 564.3289 | 122 | 564.3289
123 | 1075.4161 | 123 | 1075.4161
124 | 2821.3772 | 124 | 2821.3772
125.1 | 2288.8298 | 125 | 2288.8298
125.9 | 738.4069 | 126 | 738.4069
126.7 | 486.8229 | 127 | 486.8229
127.9 | 12811.4102 | 128 | 12811.4102
128.4 | 658.7406 | 128 | 658.7406
128.9 | 6893.6362 | 129 | 6893.6362
130 | 13538.3096 | 130 | 13538.3096
130.9 | 6082.2769 | 131 | 6082.2769
132 | 1121.7313 | 132 | 1121.7313
133.2 | 2653.8708 | 133 | 2653.8708
134 | 2213.2732 | 134 | 2213.2732
134.8 | 3892.3777 | 135 | 3892.3777
135.4 | 1898.4801 | 135 | 1898.4801
136 | 2036.0107 | 136 | 2036.0107
136.8 | 4834.978 | 137 | 4834.978
138.8 | 219.9188 | 139 | 219.9188
140.3 | 3425.1824 | 140 | 3425.1824
140.9 | 20854.9531 | 141 | 20854.9531
142.8 | 20946.6797 | 143 | 20946.6797
143.9 | 653.3989 | 144 | 653.3989
144.9 | 890.3549 | 145 | 890.3549
146 | 8595.0117 | 146 | 8595.0117
146.9 | 3613.5017 | 147 | 3613.5017
147.9 | 6451.5952 | 148 | 6451.5952
149 | 2995.4065 | 149 | 2995.4065
150.1 | 2762.9722 | 150 | 2762.9722
151.9 | 1640.2579 | 152 | 1640.2579
153 | 4180.5596 | 153 | 4180.5596
153.9 | 2595.3743 | 154 | 2595.3743
154.9 | 12270.2656 | 155 | 12270.2656
155.9 | 1618.5305 | 156 | 1618.5305
156.9 | 6763.6377 | 157 | 6763.6377
158 | 1319.5961 | 158 | 1319.5961
158.9 | 2839.353 | 159 | 2839.353
159.9 | 783.0017 | 160 | 783.0017
161 | 3756.2383 | 161 | 3756.2383
162.3 | 828.9796 | 162 | 828.9796
163.1 | 365.6845 | 163 | 365.6845
164 | 4.4e-05 | 164 | 4.4e-05
164.8 | 47.4259 | 165 | 47.4259
165.3 | 280.0924 | 165 | 280.0924
166.3 | 0.006743 | 166 | 0.006743
167.7 | 262.6771 | 168 | 262.6771
168.5 | 0.023839 | 169 | 0.023839
169.2 | 5e-06 | 169 | 5e-06
170.2 | 138.5838 | 170 | 138.5838
171 | 175.2401 | 171 | 175.2401
171.8 | 815.3143 | 172 | 815.3143
172.8 | 12063.8975 | 173 | 12063.8975
173.9 | 4533625.5 | 174 | 4533625.5
175 | 362018.4375 | 175 | 362018.4375
175.9 | 4314496.5 | 176 | 4314496.5
177 | 262960.25 | 177 | 262960.25
178 | 8190.0337 | 178 | 8190.0337
178.8 | 150.9435 | 179 | 150.9435
179.3 | 51.0604 | 179 | 51.0604
180.1 | 276.6968 | 180 | 276.6968
181.3 | 7.8553 | 181 | 7.8553
182 | 559.0389 | 182 | 559.0389
182.6 | 0.001508 | 183 | 0.001508
183.5 | 0 | 184 | 0
184.5 | 62.2576 | 185 | 62.2576
185.1 | 8.2798 | 185 | 8.2798
186.2 | 0 | 186 | 0
186.8 | 2e-05 | 187 | 2e-05
187.4 | 31.2376 | 187 | 31.2376
188.5 | 6.9015 | 189 | 6.9015
189.2 | 902.2305 | 189 | 902.2305
190.9 | 5400.2202 | 191 | 5400.2202
192 | 702.3046 | 192 | 702.3046
193 | 1387.7412 | 193 | 1387.7412
194.1 | 390.3834 | 194 | 390.3834
194.9 | 176.8678 | 195 | 176.8678
195.7 | 150.4255 | 196 | 150.4255
197 | 1.608 | 197 | 1.608
198.7 | 74.4631 | 199 | 74.4631
199.7 | 0.008583 | 200 | 0.008583
200.7 | 0.18753 | 201 | 0.18753
201.4 | 0.000202 | 201 | 0.000202
202.4 | 0.81618 | 202 | 0.81618
203.1 | 232.5139 | 203 | 232.5139
205.1 | 510.8368 | 205 | 510.8368
206.3 | 535.6074 | 206 | 535.6074
207 | 5784.1919 | 207 | 5784.1919
207.7 | 209.2083 | 208 | 209.2083
208.4 | 2586.3218 | 208 | 2586.3218
209.1 | 1425.1777 | 209 | 1425.1777
209.7 | 107.432 | 210 | 107.432
210.5 | 0.14312 | 211 | 0.14312
211.5 | 1.9e-05 | 212 | 1.9e-05
212 | 23.3756 | 212 | 23.3756
212.6 | 0 | 213 | 0
213.7 | 189.6306 | 214 | 189.6306
214.2 | 0.005518 | 214 | 0.005518
216 | 203.3676 | 216 | 203.3676
216.9 | 14.8556 | 217 | 14.8556
217.9 | 3.4353 | 218 | 3.4353
218.5 | 0.82391 | 219 | 0.82391
219.2 | 0.63129 | 219 | 0.63129
219.9 | 41.3005 | 220 | 41.3005
220.9 | 48.945 | 221 | 48.945
222.1 | 48.7924 | 222 | 48.7924
224 | 47.0276 | 224 | 47.0276
225.2 | 5e-06 | 225 | 5e-06
226.1 | 38.8146 | 226 | 38.8146
226.8 | 0 | 227 | 0
227.7 | 0 | 228 | 0
229.1 | 6e-06 | 229 | 6e-06
229.6 | 24.8582 | 230 | 24.8582
230.9 | 0.003947 | 231 | 0.003947
231.6 | 4.2679 | 232 | 4.2679
232.2 | 0.271 | 232 | 0.271
232.9 | 356.2486 | 233 | 356.2486
233.5 | 0.001348 | 234 | 0.001348
234.5 | 0.070788 | 235 | 0.070788
235 | 504.4507 | 235 | 504.4507
235.6 | 1.9e-05 | 236 | 1.9e-05
237.1 | 510.6422 | 237 | 510.6422
237.7 | 2.6535 | 238 | 2.6535
238.7 | 47.7096 | 239 | 47.7096
239.2 | 2.3e-05 | 239 | 2.3e-05
240.6 | 65.889 | 241 | 65.889
241.5 | 0 | 242 | 0
242 | 0 | 242 | 0
242.6 | 0 | 243 | 0
243.5 | 0 | 244 | 0
244.1 | 0 | 244 | 0
245.3 | 0.11365 | 245 | 0.11365
246.2 | 0.0001 | 246 | 0.0001
246.9 | 0.000341 | 247 | 0.000341
247.8 | 623.6497 | 248 | 623.6497
248.9 | 3080.0413 | 249 | 3080.0413
250.3 | 1257.0906 | 250 | 1257.0906
251 | 1101.8108 | 251 | 1101.8108
251.9 | 4.1387 | 252 | 4.1387
252.8 | 109.2701 | 253 | 109.2701
253.7 | 0 | 254 | 0
254.2 | 0 | 254 | 0
255.3 | 0.15606 | 255 | 0.15606
255.8 | 5e-06 | 256 | 5e-06
257.2 | 61.7699 | 257 | 61.7699
257.8 | 1.3e-05 | 258 | 1.3e-05
258.9 | 0.15175 | 259 | 0.15175
259.9 | 0 | 260 | 0
257.1 | 0.000177 | 257 | 0.000177
258.1 | 2.3e-05 | 258 | 2.3e-05
259.7 | 4.1155 | 260 | 4.1155
259.3 | 0
259.9 | 270.649
259.8 | 2.9e-05
257.3 | 0.004411
258.3 | 181.4352
259 | 0.043369
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[12 empty] | TV | [6 empty] | Measured | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
Name | RT | Window | Quant | Conf1 | Conf2 | Window |  | Channel | Eval. Type | IS | Cal.Type | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | RSD<20% | R^2>0.99 | Total Analytes | 81
Chloromethane (methyl chloride) | 1.47 | 0.050 AN | 50 | 52 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9285 | 1.9751 | 4.9378 | 9.6078 | 24.0219 | 100.2881 | 92.85 | 98.755 | 98.756 | 96.078 | 96.0876 | 100.2881 | Chloromethane (methyl chloride) | 1.47 | Lin | 6 | 2.0761 | 0.99984 | TRUE | TRUE | RSD Failed | 1
Chloroethene (vinyl chloride) | 1.56 | 0.050 AN | 62 | 64 | 61 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9106 | 2.0212 | 4.8554 | 9.2388 | 23.799 | 100.3841 | 91.06 | 101.06 | 97.108 | 92.388 | 95.196 | 100.3841 | Chloroethene (vinyl chloride) | 1.57 | Lin | 6 | 2.8176 | 0.9997 | TRUE | TRUE | R^2 Failed | 0
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7152 | 2.1232 | 5.3422 | 9.8184 | 25.0001 | 100.0024 | 71.52 | 106.16 | 106.844 | 98.184 | 100.0004 | 100.0024 | Bromomethane (methyl bromide) | 1.85 | Quad, WithOffset | 6 | 1.4899 | 0.99993 | TRUE | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9539 | 1.8709 | 5.0914 | 9.8932 | 24.1723 | 100.2161 | 95.39 | 93.545 | 101.828 | 98.932 | 96.6892 | 100.2161 | Chloroethane (ethyl chloride) | 1.96 | Lin | 6 | 1.6563 | 0.9999 | TRUE | TRUE
Trichlorofluoromethane | 2.2 | 0.050 AN | 101 | 103 | 105 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.876 | 1.9934 | 4.7725 | 8.9548 | 23.106 | 100.5908 | 87.6 | 99.67 | 95.45 | 89.548 | 92.424 | 100.5908 | Trichlorofluoromethane | 2.2 | Lin | 6 | 4.3176 | 0.99932 | TRUE | TRUE
Diethyl ether | 2.52 | 0.080 AN | 59 | 74 | 45 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8747 | 2.0222 | 4.8497 | 10.0799 | 24.5725 | 100.1072 | 87.47 | 101.11 | 96.994 | 100.799 | 98.29 | 100.1072 | Diethyl ether | 2.52 | Lin | 6 | 0.9213 | 0.99997 | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.915 | 2.007 | 4.8052 | 9.2302 | 23.5769 | 100.4432 | 91.5 | 100.35 | 96.104 | 92.302 | 94.3076 | 100.4432 | 1,1-Dichloroethene | 2.75 | Lin | 6 | 3.2186 | 0.99962 | TRUE | TRUE
Acetone | 2.87 | 0.088 AN | 43 | 58 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 3.0082 | 3.6089 | 8.455 | 21.1057 | 49.8103 | 200.0118 | 150.41 | 90.2225 | 84.55 | 105.5285 | 99.6206 | 100.0059 | Acetone | 2.86 | Lin, WithOffset | 6 | 2.3496 | 0.99984 | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, 1/A | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.5257 | 1.3331 | 3.9371 | 9.7826 | 27.646 | 99.6055 | 52.57 | 66.655 | 78.742 | 97.826 | 110.584 | 99.6055 | Iodomethane | 2.91 | Quad, 1/A | 6 | 31.3039 | 0.99507 | FALSE | TRUE
Carbon disulfide | 2.97 | 0.050 AN | 76 | 78 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8818 | 1.9768 | 4.8249 | 9.1604 | 23.4187 | 100.4897 | 88.18 | 98.84 | 96.498 | 91.604 | 93.6748 | 100.4897 | Carbon disulfide | 2.97 | Lin | 6 | 3.5616 | 0.99953 | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8446 | 1.7465 | 4.8572 | 9.979 | 25.0759 | 99.9966 | 84.46 | 87.325 | 97.144 | 99.79 | 100.3036 | 99.9966 | 3-Chloropropene (allyl chloride) | 3.21 | Quad | 6 | 0.6437 | 0.99999 | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0824 | 2.278 | 5.1978 | 10.5829 | 25.4397 | 99.8155 | 108.24 | 113.9 | 103.956 | 105.829 | 101.7588 | 99.8155 | Methylene chloride (DCM) | 3.38 | Lin | 6 | 1.5441 | 0.99991 | TRUE | TRUE
trans-1,2-Dichloroethene | 3.68 | 0.100 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9078 | 2.0811 | 4.8522 | 9.7005 | 23.8883 | 100.3146 | 90.78 | 104.055 | 97.044 | 97.005 | 95.5532 | 100.3146 | trans-1,2-Dichloroethene | 3.7 | Lin | 6 | 2.2885 | 0.9998 | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.892 | 1.9143 | 4.7177 | 10.0119 | 24.476 | 100.1467 | 89.2 | 95.715 | 94.354 | 100.119 | 97.904 | 100.1467 | Methyl tert-butyl ether (MTBE) | 3.71 | Lin | 6 | 1.1862 | 0.99995 | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9235 | 2.0589 | 4.9214 | 10.0394 | 24.1706 | 100.2069 | 92.35 | 102.945 | 98.428 | 100.394 | 96.6824 | 100.2069 | 1,1-Dichloroethane | 4.21 | Lin | 6 | 1.6303 | 0.9999 | TRUE | TRUE
2,2-Dichloropropane | 4.822 | 0.060 AN | 77 | 41 | 79 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7334 | 1.5822 | 4.0558 | 8.2344 | 21.6512 | 101.072 | 73.34 | 79.11 | 81.116 | 82.344 | 86.6048 | 101.072 | 2,2-Dichloropropane | 4.83 | Lin | 6 | 7.9651 | 0.99782 | TRUE | TRUE
cis-1,2-Dichloroethene | 4.83 | 0.060 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9329 | 1.9558 | 4.8173 | 9.9154 | 23.8327 | 100.311 | 93.29 | 97.79 | 96.346 | 99.154 | 95.3308 | 100.311 | cis-1,2-Dichloroethene | 4.84 | Lin | 6 | 2.3231 | 0.9998 | TRUE | TRUE
2-Butanone (MEK) | 4.86 | 0.060 AN | 43 | 72 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 1.8114 | 3.2409 | 8.5824 | 20.6024 | 50.3505 | 199.9401 | 90.57 | 81.0225 | 85.824 | 103.012 | 100.701 | 99.9701 | 2-Butanone (MEK) | 4.86 | Lin | 6 | 1.6634 | 0.99989 | TRUE | TRUE
Methyl acrylate | 4.953 | 0.060 AN | 55 | 85 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.391 | 2.0168 | 4.6648 | 9.7707 | 25.162 | 99.9943 | 139.1 | 100.84 | 93.296 | 97.707 | 100.648 | 99.9943 | Methyl acrylate | 4.96 | Quad, WithOffset | 6 | 1.2945 | 0.99997 | TRUE | TRUE
Methacrylonitrile | 5.073 | 0.060 AN | 49 | 130 | 128 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | 1.7387 | 4.1869 | 9.7104 | 24.1651 | 100.2868 | #VALUE! | 86.935 | 83.738 | 97.104 | 96.6604 | 100.2868 | Methacrylonitrile | 5.08 | Lin | 6 | 2.485 | 0.99977 | TRUE | TRUE
Bromochloromethane | 5.076 | 0.060 AN | 67 | 52 | 40 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9231 | 1.9961 | 4.8018 | 10.1663 | 24.5075 | 100.1173 | 92.31 | 99.805 | 96.036 | 101.663 | 98.03 | 100.1173 | Bromochloromethane | 5.09 | Lin | 6 | 1.0804 | 0.99996 | TRUE | TRUE
Tetrahydrofuran | 5.1 | 0.060 AN | 42 | 72 | 71 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7857 | 1.9825 | 4.4234 | 9.9315 | 25.0109 | 100.0354 | 78.57 | 99.125 | 88.468 | 99.315 | 100.0436 | 100.0354 | Tetrahydrofuran | 5.12 | Lin | 6 | 1.1707 | 0.99995 | TRUE | TRUE
Trichloromethane (chloroform) | 5.211 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9831 | 2.1441 | 5.0841 | 10.2909 | 24.6932 | 100.0407 | 98.31 | 107.205 | 101.682 | 102.909 | 98.7728 | 100.0407 | Trichloromethane (chloroform) | 5.22 | Lin | 6 | 0.8555 | 0.99997 | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8399 | 1.9045 | 4.5274 | 9.5429 | 23.82 | 100.3678 | 83.99 | 95.225 | 90.548 | 95.429 | 95.28 | 100.3678 | 1,1,1-Trichloroethane | 5.35 | Lin | 6 | 2.6877 | 0.99973 | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 19.4305 | 19.0288 | 19.7158 | 20.2783 | 20.7682 | 20.7785 | 97.1525 | 95.144 | 98.579 | 101.3915 | 103.841 | 103.8925 | Dibromofluoromethane [SS1] | 5.38 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Pentafluorobenzene [IS1] | 5.44 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.486 | 0.040 AN | 56 | 41 | 43 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8775 | 2.0601 | 4.7813 | 9.4316 | 24.2419 | 100.2573 | 87.75 | 103.005 | 95.626 | 94.316 | 96.9676 | 100.2573 | 1-Chlorobutane (butyl chloride) | 5.49 | Lin | 6 | 1.9228 | 0.99986 | TRUE | TRUE
Carbon tetrachloride | 5.501 | 0.040 AN | 119 | 121 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9989 | 2.1844 | 4.932 | 9.7969 | 25.0907 | 99.9971 | 99.89 | 109.22 | 98.64 | 97.969 | 100.3628 | 99.9971 | Carbon tetrachloride | 5.51 | Quad, WithOffset | 6 | 0.6509 | 0.99999 | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8712 | 2.0205 | 4.8255 | 9.4654 | 24.456 | 100.1991 | 87.12 | 101.025 | 96.51 | 94.654 | 97.824 | 100.1991 | 1,1-Dichloropropene | 5.52 | Lin | 6 | 1.5471 | 0.99991 | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8827 | 1.9145 | 4.4235 | 9.7561 | 23.7717 | 100.3632 | 88.27 | 95.725 | 88.47 | 97.561 | 95.0868 | 100.3632 | Benzene | 5.72 | Lin | 6 | 2.7249 | 0.99972 | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.039 | 2.0961 | 4.9429 | 11.0444 | 25.3265 | 99.8145 | 103.9 | 104.805 | 98.858 | 110.444 | 101.306 | 99.8145 | 1,2-Dichloroethane | 5.79 | Lin | 6 | 2.076 | 0.99983 | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.181 | 0.080 AN | 114 | 88 | 63 | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Difluorobenzene [IS2] | 6.18 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9685 | 2.0519 | 4.9627 | 10.4203 | 24.4346 | 100.1005 | 96.85 | 102.595 | 99.254 | 104.203 | 97.7384 | 100.1005 | Trichloroethene | 6.4 | Lin | 6 | 1.3426 | 0.99993 | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9491 | 2.009 | 4.7099 | 10.5924 | 25.0001 | 99.9556 | 94.91 | 100.45 | 94.198 | 105.924 | 100.0004 | 99.9556 | 1,2-Dichloropropane | 6.66 | Lin | 6 | 1.2424 | 0.99994 | TRUE | TRUE
Dibromomethane | 6.729 | 0.060 AN | 174 | 93 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0103 | 2.1685 | 4.9606 | 11.3671 | 25.9621 | 99.6213 | 101.03 | 108.425 | 99.212 | 113.671 | 103.8484 | 99.6213 | Dibromomethane | 6.74 | Lin | 6 | 3.1862 | 0.99959 | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2789 | 2.0536 | 4.4133 | 10.2477 | 25.0067 | 99.9988 | 127.89 | 102.68 | 88.266 | 102.477 | 100.0268 | 99.9988 | Methyl methacrylate | 6.76 | Quad, WithOffset | 6 | 1.5322 | 0.99995 | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8498 | 1.7659 | 4.2879 | 9.8238 | 23.8134 | 100.356 | 84.98 | 88.295 | 85.758 | 98.238 | 95.2536 | 100.356 | Bromodichloromethane | 6.93 | Lin | 6 | 2.7926 | 0.99971 | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | 1.4934 | 4.0466 | 10.4903 | 25.0428 | 99.9966 | #VALUE! | 74.67 | 80.932 | 104.903 | 100.1712 | 99.9966 | 2-Nitropropane | n.a. | Quad | 5 | 1.9787 | 0.99988 | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8414 | 1.7428 | 4.3066 | 10.3194 | 25.0503 | 99.9965 | 84.14 | 87.14 | 86.132 | 103.194 | 100.2012 | 99.9965 | cis-1,3-Dichloropropene | 7.37 | Quad | 6 | 1.5484 | 0.99993 | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.532 | 0.049 AN | 43 | 58 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 1.7551 | 3.8527 | 9.1384 | 21.2446 | 49.8888 | 199.9518 | 87.755 | 96.3175 | 91.384 | 106.223 | 99.7776 | 99.9759 | 4-Methyl-2-pentanone (MIBK) | 7.54 | Lin | 6 | 1.4506 | 0.99992 | TRUE | TRUE
Toluene-d8 [SS2] | 7.62 | 0.058 AN | 98 | 100 | 70 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20.0118 | 19.0368 | 19.3913 | 21.058 | 19.8253 | 20.6769 | 100.059 | 95.184 | 96.9565 | 105.29 | 99.1265 | 103.3845 | Toluene-d8 [SS2] | 7.62 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Toluene | 7.684 | 0.047 AN | 91 | 92 | 65 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8689 | 1.8252 | 4.3836 | 9.3843 | 22.476 | 100.7282 | 86.89 | 91.26 | 87.672 | 93.843 | 89.904 | 100.7282 | Toluene | 7.69 | Lin | 6 | 5.3337 | 0.99897 | TRUE | TRUE
trans-1,3-Dichloropropene | 7.931 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8221 | 1.6319 | 4.1081 | 10.0255 | 25.231 | 99.9911 | 82.21 | 81.595 | 82.162 | 100.255 | 100.924 | 99.9911 | trans-1,3-Dichloropropene | 7.94 | Quad | 6 | 1.7771 | 0.99992 | TRUE | TRUE
Ethyl methacrylate | 8 | 0.060 AN | 69 | 41 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.6621 | 1.4329 | 3.8329 | 9.6021 | 25.5198 | 99.9801 | 66.21 | 71.645 | 76.658 | 96.021 | 102.0792 | 99.9801 | Ethyl methacrylate | 8.01 | Quad | 6 | 2.7951 | 0.99979 | TRUE | TRUE
1,1,2-Trichloroethane | 8.11 | 0.060 AN | 97 | 83 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.987 | 1.9515 | 4.5416 | 10.3129 | 24.7465 | 100.0561 | 98.7 | 97.575 | 90.832 | 103.129 | 98.986 | 100.0561 | 1,1,2-Trichloroethane | 8.11 | Lin | 6 | 1.1567 | 0.99995 | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1359 | 2.3497 | 5.2389 | 10.6718 | 24.5535 | 100.0257 | 113.59 | 117.485 | 104.778 | 106.718 | 98.214 | 100.0257 | Tetrachloroethene | 8.17 | Quad | 6 | 2.0658 | 0.99985 | TRUE | TRUE
1,3-Dichloropropane | 8.25 | 0.060 AN | 76 | 41 | 78 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9404 | 1.9683 | 4.5223 | 10.4932 | 24.4345 | 100.1172 | 94.04 | 98.415 | 90.446 | 104.932 | 97.738 | 100.1172 | 1,3-Dichloropropane | 8.25 | Lin | 6 | 1.6944 | 0.99989 | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 2 | 4 | 10 | 20 | 50 | 200 | ppb | 1.7639 | 3.661 | 9.1126 | 21.7258 | 50.7762 | 199.6869 | 88.195 | 91.525 | 91.126 | 108.629 | 101.5524 | 99.8435 | 2-Hexanone | 8.33 | Lin | 6 | 2.0152 | 0.99984 | TRUE | TRUE
Dibromochloromethane | 8.435 | 0.050 AN | 129 | 127 | 131 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.708 | 1.455 | 3.6103 | 8.5456 | 22.4116 | 100.8759 | 70.8 | 72.75 | 72.206 | 85.456 | 89.6464 | 100.8759 | Dibromochloromethane | 8.44 | Lin | 6 | 6.7255 | 0.99844 | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.52 | 0.049 AN | 107 | 109 | 93 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9428 | 1.9993 | 4.6241 | 10.829 | 25.5685 | 99.7944 | 94.28 | 99.965 | 92.482 | 108.29 | 102.274 | 99.7944 | 1,2-Dibromoethane (EDB) | 8.53 | Lin | 6 | 2.0423 | 0.99984 | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.92 | 0.040 AF | 117 | 82 | 52 | 0.1 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Chlorobenzene-d5 [IS3] | 8.92 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Chlorobenzene | 8.94 | 0.040 AN | 112 | 77 | 114 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9643 | 1.9969 | 4.6566 | 10.1569 | 23.9956 | 100.253 | 96.43 | 99.845 | 93.132 | 101.569 | 95.9824 | 100.253 | Chlorobenzene | 8.95 | Lin | 6 | 2.084 | 0.99984 | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.02 | 0.040 AN | 131 | 133 | 117 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7767 | 1.5766 | 3.9386 | 8.9478 | 22.4018 | 100.8185 | 77.67 | 78.83 | 78.772 | 89.478 | 89.6072 | 100.8185 | 1,1,1,2-Tetrachloroethane | 9.02 | Lin | 6 | 6.0925 | 0.99869 | TRUE | TRUE
Ethylbenzene | 9.03 | 0.040 AN | 91 | 106 | 51 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9814 | 2.0592 | 4.8513 | 10.0724 | 23.9817 | 100.2538 | 98.14 | 102.96 | 97.026 | 100.724 | 95.9268 | 100.2538 | Ethylbenzene | 9.03 | Lin | 6 | 2.0082 | 0.99985 | TRUE | TRUE
m/p-Xylene | 9.138 | 0.060 AN | 91 | 106 | 105 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9758 | 2.0394 | 4.7933 | 10.2117 | 24.634 | 100.0801 | 97.58 | 101.97 | 95.866 | 102.117 | 98.536 | 100.0801 | m/p-Xylene | 9.14 | Lin | 6 | 0.9017 | 0.99997 | TRUE | TRUE
o-Xylene | 9.435 | 0.060 AN | 91 | 106 | 105 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9965 | 2.0396 | 4.7297 | 10.2096 | 24.3541 | 100.1533 | 99.65 | 101.98 | 94.594 | 102.096 | 97.4164 | 100.1533 | o-Xylene | 9.44 | Lin | 6 | 1.4084 | 0.99992 | TRUE | TRUE
Styrene | 9.451 | 0.060 AN | 104 | 78 | 103 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9297 | 1.9854 | 4.6446 | 10.2216 | 23.8982 | 100.2721 | 92.97 | 99.27 | 92.892 | 102.216 | 95.5928 | 100.2721 | Styrene | 9.46 | Lin | 6 | 2.2908 | 0.9998 | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8699 | 1.5988 | 4.1044 | 10.8653 | 24.862 | 100.0019 | 86.99 | 79.94 | 82.088 | 108.653 | 99.448 | 100.0019 | Bromoform | 9.59 | Quad | 6 | 2.2942 | 0.99986 | TRUE | TRUE
Isopropylbenzene (cumene) | 9.72 | 0.060 AN | 105 | 120 | 79 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8788 | 1.9269 | 4.5113 | 10.1787 | 22.4191 | 100.6545 | 87.88 | 96.345 | 90.226 | 101.787 | 89.6764 | 100.6545 | Isopropylbenzene (cumene) | 9.72 | Lin | 6 | 5.1858 | 0.99901 | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.843 | 0.060 AN | 95 | 174 | 176 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 19.9658 | 19.1828 | 19.6126 | 21.2255 | 19.2257 | 20.7876 | 99.829 | 95.914 | 98.063 | 106.1275 | 96.1285 | 103.938 | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.85 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
Bromobenzene | 9.943 | 0.040 AN | 77 | 156 | 158 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9879 | 1.9711 | 4.5971 | 11.1131 | 23.0216 | 100.4041 | 98.79 | 98.555 | 91.942 | 111.131 | 92.0864 | 100.4041 | Bromobenzene | 9.95 | Lin | 6 | 4.4204 | 0.99926 | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.954 | 0.040 AN | 83 | 85 | 95 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8609 | 1.7334 | 4.4057 | 11.3207 | 24.2698 | 100.0869 | 86.09 | 86.67 | 88.114 | 113.207 | 97.0792 | 100.0869 | 1,1,2,2-Tetrachloroethane | 9.96 | Lin | 6 | 3.1065 | 0.99963 | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.988 | 0.040 AN | 77 | 110 | 61 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9231 | 1.8313 | 4.4621 | 11.1843 | 23.3608 | 100.3224 | 92.31 | 91.565 | 89.242 | 111.843 | 93.4432 | 100.3224 | 1,2,3-Trichloropropane (TCP) | 9.99 | Lin | 6 | 4.0138 | 0.99939 | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.988 | 0.040 AN | 75 | 53 | 89 | 0.05 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8723 | 1.7555 | 4.1778 | 10.8429 | 23.5128 | 100.3348 | 87.23 | 87.775 | 83.556 | 108.429 | 94.0512 | 100.3348 | trans-1,4-Dichloro-2-butene | 9.99 | Lin | 6 | 3.6901 | 0.99949 | TRUE | TRUE
n-Propylbenzene | 10.023 | 0.026 AN | 91 | 120 | 65 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8756 | 1.8681 | 4.4623 | 10.4567 | 22.3635 | 100.6442 | 87.56 | 93.405 | 89.246 | 104.567 | 89.454 | 100.6442 | n-Propylbenzene | 10.03 | Lin | 6 | 5.3603 | 0.99894 | TRUE | TRUE
2-Chlorotoluene | 10.08 | 0.031 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9097 | 1.9241 | 4.5165 | 10.8012 | 22.6232 | 100.5407 | 90.97 | 96.205 | 90.33 | 108.012 | 90.4928 | 100.5407 | 2-Chlorotoluene | 10.09 | Lin | 6 | 4.9628 | 0.99908 | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.159 | 0.040 AN | 105 | 120 | 119 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8409 | 1.8651 | 4.3821 | 10.4792 | 22.3285 | 100.6552 | 84.09 | 93.255 | 87.642 | 104.792 | 89.314 | 100.6552 | 1,3,5-Trimethylbenzene | 10.16 | Lin | 6 | 5.4739 | 0.9989 | TRUE | TRUE
4-Chlorotoluene | 10.176 | 0.040 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9351 | 2.012 | 4.8482 | 11.5174 | 24.6022 | 99.9557 | 93.51 | 100.6 | 96.964 | 115.174 | 98.4088 | 99.9557 | 4-Chlorotoluene | 10.18 | Lin | 6 | 2.9431 | 0.99966 | TRUE | TRUE
tert-Butylbenzene | 10.38 | 0.040 AN | 119 | 91 | 134 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8222 | 1.7745 | 4.3127 | 10.1284 | 21.9625 | 100.7872 | 82.22 | 88.725 | 86.254 | 101.284 | 87.85 | 100.7872 | tert-Butylbenzene | 10.38 | Lin | 6 | 6.1955 | 0.99861 | TRUE | TRUE
Pentachloroethane | 10.398 | 0.030 AN | 167 | 130 | 132 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.4528 | 2.0206 | 4.1755 | 10.3395 | 25.0081 | 99.999 | 145.28 | 101.03 | 83.51 | 103.395 | 100.0324 | 99.999 | Pentachloroethane | 10.4 | Quad, WithOffset | 6 | 1.7883 | 0.99994 | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.422 | 0.030 AN | 105 | 120 | 77 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8513 | 1.8306 | 4.3641 | 10.5819 | 22.2574 | 100.6641 | 85.13 | 91.53 | 87.282 | 105.819 | 89.0296 | 100.6641 | 1,2,4-Trimethylbenzene | 10.42 | Lin | 6 | 5.6495 | 0.99883 | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.535 | 0.037 AN | 105 | 134 | 91 | 0.08 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8421 | 1.8248 | 4.4599 | 10.2883 | 22.5045 | 100.6271 | 84.21 | 91.24 | 89.198 | 102.883 | 90.018 | 100.6271 | 1-Methylpropylbenzene (sec-butylbenzene) | 10.54 | Lin | 6 | 5.0697 | 0.99906 | TRUE | TRUE
1,3-Dichlorobenzene | 10.61 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9575 | 1.9597 | 4.6768 | 11.3461 | 23.7942 | 100.1842 | 95.75 | 97.985 | 93.536 | 113.461 | 95.1768 | 100.1842 | 1,3-Dichlorobenzene | 10.62 | Lin | 6 | 3.4659 | 0.99954 | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8668 | 1.8584 | 4.5655 | 10.6046 | 23.0488 | 100.4532 | 86.68 | 92.92 | 91.31 | 106.046 | 92.1952 | 100.4532 | 4-Isopropyltoluene (p-cymene) | 10.64 | Lin | 6 | 4.0722 | 0.99938 | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 0.036 AN | 152 | 115 |  | 0.07 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | Lin, Avg | 1 | n.a. | n.a. | TRUE | TRUE
1,4-Dichlorobenzene | 10.687 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9599 | 2.0227 | 4.9069 | 11.7356 | 24.3221 | 100.0005 | 95.99 | 101.135 | 98.138 | 117.356 | 97.2884 | 100.0005 | 1,4-Dichlorobenzene | 10.69 | Lin | 6 | 3.4787 | 0.99952 | TRUE | TRUE
n-Butylbenzene | 10.923 | 0.040 AN | 91 | 92 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8159 | 1.7765 | 4.3292 | 9.8988 | 22.5435 | 100.6641 | 81.59 | 88.825 | 86.584 | 98.988 | 90.174 | 100.6641 | n-Butylbenzene | 10.92 | Lin | 6 | 5.0769 | 0.99906 | TRUE | TRUE
1,2-Dichlorobenzene | 10.93 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0271 | 2.0863 | 5.0349 | 12.3767 | 25.2928 | 99.6854 | 102.71 | 104.315 | 100.698 | 123.767 | 101.1712 | 99.6854 | 1,2-Dichlorobenzene | 10.93 | Lin | 6 | 4.4576 | 0.99919 | TRUE | TRUE
Hexachloroethane | 11.11 | 0.040 AN | 117 | 119 | 201 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7103 | 1.7508 | 4.4535 | 10.9077 | 24.7651 | 100.0049 | 71.03 | 87.54 | 89.07 | 109.077 | 99.0604 | 100.0049 | Hexachloroethane | 11.11 | Quad | 6 | 1.8962 | 0.99991 | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 0.040 AN | 157 | 155 | 75 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8045 | 1.6148 | 4.0874 | 10.6952 | 24.9438 | 99.9994 | 80.45 | 80.74 | 81.748 | 106.952 | 99.7752 | 99.9994 | 1,2-Dibromo-3-chloropropane (DBCP) | 11.46 | Quad | 6 | 2.266 | 0.99986 | TRUE | TRUE
Nitrobenzene | 11.58 | 0.040 AN | 77 | 51 | 123 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | n.a./n.r. | 1.9366 | 3.9389 | 10.9864 | 24.81 | 100.0035 | #VALUE! | 96.83 | 78.778 | 109.864 | 99.24 | 100.0035 | Nitrobenzene | n.a. | Quad | 5 | 2.412 | 0.99983 | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.98 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9297 | 1.8901 | 4.6351 | 11.4471 | 23.5282 | 100.2444 | 92.97 | 94.505 | 92.702 | 114.471 | 94.1128 | 100.2444 | 1,2,4-Trichlorobenzene | 11.99 | Lin | 6 | 3.9763 | 0.99939 | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.07 | 0.040 AN | 225 | 227 | 223 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8096 | 1.8107 | 4.3569 | 10.3695 | 22.4504 | 100.6383 | 80.96 | 90.535 | 87.138 | 103.695 | 89.8016 | 100.6383 | Hexachloro-1,3-butadiene | 12.08 | Lin | 6 | 5.243 | 0.99899 | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8399 | 1.7106 | 4.2653 | 11.0177 | 23.378 | 100.3478 | 83.99 | 85.53 | 85.306 | 110.177 | 93.512 | 100.3478 | Naphthalene | 12.16 | Lin | 6 | 3.9926 | 0.9994 | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.29 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9097 | 1.9418 | 4.718 | 11.5904 | 24.486 | 99.9856 | 90.97 | 97.09 | 94.36 | 115.904 | 97.944 | 99.9856 | 1,2,3-Trichlorobenzene | 12.3 | Lin | 6 | 3.1731 | 0.99961 | TRUE | TRUE

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 2
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 2
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 2
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | 20
Pentafluorobenzene [IS1] | 20
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | 20
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 2
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 2
Toluene-d8 [SS2] | 20
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 2
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | 20
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 1
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | 20
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | 20
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 2
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.47 | 352 | 0.02 | 0.054 | Not confirmed | 50 | 52 | 32.96 | n.a. | Not confirmed | 49 | 9.91 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | 1.57 | 105 | 0.01 | 0.013 | Not confirmed | 62 | 64 | 30.99 | n.a. | Not confirmed | 61 | 7.69 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.85 | 1286 | 0.08 | n.a. | Confirmed | 94 | 96 | 93.92 | 108.69 | Confirmed | 93 | 19.56 | 15.74 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 31.93 | n.a. | Not confirmed | 49 | 20.34 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 101 | 103 | 63.27 | n.a. | Not confirmed | 105 | 9.45 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 77.49 | n.a. | Not confirmed | 45 | 67.73 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 61 | 96 | 63.3 | n.a. | Not confirmed | 98 | 39.27 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 38.26 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Iodomethane | 2.91 | 588 | 0.04 | 0.11 | Confirmed | 142 | 127 | 29.9 | 30 | Confirmed | 141 | 13.59 | 16.26 | Confirmed
1 | TRUE | TRUE | FALSE | Carbon disulfide | 2.97 | 1141 | 0.07 | 0.063 | Not confirmed | 76 | 78 | 8.59 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 55.68 | n.a. | Not confirmed | 76 | 43.16 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Methylene chloride (DCM) | 3.36 | 507 | 0.03 | 0.069 | Confirmed | 49 | 84 | 93.47 | 71.77 | Confirmed | 86 | 58.49 | 58.63 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,2-Dichloroethene | 3.69 | 149 | 0.01 | 0.018 | Not confirmed | 61 | 96 | 66.42 | n.a. | Not confirmed | 98 | 41.02 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 28.43 | n.a. | Not confirmed | 57 | 22.43 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 65 | 32.05 | n.a. | Not confirmed | 83 | 11.82 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 79.1 | n.a. | Not confirmed | 79 | 32.28 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,2-Dichloroethene | 4.84 | 143 | 0.01 | 0.015 | Not confirmed | 61 | 96 | 69.52 | n.a. | Not confirmed | 98 | 42.6 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 29.99 | n.a. | Not confirmed | 57 | 8.44 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 16.02 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 30.15 | n.a. | Not confirmed | 40 | 33.08 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | Bromochloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 49 | 130 | 66.93 | n.a. | Not confirmed | 128 | 51.66 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 43.97 | n.a. | Not confirmed | 71 | 46.41 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 64.37 | n.a. | Not confirmed | 47 | 17.68 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 99 | 62.62 | n.a. | Not confirmed | 61 | 50.04 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.37 | 81453 | 5.01 | 20.246 | Confirmed | 113 | 111 | 103.67 | 102.13 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.44 | 140346 | 8.63 | 20 | Confirmed | 168 | 99 | 72.03 | 73.32 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 56 | 41 | 53.61 | n.a. | Not confirmed | 43 | 25.43 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon tetrachloride | n.a. | n.a. | n.a. | n.a. | Not confirmed | 119 | 121 | 30.81 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,1-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 77 | 31.3 | n.a. | Not confirmed | 110 | 32.93 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Benzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 78 | 77 | 26 | n.a. | Not confirmed | 52 | 14.41 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 32.31 | n.a. | Not confirmed | 49 | 28.7 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.18 | 266720 | 16.4 | 20 | Confirmed | 114 | 88 | 22.56 | 22.97 | Confirmed | 63 | 23.48 | 23.1 | Confirmed
1 | TRUE | TRUE | FALSE | Trichloroethene | 6.41 | 86 | 0.01 | 0.017 | Not confirmed | 130 | 132 | 95.16 | 98.27 | Confirmed | 95 | 119.47 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 68.23 | n.a. | Not confirmed | 41 | 37.13 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 103.16 | n.a. | Not confirmed | 39 | 44.55 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.74 | 67 | 0 | 0.025 | Not confirmed | 174 | 93 | 134.18 | n.a. | Not confirmed | 95 | 113.69 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 63.35 | n.a. | Not confirmed | 47 | 15.24 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 83.22 | n.a. | Not confirmed | 39 | 28.51 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 39 | 36.99 | n.a. | Not confirmed | 77 | 31.18 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 41.42 | n.a. | Not confirmed | 41 | 22.29 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.62 | 403647 | 24.82 | 19.705 | Confirmed | 98 | 100 | 57.21 | 57.99 | Confirmed | 70 | 12.49 | 12.8 | Confirmed
1 | TRUE | TRUE | FALSE | Toluene | 7.68 | 663 | 0.04 | 0.023 | Not confirmed | 91 | 92 | 52.14 | 48.23 | Confirmed | 65 | 12.46 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | trans-1,3-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 39 | 37.28 | n.a. | Not confirmed | 77 | 30.85 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 69 | 41 | 53.73 | n.a. | Not confirmed | 99 | 21.17 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 83 | 95.21 | n.a. | Not confirmed | 99 | 61.82 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.17 | 116 | 0.01 | 0.02 | Confirmed | 166 | 164 | 78.86 | 67.88 | Confirmed | 129 | 87.86 | 97.89 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 76 | 41 | 52.69 | n.a. | Not confirmed | 78 | 31.81 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 59.02 | n.a. | Not confirmed | 57 | 20.51 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 129 | 127 | 79.05 | n.a. | Not confirmed | 131 | 23.9 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 93.79 | n.a. | Not confirmed | 93 | 4.63 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.92 | 247299 | 15.2 | 20 | Confirmed | 117 | 82 | 75.82 | 75.93 | Confirmed | 52 | 17.58 | 17.04 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.95 | 333 | 0.02 | 0.023 | Not confirmed | 112 | 77 | 82.16 | 640.52 | Not confirmed | 114 | 31.02 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 131 | 133 | 95.8 | n.a. | Not confirmed | 117 | 87.39 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethylbenzene | 9.03 | 824 | 0.05 | 0.03 | Not confirmed | 91 | 106 | 30.02 | 25.16 | Confirmed | 51 | 9.05 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | m/p-Xylene | 9.14 | 1751 | 0.11 | 0.038 | Not confirmed | 91 | 106 | 43.16 | 40.31 | Confirmed | 105 | 19.77 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | o-Xylene | 9.44 | 877 | 0.05 | 0.036 | Not confirmed | 91 | 106 | 41.21 | 30.54 | Confirmed | 105 | 22.64 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.45 | 507 | 0.03 | 0.03 | Confirmed | 104 | 78 | 67.4 | 66.74 | Confirmed | 103 | 59.05 | 42.6 | Confirmed
1 | TRUE | TRUE | FALSE | Bromoform | 9.58 | 19 | 0 | 0.011 | Not confirmed | 173 | 171 | 52.49 | n.a. | Not confirmed | 175 | 48.09 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Isopropylbenzene (cumene) | 9.72 | 484 | 0.03 | 0.019 | Not confirmed | 105 | 120 | 25.71 | 20.36 | Confirmed | 79 | 19.16 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.85 | 155333 | 9.55 | 20.058 | Confirmed | 95 | 174 | 50.37 | 51.1 | Confirmed | 176 | 47.8 | 49.32 | Confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.95 | 383 | 0.02 | 0.03 | Confirmed | 77 | 156 | 39.29 | 47.87 | Confirmed | 158 | 38.32 | 43.57 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 63.98 | n.a. | Not confirmed | 95 | 12.16 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | trans-1,4-Dichloro-2-butene | 9.99 | 133 | 0.01 | 0.016 | Not confirmed | 75 | 53 | 18.84 | n.a. | Not confirmed | 89 | 7.56 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 110 | 67.29 | n.a. | Not confirmed | 61 | 62.83 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | n-Propylbenzene | 10.02 | 1101 | 0.07 | 0.033 | Not confirmed | 91 | 120 | 20.55 | 19.48 | Confirmed | 65 | 11.42 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2-Chlorotoluene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 91 | 126 | 27.9 | n.a. | Not confirmed | 89 | 18.25 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,3,5-Trimethylbenzene | 10.16 | 556 | 0.03 | 0.023 | Not confirmed | 105 | 120 | 41.07 | 34.63 | Confirmed | 119 | 10.74 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 4-Chlorotoluene | 10.18 | 908 | 0.06 | 0.041 | Not confirmed | 91 | 126 | 24.92 | 19.55 | Confirmed | 89 | 12.03 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | tert-Butylbenzene | 10.38 | 319 | 0.02 | 0.016 | Not confirmed | 119 | 91 | 84.03 | n.a. | Not confirmed | 134 | 21.74 | 22.97 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 167 | 130 | 86.52 | n.a. | Not confirmed | 132 | 91.7 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,4-Trimethylbenzene | 10.42 | 708 | 0.04 | 0.029 | Not confirmed | 105 | 120 | 39.32 | 36.94 | Confirmed | 77 | 13.41 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 759 | 0.05 | 0.026 | Not confirmed | 105 | 134 | 17.36 | 19.26 | Confirmed | 91 | 16.61 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.62 | 471 | 0.03 | 0.047 | Confirmed | 146 | 148 | 62.72 | 64.85 | Confirmed | 111 | 54.91 | 58.62 | Confirmed
1 | TRUE | TRUE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.64 | 607 | 0.04 | 0.028 | Not confirmed | 119 | 91 | 33.61 | 78.15 | Not confirmed | 134 | 25.92 | 24.41 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 113091 | 6.95 | 20 | Confirmed | 152 | 115 | 80.72 | 79.31 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.69 | 598 | 0.04 | 0.062 | Not confirmed | 146 | 148 | 63.83 | 100.72 | Not confirmed | 111 | 59.4 | 295.7 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.92 | 1263 | 0.08 | 0.049 | Confirmed | 91 | 92 | 51.09 | 57.89 | Confirmed | 134 | 20.84 | 23.34 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.93 | 480 | 0.03 | 0.052 | Confirmed | 146 | 148 | 62.49 | 65.83 | Confirmed | 111 | 55.09 | 57.87 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 96.74 | n.a. | Not confirmed | 201 | 63.53 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 157 | 155 | 75.33 | n.a. | Not confirmed | 75 | 138.48 | n.a. | Not confirmed
2 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 47.09 | n.a. | Not confirmed | 123 | 32.69 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.99 | 626 | 0.04 | 0.102 | Confirmed | 180 | 182 | 95.01 | 99.13 | Confirmed | 145 | 40.14 | 32.27 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.08 | 137 | 0.01 | 0.053 | Confirmed | 225 | 227 | 64.41 | 56.69 | Confirmed | 223 | 62.91 | 55.19 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.16 | 1998 | 0.12 | 0.089 | Confirmed | 128 | 127 | 13.76 | 13.02 | Confirmed | 129 | 10.59 | 8.77 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.3 | 575 | 0.04 | 0.097 | Confirmed | 180 | 182 | 96.39 | 89.21 | Confirmed | 145 | 41.51 | 41.1 | Confirmed

## CCV
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? |  |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.44 | 140428 | 5.44 | 133044 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.18 | 271106 | 6.18 | 263138 | TRUE | TRUE | 82.09 | 10 | TRUE | Chloromethane (methyl chloride) | 1.48 | 53749 | 0.33 | 8.209 | Confirmed | 50 | 52 | 32.96 | 32.53 | Confirmed | 49 | 9.91 | 9.77 | Confirmed
Chlorobenzene-d5 [IS3] | 8.92 | 257992 | 8.92 | 236091 | TRUE | TRUE | 75.6 | 10 | FALSE | Chloroethene (vinyl chloride) | 1.57 | 62935 | 0.39 | 7.56 | Confirmed | 62 | 64 | 30.99 | 33.04 | Confirmed | 61 | 7.69 | 8.53 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 123557 | 10.67 | 100630 | TRUE | TRUE | 117.57 | 10 | TRUE | Bromomethane (methyl bromide) | 1.85 | 73292 | 0.46 | 11.757 | Confirmed | 94 | 96 | 93.92 | 88.62 | Confirmed | 93 | 19.56 | 19.31 | Confirmed
[7 empty] | 84.45 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.96 | 45314 | 0.28 | 8.445 | Confirmed | 64 | 66 | 31.93 | 32.85 | Confirmed | 49 | 20.34 | 21.33 | Confirmed
[7 empty] | 81.2 | 10 | TRUE | Trichlorofluoromethane | 2.2 | 54745 | 0.34 | 8.12 | Confirmed | 101 | 103 | 63.27 | 63.73 | Confirmed | 105 | 9.45 | 8.72 | Confirmed
Total Analytes | 81 | [5 empty] | 89.69 | 10 | TRUE | Diethyl ether | 2.52 | 47078 | 0.29 | 8.969 | Confirmed | 59 | 74 | 77.49 | 77.58 | Confirmed | 45 | 67.73 | 68.95 | Confirmed
Failed | 21 | [5 empty] | 78.49 | 10 | FALSE | 1,1-Dichloroethene | 2.75 | 62834 | 0.39 | 7.849 | Confirmed | 61 | 96 | 63.3 | 66.39 | Confirmed | 98 | 39.27 | 42.69 | Confirmed
[7 empty] | 107.7056 | 18 | TRUE | Acetone | 2.84 | 32097 | 0.2 | 19.387 | Confirmed | 43 | 58 | 38.26 | 42.72 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 46.87 | 10 | FALSE | Iodomethane | 2.91 | 25462 | 0.16 | 4.687 | Confirmed | 142 | 127 | 29.9 | 32.37 | Confirmed | 141 | 13.59 | 10.58 | Confirmed
[7 empty] | 74.5 | 10 | FALSE | Carbon disulfide | 2.97 | 135168 | 0.84 | 7.45 | Confirmed | 76 | 78 | 8.59 | 9.26 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 83.67 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.21 | 59809 | 0.37 | 8.367 | Confirmed | 41 | 39 | 55.68 | 56.63 | Confirmed | 76 | 43.16 | 42.56 | Confirmed
[7 empty] | 95.24 | 10 | TRUE | Methylene chloride (DCM) | 3.38 | 70422 | 0.44 | 9.524 | Confirmed | 49 | 84 | 93.47 | 94.35 | Confirmed | 86 | 58.49 | 58.72 | Confirmed
[7 empty] | 83.3 | 10 | TRUE | trans-1,2-Dichloroethene | 3.69 | 68318 | 0.42 | 8.33 | Confirmed | 61 | 96 | 66.42 | 67.28 | Confirmed | 98 | 41.02 | 42.23 | Confirmed
[7 empty] | 89.3 | 10 | TRUE | Methyl tert-butyl ether (MTBE) | 3.7 | 125047 | 0.78 | 8.93 | Confirmed | 73 | 41 | 28.43 | 28.03 | Confirmed | 57 | 22.43 | 22.38 | Confirmed
[7 empty] | 84.63 | 10 | TRUE | 1,1-Dichloroethane | 4.2 | 93071 | 0.58 | 8.463 | Confirmed | 63 | 65 | 32.05 | 32.11 | Confirmed | 83 | 11.82 | 12.72 | Confirmed
[7 empty] | 58.27 | 10 | FALSE | 2,2-Dichloropropane | 4.83 | 34938 | 0.22 | 5.827 | Confirmed | 77 | 41 | 79.1 | 79.57 | Confirmed | 79 | 32.28 | 31.46 | Confirmed
[7 empty] | 82.2 | 10 | TRUE | cis-1,2-Dichloroethene | 4.83 | 78084 | 0.48 | 8.22 | Confirmed | 61 | 96 | 69.52 | 73.23 | Confirmed | 98 | 42.6 | 45.95 | Confirmed
[7 empty] | 95.4667 | 18 | TRUE | 2-Butanone (MEK) | 4.85 | 48756 | 0.3 | 17.184 | Confirmed | 43 | 72 | 29.99 | 29.74 | Confirmed | 57 | 8.44 | 9.44 | Confirmed
[7 empty] | 92.05 | 10 | TRUE | Methyl acrylate | 4.95 | 41441 | 0.26 | 9.205 | Confirmed | 55 | 85 | 16.02 | 16.33 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 86.46 | 10 | TRUE | Methacrylonitrile | 5.07 | 28263 | 0.18 | 8.646 | Confirmed | 67 | 52 | 30.15 | 30.66 | Confirmed | 40 | 33.08 | 36.88 | Confirmed
[7 empty] | 91.07 | 10 | TRUE | Bromochloromethane | 5.08 | 46935 | 0.29 | 9.107 | Confirmed | 49 | 130 | 66.93 | 70.24 | Confirmed | 128 | 51.66 | 55.41 | Confirmed
[7 empty] | 90.14 | 10 | TRUE | Tetrahydrofuran | 5.1 | 18600 | 0.12 | 9.014 | Confirmed | 42 | 72 | 43.97 | 45.46 | Confirmed | 71 | 46.41 | 45.88 | Confirmed
[7 empty] | 92.74 | 10 | TRUE | Trichloromethane (chloroform) | 5.21 | 89402 | 0.56 | 9.274 | Confirmed | 83 | 85 | 64.37 | 63.53 | Confirmed | 47 | 17.68 | 17.53 | Confirmed
[7 empty] | 82.46 | 10 | TRUE | 1,1,1-Trichloroethane | 5.35 | 52041 | 0.32 | 8.246 | Confirmed | 97 | 99 | 62.62 | 62.76 | Confirmed | 61 | 50.04 | 49.41 | Confirmed
[7 empty] | 101.67 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.37 | 81854 | 0.51 | 20.334 | Confirmed | 113 | 111 | 103.67 | 103.49 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.44 | 140428 | 0.87 | 20 | Confirmed | 168 | 99 | 72.03 | 70.79 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 80.76 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.49 | 79001 | 0.49 | 8.076 | Confirmed | 56 | 41 | 53.61 | 54.17 | Confirmed | 43 | 25.43 | 24.65 | Confirmed
[7 empty] | 88.37 | 10 | TRUE | Carbon tetrachloride | 5.5 | 32435 | 0.2 | 8.837 | Confirmed | 119 | 121 | 30.81 | 31.36 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 81.03 | 10 | TRUE | 1,1-Dichloropropene | 5.52 | 61936 | 0.38 | 8.103 | Confirmed | 75 | 77 | 31.3 | 30.57 | Confirmed | 110 | 32.93 | 33.93 | Confirmed
[7 empty] | 79.75 | 10 | FALSE | Benzene | 5.71 | 220398 | 1.37 | 7.975 | Confirmed | 78 | 77 | 26 | 25.36 | Confirmed | 52 | 14.41 | 14.58 | Confirmed
[7 empty] | 96.03 | 10 | TRUE | 1,2-Dichloroethane | 5.79 | 67931 | 0.42 | 9.603 | Confirmed | 62 | 64 | 32.31 | 32.64 | Confirmed | 49 | 28.7 | 27.97 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.18 | 271106 | 1.68 | 20 | Confirmed | 114 | 88 | 22.56 | 22.28 | Confirmed | 63 | 23.48 | 22.67 | Confirmed
[7 empty] | 87.22 | 10 | TRUE | Trichloroethene | 6.4 | 44839 | 0.28 | 8.722 | Confirmed | 130 | 132 | 95.16 | 95.57 | Confirmed | 95 | 119.47 | 118.14 | Confirmed
[7 empty] | 86.57 | 10 | TRUE | 1,2-Dichloropropane | 6.65 | 59695 | 0.37 | 8.657 | Confirmed | 63 | 62 | 68.23 | 67.97 | Confirmed | 41 | 37.13 | 37.85 | Confirmed
[7 empty] | 102.3 | 10 | TRUE | Dibromomethane | 6.74 | 28048 | 0.17 | 10.23 | Confirmed | 174 | 93 | 134.18 | 131.08 | Confirmed | 95 | 113.69 | 109.72 | Confirmed
[7 empty] | 92.91 | 10 | TRUE | Methyl methacrylate | 6.75 | 35797 | 0.22 | 9.291 | Confirmed | 41 | 69 | 103.16 | 100.57 | Confirmed | 39 | 44.55 | 43.85 | Confirmed
[7 empty] | 83.19 | 10 | TRUE | Bromodichloromethane | 6.93 | 57480 | 0.36 | 8.319 | Confirmed | 83 | 85 | 63.35 | 61.59 | Confirmed | 47 | 15.24 | 15.35 | Confirmed
[7 empty] | 90.85 | 10 | TRUE | 2-Nitropropane | 7.16 | 6843 | 0.04 | 9.085 | Confirmed | 43 | 41 | 83.22 | 81.09 | Confirmed | 39 | 28.51 | 28.7 | Confirmed
[7 empty] | 77.56 | 10 | FALSE | cis-1,3-Dichloropropene | 7.37 | 63600 | 0.39 | 7.756 | Confirmed | 75 | 39 | 36.99 | 37.47 | Confirmed | 77 | 31.18 | 32.07 | Confirmed
[7 empty] | 101.2778 | 18 | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.53 | 117352 | 0.73 | 18.23 | Confirmed | 43 | 58 | 41.42 | 41.94 | Confirmed | 41 | 22.29 | 22.17 | Confirmed
[7 empty] | 96.995 | 20 | TRUE | Toluene-d8 [SS2] | 7.62 | 414553 | 2.57 | 19.399 | Confirmed | 98 | 100 | 57.21 | 56.97 | Confirmed | 70 | 12.49 | 12.05 | Confirmed
[7 empty] | 76.29 | 10 | FALSE | Toluene | 7.69 | 233085 | 1.45 | 7.629 | Confirmed | 91 | 92 | 52.14 | 52.66 | Confirmed | 65 | 12.46 | 12.12 | Confirmed
[7 empty] | 78.58 | 10 | FALSE | trans-1,3-Dichloropropene | 7.93 | 47705 | 0.3 | 7.858 | Confirmed | 75 | 39 | 37.28 | 38.26 | Confirmed | 77 | 30.85 | 30.99 | Confirmed
[7 empty] | 89.13 | 10 | TRUE | Ethyl methacrylate | 8 | 68220 | 0.42 | 8.913 | Confirmed | 69 | 41 | 53.73 | 53.73 | Confirmed | 99 | 21.17 | 20.7 | Confirmed
[7 empty] | 90.87 | 10 | TRUE | 1,1,2-Trichloroethane | 8.11 | 50752 | 0.32 | 9.087 | Confirmed | 97 | 83 | 95.21 | 92.93 | Confirmed | 99 | 61.82 | 61.41 | Confirmed
[7 empty] | 86.91 | 10 | TRUE | Tetrachloroethene | 8.16 | 52311 | 0.32 | 8.691 | Confirmed | 166 | 164 | 78.86 | 79.9 | Confirmed | 129 | 87.86 | 87.72 | Confirmed
[7 empty] | 86.7 | 10 | TRUE | 1,3-Dichloropropane | 8.25 | 90255 | 0.56 | 8.67 | Confirmed | 76 | 41 | 52.69 | 53.88 | Confirmed | 78 | 31.81 | 32.01 | Confirmed
[7 empty] | 101.9611 | 18 | TRUE | 2-Hexanone | 8.32 | 80084 | 0.5 | 18.353 | Confirmed | 43 | 58 | 59.02 | 59.75 | Confirmed | 57 | 20.51 | 21.62 | Confirmed
[7 empty] | 73.35 | 10 | FALSE | Dibromochloromethane | 8.44 | 30978 | 0.19 | 7.335 | Confirmed | 129 | 127 | 79.05 | 78.6 | Confirmed | 131 | 23.9 | 23.84 | Confirmed
[7 empty] | 95.57 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.53 | 46606 | 0.29 | 9.557 | Confirmed | 107 | 109 | 93.79 | 93.67 | Confirmed | 93 | 4.63 | 4.76 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.92 | 257992 | 1.6 | 20 | Confirmed | 117 | 82 | 75.82 | 73.11 | Confirmed | 52 | 17.58 | 16.86 | Confirmed
[7 empty] | 86.19 | 10 | TRUE | Chlorobenzene | 8.95 | 130146 | 0.81 | 8.619 | Confirmed | 112 | 77 | 82.16 | 76.39 | Confirmed | 114 | 31.02 | 31.16 | Confirmed
[7 empty] | 79.55 | 10 | FALSE | 1,1,1,2-Tetrachloroethane | 9.02 | 31421 | 0.2 | 7.955 | Confirmed | 131 | 133 | 95.8 | 95.66 | Confirmed | 117 | 87.39 | 84.62 | Confirmed
[7 empty] | 84.74 | 10 | TRUE | Ethylbenzene | 9.03 | 240155 | 1.49 | 8.474 | Confirmed | 91 | 106 | 30.02 | 30.96 | Confirmed | 51 | 9.05 | 9.18 | Confirmed
[7 empty] | 86.95 | 10 | TRUE | m/p-Xylene | 9.14 | 419278 | 2.6 | 8.695 | Confirmed | 91 | 106 | 43.16 | 43.54 | Confirmed | 105 | 19.77 | 20.29 | Confirmed
[7 empty] | 88.46 | 10 | TRUE | o-Xylene | 9.44 | 223515 | 1.39 | 8.846 | Confirmed | 91 | 106 | 41.21 | 41.94 | Confirmed | 105 | 22.64 | 22.76 | Confirmed
[7 empty] | 87.88 | 10 | TRUE | Styrene | 9.45 | 153584 | 0.95 | 8.788 | Confirmed | 104 | 78 | 67.4 | 65.07 | Confirmed | 103 | 59.05 | 59.33 | Confirmed
[7 empty] | 81.97 | 10 | TRUE | Bromoform | 9.58 | 15770 | 0.1 | 8.197 | Confirmed | 173 | 171 | 52.49 | 52.81 | Confirmed | 175 | 48.09 | 48.24 | Confirmed
[7 empty] | 79.13 | 10 | FALSE | Isopropylbenzene (cumene) | 9.72 | 221586 | 1.38 | 7.913 | Confirmed | 105 | 120 | 25.71 | 25.98 | Confirmed | 79 | 19.16 | 18.11 | Confirmed
[7 empty] | 93.765 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 158670 | 0.99 | 18.753 | Confirmed | 95 | 174 | 50.37 | 50.9 | Confirmed | 176 | 47.8 | 50.59 | Confirmed
[7 empty] | 80.95 | 10 | TRUE | Bromobenzene | 9.94 | 112189 | 0.7 | 8.095 | Confirmed | 77 | 156 | 39.29 | 43.08 | Confirmed | 158 | 38.32 | 41.54 | Confirmed
[7 empty] | 88.54 | 10 | TRUE | 1,1,2,2-Tetrachloroethane | 9.95 | 70125 | 0.44 | 8.854 | Confirmed | 83 | 85 | 63.98 | 62.3 | Confirmed | 95 | 12.16 | 12.3 | Confirmed
[7 empty] | 82.42 | 10 | TRUE | 1,2,3-Trichloropropane (TCP) | 9.99 | 24902 | 0.15 | 8.242 | Confirmed | 77 | 110 | 67.29 | 73.65 | Confirmed | 61 | 62.83 | 65.02 | Confirmed
[7 empty] | 79.68 | 10 | FALSE | trans-1,4-Dichloro-2-butene | 9.99 | 73661 | 0.46 | 7.968 | Confirmed | 75 | 53 | 18.84 | 18.85 | Confirmed | 89 | 7.56 | 7.42 | Confirmed
[7 empty] | 78.17 | 10 | FALSE | n-Propylbenzene | 10.02 | 286285 | 1.78 | 7.817 | Confirmed | 91 | 120 | 20.55 | 20.81 | Confirmed | 65 | 11.42 | 11.15 | Confirmed
[7 empty] | 81.76 | 10 | TRUE | 2-Chlorotoluene | 10.09 | 181690 | 1.13 | 8.176 | Confirmed | 91 | 126 | 27.9 | 28.89 | Confirmed | 89 | 18.25 | 18.6 | Confirmed
[7 empty] | 77.32 | 10 | FALSE | 1,3,5-Trimethylbenzene | 10.16 | 208461 | 1.29 | 7.732 | Confirmed | 105 | 120 | 41.07 | 41 | Confirmed | 119 | 10.74 | 10.68 | Confirmed
[7 empty] | 86.94 | 10 | TRUE | 4-Chlorotoluene | 10.17 | 212385 | 1.32 | 8.694 | Confirmed | 91 | 126 | 24.92 | 25.25 | Confirmed | 89 | 12.03 | 11.88 | Confirmed
[7 empty] | 79.29 | 10 | FALSE | tert-Butylbenzene | 10.38 | 167959 | 1.04 | 7.929 | Confirmed | 119 | 91 | 84.03 | 81.53 | Confirmed | 134 | 21.74 | 22.58 | Confirmed
[7 empty] | 96.52 | 10 | TRUE | Pentachloroethane | 10.4 | 11009 | 0.07 | 9.652 | Confirmed | 167 | 130 | 86.52 | 83.26 | Confirmed | 132 | 91.7 | 86.59 | Confirmed
[7 empty] | 78.34 | 10 | FALSE | 1,2,4-Trimethylbenzene | 10.42 | 212281 | 1.32 | 7.834 | Confirmed | 105 | 120 | 39.32 | 39.76 | Confirmed | 77 | 13.41 | 13.01 | Confirmed
[7 empty] | 76.97 | 10 | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 247501 | 1.54 | 7.697 | Confirmed | 105 | 134 | 17.36 | 17.6 | Confirmed | 91 | 16.61 | 16.91 | Confirmed
[7 empty] | 86.75 | 10 | TRUE | 1,3-Dichlorobenzene | 10.61 | 94511 | 0.59 | 8.675 | Confirmed | 146 | 148 | 62.72 | 62.9 | Confirmed | 111 | 54.91 | 54.58 | Confirmed
[7 empty] | 79.16 | 10 | FALSE | 4-Isopropyltoluene (p-cymene) | 10.64 | 187955 | 1.17 | 7.916 | Confirmed | 119 | 91 | 33.61 | 34.04 | Confirmed | 134 | 25.92 | 25.91 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 123557 | 0.77 | 20 | Confirmed | 152 | 115 | 80.72 | 68.32 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 89.94 | 10 | TRUE | 1,4-Dichlorobenzene | 10.69 | 95542 | 0.59 | 8.994 | Confirmed | 146 | 148 | 63.83 | 61.91 | Confirmed | 111 | 59.4 | 55.5 | Confirmed
[7 empty] | 71.5 | 10 | FALSE | n-Butylbenzene | 10.92 | 199789 | 1.24 | 7.15 | Confirmed | 91 | 92 | 51.09 | 51.22 | Confirmed | 134 | 20.84 | 21.91 | Confirmed
[7 empty] | 97.25 | 10 | TRUE | 1,2-Dichlorobenzene | 10.93 | 98399 | 0.61 | 9.725 | Confirmed | 146 | 148 | 62.49 | 62.05 | Confirmed | 111 | 55.09 | 54.36 | Confirmed
[7 empty] | 89.46 | 10 | TRUE | Hexachloroethane | 11.11 | 18945 | 0.12 | 8.946 | Confirmed | 117 | 119 | 96.74 | 93.87 | Confirmed | 201 | 63.53 | 61.48 | Confirmed
[7 empty] | 88.37 | 10 | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 9491 | 0.06 | 8.837 | Confirmed | 157 | 155 | 75.33 | 78.91 | Confirmed | 75 | 138.48 | 138.06 | Confirmed
[7 empty] | 83.1 | 10 | TRUE | Nitrobenzene | 11.58 | 2084 | 0.01 | 8.31 | Confirmed | 77 | 51 | 47.09 | 47.9 | Confirmed | 123 | 32.69 | 37.63 | Confirmed
[7 empty] | 84.16 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.98 | 56200 | 0.35 | 8.416 | Confirmed | 180 | 182 | 95.01 | 92.19 | Confirmed | 145 | 40.14 | 39.07 | Confirmed
[7 empty] | 75.36 | 10 | FALSE | Hexachloro-1,3-butadiene | 12.07 | 21252 | 0.13 | 7.536 | Confirmed | 225 | 227 | 64.41 | 64.63 | Confirmed | 223 | 62.91 | 61.64 | Confirmed
[7 empty] | 82.96 | 10 | TRUE | Naphthalene | 12.15 | 202872 | 1.26 | 8.296 | Confirmed | 128 | 127 | 13.76 | 13.35 | Confirmed | 129 | 10.59 | 10.17 | Confirmed
[7 empty] | 88.52 | 10 | TRUE | 1,2,3-Trichlorobenzene | 12.3 | 57465 | 0.36 | 8.852 | Confirmed | 180 | 182 | 96.39 | 94.21 | Confirmed | 145 | 41.51 | 40.15 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2024
240111
Peak Name | Amount
ppb
First Injection | H2O-R1 | H2O-R2 | H2O-R3 | H2O-R4 | H2O-R5 | H2O-R6 | H2O-R7 | H2O-V1 | Cal1-1 | Cal2-2 | CAL3-5 | CAL4-10 | CAL5-25 | CAL6-100 | H2O-V2 | H2O-V3 | MB | 240111-Holmes | 12-12-23-JI-01 | 12-12-23-JI-02 | 12-12-23-JI-05 | 12-12-23-JI-06 | 12-13-23-JI-01 | 12-13-23-JI-02 | 12-13-23-JI-03 | 12-14-23-JI-01 | 12-14-23-JI-02 | 12-28-23-JI-02 | 12-28-23-JI-03 | 12-28-23-JI-04 | 12-28-23-JI-05 | 12-28-23-JI-06 | 12-28-23-JI-07 | 12-29-23-JI-01 | 12-29-23-JI-02 | CCV
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9285 | 1.9751 | 4.9378 | 9.6078 | 24.0219 | 100.2881 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2089
Chloroethene (vinyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9106 | 2.0212 | 4.8554 | 9.2388 | 23.799 | 100.3841 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.6691 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.5603
Bromomethane (methyl bromide) | n.a./n.r. | n.a. | n.a. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7152 | 2.1232 | 5.3422 | 9.8184 | 25.0001 | 100.0024 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a./n.r. | n.a. | n.a./n.r. | n.a./n.r. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a./n.r. | n.a. | n.a. | 11.7574
Chloroethane (ethyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9539 | 1.8709 | 5.0914 | 9.8932 | 24.1723 | 100.2161 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4446
Trichlorofluoromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.876 | 1.9934 | 4.7725 | 8.9548 | 23.106 | 100.5908 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1202
Diethyl ether | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8747 | 2.0222 | 4.8497 | 10.0799 | 24.5725 | 100.1072 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9685
1,1-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.915 | 2.007 | 4.8052 | 9.2302 | 23.5769 | 100.4432 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8489
Acetone | 1.7343 | 5.3895 | 6.3148 | 5.4442 | 5.9804 | 7.2145 | 5.8855 | n.a./n.r. | 3.0082 | 3.6089 | 8.455 | 21.1057 | 49.8103 | 200.0118 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 79.0867 | 106.3395 | 50.0148 | 101.3358 | 77.4382 | 68.3729 | 91.5264 | 75.4725 | 74.0075 | 35.7211 | 3.3374 | n.a./n.r. | 8.3951 | 19.495 | 26.2868 | 25.8537 | 8.4975 | 19.3875
Iodomethane | 0.7408 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1127 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.5257 | 1.3331 | 3.9371 | 9.7826 | 27.646 | 99.6055 | 0.2493 | 0.1202 | 0.1102 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 4.6866
Carbon disulfide | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8818 | 1.9768 | 4.8249 | 9.1604 | 23.4187 | 100.4897 | 0.2846 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.4497
3-Chloropropene (allyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8446 | 1.7465 | 4.8572 | 9.979 | 25.0759 | 99.9966 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3671
Methylene chloride (DCM) | 0.201 | 0.1501 | 0.1436 | 0.1461 | 0.1333 | 0.1573 | 0.1538 | 0.0561 | 1.0824 | 2.278 | 5.1978 | 10.5829 | 25.4397 | 99.8155 | n.a./n.r. | 0.0813 | 0.0686 | 0.0809 | 0.1524 | 0.0669 | 0.0686 | 0.0673 | 0.0989 | 0.1369 | 0.1528 | n.a./n.r. | n.a./n.r. | 0.0724 | 0.0809 | n.a./n.r. | 0.0778 | 0.5139 | 0.0906 | 0.0654 | 0.0748 | 9.5243
trans-1,2-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9078 | 2.0811 | 4.8522 | 9.7005 | 23.8883 | 100.3146 | 0.0803 | 0.0414 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3305
Methyl tert-butyl ether (MTBE) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.892 | 1.9143 | 4.7177 | 10.0119 | 24.476 | 100.1467 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9296
1,1-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9235 | 2.0589 | 4.9214 | 10.0394 | 24.1706 | 100.2069 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4626
2,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7334 | 1.5822 | 4.0558 | 8.2344 | 21.6512 | 101.072 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.8272
cis-1,2-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9329 | 1.9558 | 4.8173 | 9.9154 | 23.8327 | 100.311 | 0.0564 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2197
2-Butanone (MEK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.8114 | 3.2409 | 8.5824 | 20.6024 | 50.3505 | 199.9401 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 3.005 | 6.8379 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 17.1835
Methyl acrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.391 | 2.0168 | 4.6648 | 9.7707 | 25.162 | 99.9943 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2055
Methacrylonitrile | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.7387 | 4.1869 | 9.7104 | 24.1651 | 100.2868 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.646
Bromochloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9231 | 1.9961 | 4.8018 | 10.1663 | 24.5075 | 100.1173 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0836 | 0.1623 | 0.1162 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1065
Tetrahydrofuran | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7857 | 1.9825 | 4.4234 | 9.9315 | 25.0109 | 100.0354 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.7213 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0139
Trichloromethane (chloroform) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9831 | 2.1441 | 5.0841 | 10.2909 | 24.6932 | 100.0407 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.1344 | 6.6138 | n.a./n.r. | 0.0583 | 15.159 | 11.6138 | 14.5194 | 0.3504 | 12.7838 | 46.6734 | n.a./n.r. | n.a./n.r. | 13.9962 | 0.1822 | 31.4676 | 74.8262 | 0.5273 | 9.2742
1,1,1-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8399 | 1.9045 | 4.5274 | 9.5429 | 23.82 | 100.3678 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2463
Dibromofluoromethane [SS1] | 22.7558 | 20.5657 | 20.8578 | 20.1807 | 19.886 | 19.9999 | 19.4302 | 19.5659 | 19.4305 | 19.0288 | 19.7158 | 20.2783 | 20.7682 | 20.7785 | 20.3119 | 20.6568 | 20.246 | 21.0827 | 20.3389 | 20.7271 | 20.2314 | 20.2124 | 20.1537 | 20.6859 | 20.1219 | 20.6846 | 20.6959 | 20.3695 | 20.7419 | 21.1752 | 20.3937 | 20.2272 | 20.2226 | 21.0451 | 19.9657 | 20.3338
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8775 | 2.0601 | 4.7813 | 9.4316 | 24.2419 | 100.2573 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0761
Carbon tetrachloride | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9989 | 2.1844 | 4.932 | 9.7969 | 25.0907 | 99.9971 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8372
1,1-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8712 | 2.0205 | 4.8255 | 9.4654 | 24.456 | 100.1991 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1034
Benzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8827 | 1.9145 | 4.4235 | 9.7561 | 23.7717 | 100.3632 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9748
1,2-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.039 | 2.0961 | 4.9429 | 11.0444 | 25.3265 | 99.8145 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6027
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9685 | 2.0519 | 4.9627 | 10.4203 | 24.4346 | 100.1005 | 0.0613 | 0.0274 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7218
1,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9491 | 2.009 | 4.7099 | 10.5924 | 25.0001 | 99.9556 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6575
Dibromomethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0103 | 2.1685 | 4.9606 | 11.3671 | 25.9621 | 99.6213 | 0.0711 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.2564 | 0.1193 | 1.0689 | 0.4418 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.2297
Methyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.2789 | 2.0536 | 4.4133 | 10.2477 | 25.0067 | 99.9988 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2914
Bromodichloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8498 | 1.7659 | 4.2879 | 9.8238 | 23.8134 | 100.356 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.2137 | 0.154 | n.a./n.r. | 5.8703 | 8.1751 | 4.6318 | 0.52 | 3.5263 | 14.2597 | n.a./n.r. | n.a./n.r. | 2.8244 | n.a./n.r. | 12.1955 | 11.549 | 0.1559 | 8.3194
2-Nitropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.4934 | 4.0466 | 10.4903 | 25.0428 | 99.9966 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0846
cis-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8414 | 1.7428 | 4.3066 | 10.3194 | 25.0503 | 99.9965 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7555
4-Methyl-2-pentanone (MIBK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.7551 | 3.8527 | 9.1384 | 21.2446 | 49.8888 | 199.9518 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 18.2297
Toluene-d8 [SS2] | 19.2666 | 18.9529 | 19.6705 | 19.752 | 19.3931 | 19.8146 | 19.6493 | 19.4242 | 20.0118 | 19.0368 | 19.3913 | 21.058 | 19.8253 | 20.6769 | 20.5465 | 20.6262 | 19.7053 | 19.2946 | 19.8128 | 20.1984 | 19.705 | 19.4306 | 19.7625 | 19.4275 | 19.3991 | 19.3176 | 19.1752 | 19.6842 | 19.2331 | 19.013 | 19.5479 | 19.2855 | 19.3565 | 19.3505 | 19.2986 | 19.3989
Toluene | n.a./n.r. | 0.0606 | 0.0739 | 0.0765 | 0.068 | n.a./n.r. | 0.0667 | n.a./n.r. | 0.8689 | 1.8252 | 4.3836 | 9.3843 | 22.476 | 100.7282 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0503 | 0.0684 | 0.0719 | 0.0531 | n.a./n.r. | n.a./n.r. | 0.0571 | n.a./n.r. | n.a./n.r. | 0.098 | 0.1064 | n.a./n.r. | 0.0442 | 0.3417 | n.a./n.r. | n.a./n.r. | 0.0385 | 7.6294
trans-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8221 | 1.6319 | 4.1081 | 10.0255 | 25.231 | 99.9911 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8582
Ethyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6621 | 1.4329 | 3.8329 | 9.6021 | 25.5198 | 99.9801 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9126
1,1,2-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.987 | 1.9515 | 4.5416 | 10.3129 | 24.7465 | 100.0561 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0872
Tetrachloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.1359 | 2.3497 | 5.2389 | 10.6718 | 24.5535 | 100.0257 | 0.0936 | 0.0285 | 0.0199 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6905
1,3-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9404 | 1.9683 | 4.5223 | 10.4932 | 24.4345 | 100.1172 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6698
2-Hexanone | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.7639 | 3.661 | 9.1126 | 21.7258 | 50.7762 | 199.6869 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.2015 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 18.3531
Dibromochloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.708 | 1.455 | 3.6103 | 8.5456 | 22.4116 | 100.8759 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 3.3696 | 0.534 | n.a./n.r. | 2.5863 | 5.1171 | 3.5132 | 1.6961 | 1.37 | 5.1963 | n.a./n.r. | n.a./n.r. | 0.7988 | n.a./n.r. | 5.423 | 1.7831 | n.a./n.r. | 7.3348
1,2-Dibromoethane (EDB) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9428 | 1.9993 | 4.6241 | 10.829 | 25.5685 | 99.7944 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5569
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9643 | 1.9969 | 4.6566 | 10.1569 | 23.9956 | 100.253 | 0.1069 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6186
1,1,1,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7767 | 1.5766 | 3.9386 | 8.9478 | 22.4018 | 100.8185 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9546
Ethylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9814 | 2.0592 | 4.8513 | 10.0724 | 23.9817 | 100.2538 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4737
m/p-Xylene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9758 | 2.0394 | 4.7933 | 10.2117 | 24.634 | 100.0801 | 0.0909 | 0.043 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.026 | n.a./n.r. | n.a./n.r. | 0.0515 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6947
o-Xylene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9965 | 2.0396 | 4.7297 | 10.2096 | 24.3541 | 100.1533 | 0.0883 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8463
Styrene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9297 | 1.9854 | 4.6446 | 10.2216 | 23.8982 | 100.2721 | 0.0826 | 0.0415 | 0.0302 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7885
Bromoform | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8699 | 1.5988 | 4.1044 | 10.8653 | 24.862 | 100.0019 | 0.0628 | n.a./n.r. | n.a./n.r. | 0.8465 | n.a./n.r. | 0.8888 | 1.3428 | n.a./n.r. | 1.0665 | 1.1941 | 9.6623 | 16.1471 | 0.3467 | 1.3992 | n.a./n.r. | 0.152 | 0.3783 | n.a./n.r. | 1.3856 | 0.1452 | n.a./n.r. | 8.1974
Isopropylbenzene (cumene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8788 | 1.9269 | 4.5113 | 10.1787 | 22.4191 | 100.6545 | 0.0894 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.913
1-Bromo-4-fluorobenzene (BFB) [SS3] | 26.0069 | 20.3479 | 20.4413 | 20.3676 | 20.4685 | 20.2132 | 20.0694 | 20.2561 | 19.9658 | 19.1828 | 19.6126 | 21.2255 | 19.2257 | 20.7876 | 22.9111 | 21.0499 | 20.058 | 20.2553 | 20.8409 | 22.0509 | 18.9241 | 19.7307 | 20.4693 | 20.0854 | 20.0904 | 20.353 | 19.4061 | 20.6646 | 19.5015 | 20.7698 | 20.0921 | 20.7368 | 20.1732 | 19.4948 | 19.6984 | 18.7533
Bromobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9879 | 1.9711 | 4.5971 | 11.1131 | 23.0216 | 100.4041 | 0.1629 | 0.0553 | 0.0302 | n.a./n.r. | 0.0244 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.095
1,1,2,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8609 | 1.7334 | 4.4057 | 11.3207 | 24.2698 | 100.0869 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8538
1,2,3-Trichloropropane (TCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9231 | 1.8313 | 4.4621 | 11.1843 | 23.3608 | 100.3224 | 0.1127 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2419
trans-1,4-Dichloro-2-butene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8723 | 1.7555 | 4.1778 | 10.8429 | 23.5128 | 100.3348 | 0.0979 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9675
n-Propylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8756 | 1.8681 | 4.4623 | 10.4567 | 22.3635 | 100.6442 | 0.1332 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8166
2-Chlorotoluene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9097 | 1.9241 | 4.5165 | 10.8012 | 22.6232 | 100.5407 | 0.1214 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1757
1,3,5-Trimethylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8409 | 1.8651 | 4.3821 | 10.4792 | 22.3285 | 100.6552 | 0.1013 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7319
4-Chlorotoluene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9351 | 2.012 | 4.8482 | 11.5174 | 24.6022 | 99.9557 | 0.187 | 0.0709 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6941
tert-Butylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8222 | 1.7745 | 4.3127 | 10.1284 | 21.9625 | 100.7872 | 0.1221 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.929
Pentachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.4528 | 2.0206 | 4.1755 | 10.3395 | 25.0081 | 99.999 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6523
1,2,4-Trimethylbenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8513 | 1.8306 | 4.3641 | 10.5819 | 22.2574 | 100.6641 | 0.1188 | 0.0471 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8337
1-Methylpropylbenzene (sec-butylbenzene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8421 | 1.8248 | 4.4599 | 10.2883 | 22.5045 | 100.6271 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.6968
1,3-Dichlorobenzene | 0.0348 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0116 | n.a./n.r. | n.a./n.r. | 0.9575 | 1.9597 | 4.6768 | 11.3461 | 23.7942 | 100.1842 | 0.2633 | 0.0921 | 0.0472 | 0.0371 | 0.0288 | 0.0304 | 0.0203 | 0.0225 | 0.0158 | 0.0116 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0114 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6749
4-Isopropyltoluene (p-cymene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8668 | 1.8584 | 4.5655 | 10.6046 | 23.0488 | 100.4532 | 0.1418 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9158
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9599 | 2.0227 | 4.9069 | 11.7356 | 24.3221 | 100.0005 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9939
n-Butylbenzene | 0.0456 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8159 | 1.7765 | 4.3292 | 9.8988 | 22.5435 | 100.6641 | 0.1904 | 0.0831 | 0.0494 | 0.037 | 0.027 | 0.0255 | n.a./n.r. | 0.0225 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.1505
1,2-Dichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.0271 | 2.0863 | 5.0349 | 12.3767 | 25.2928 | 99.6854 | 0.2715 | 0.0818 | 0.0518 | 0.0362 | 0.0301 | 0.0208 | 0.0155 | 0.0229 | 0.0157 | n.a./n.r. | n.a./n.r. | 0.0176 | n.a./n.r. | n.a./n.r. | 0.0115 | n.a./n.r. | 0.0123 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7253
Hexachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.7103 | 1.7508 | 4.4535 | 10.9077 | 24.7651 | 100.0049 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.9455
1,2-Dibromo-3-chloropropane (DBCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8045 | 1.6148 | 4.0874 | 10.6952 | 24.9438 | 99.9994 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.837
Nitrobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.9366 | 3.9389 | 10.9864 | 24.81 | 100.0035 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3098
1,2,4-Trichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0201 | n.a./n.r. | 0.9297 | 1.8901 | 4.6351 | 11.4471 | 23.5282 | 100.2444 | 0.5972 | 0.1751 | 0.1024 | 0.0701 | 0.0495 | 0.0505 | 0.0335 | 0.0334 | n.a./n.r. | n.a./n.r. | 0.0222 | n.a./n.r. | n.a./n.r. | 0.0167 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4164
Hexachloro-1,3-butadiene | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0112 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8096 | 1.8107 | 4.3569 | 10.3695 | 22.4504 | 100.6383 | 0.2418 | 0.0841 | 0.0533 | 0.0437 | 0.0294 | n.a./n.r. | n.a./n.r. | 0.0173 | 0.0113 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0055 | 0.0109 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.5359
Naphthalene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.8399 | 1.7106 | 4.2653 | 11.0177 | 23.378 | 100.3478 | 0.6867 | 0.1758 | 0.0893 | 0.064 | 0.043 | n.a./n.r. | n.a./n.r. | 0.0245 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2964
1,2,3-Trichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9097 | 1.9418 | 4.718 | 11.5904 | 24.486 | 99.9856 | 0.6874 | 0.1965 | 0.0968 | 0.0674 | 0.0504 | 0.045 | n.a./n.r. | n.a./n.r. | 0.0231 | n.a./n.r. | 0.0221 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8522

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library | Area | Estimated | 20 ppb (TIC) | RT | Area
| No. | min | SI |  |  | SI |  |  | SI |  |  | counts*min | Concentration | Pentafluorobenzene [IS1] | 5.43 | 427045
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC |  | (ppb) | 1,4-Difluorobenzene [IS2] | 6.18 | 612685
12-14-23-JI-02 | 16 | 9.47 | 747 | 3-Heptanone | mainlib | 739 | 3-Hexanone, 5-methyl- | mainlib | 688 | 2-Butanone, 3,3-dimethyl-1-thiocyanato- | mainlib | 16177.338 | 0.4234 | Chlorobenzene-d5 [IS3] | 8.92 | 764164
| 18 | 10.13 | 795 | Hexanal, 2-ethyl- | mainlib | 755 | 2-Ethylhexanal | mainlib | 731 | 1-Octen-3-ol | mainlib | 44748.336 | 1.195 | 1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 748921
| 19 | 10.42 | 853 | Benzaldehyde | mainlib | 813 | ethanone, 2,2-dichloro-2-fluoro-1-phenyl- | mainlib | 789 | 2,4-Heptadien-6-ynal, (E,E)- | mainlib | 56735.239 | 1.5151
| 21 | 10.82 | 761 | 1-Hexanol, 2-ethyl- | mainlib | 749 | 2-Propyl-1-pentanol | mainlib | 729 | 2-Ethyl-1-hexanol | mainlib | 52664.89 | 1.4064