## DFTPP
n.a.

## ICAL
Name | Ret.Time | Quant | Conf1 | Conf2 | IS | Level 1 | Level 2 | Level 3 | Level 4 | Unit | Cal.Type | Number | Rel.Std.Dev. | Coeff.of
| min | [10 empty] | of Points | % | Determination | RSD<20% | R^2>0.99 | Total Analytes | 81
Phenol | 4.87 | 94 | 66 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.2338 | 0.9992 | TRUE | TRUE | Failed | 1
Aniline | 4.918 | 93 | 66 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 7.6873 | 0.99294 | TRUE | TRUE | Allowance | 8.1
Bis(2-chloroethyl) ether | 4.956 | 93 | 63 | 95 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 4.6538 | 0.99694 | TRUE | TRUE | Pass? | TRUE
2-Chlorophenol | 5.021 | 128 | 64 | 130 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 4.5999 | 0.99829 | TRUE | TRUE
1,3-Dichlorobenzene | 5.146 | 146 | 148 | 111 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.6586 | 0.99888 | TRUE | TRUE
1,4-Dichlorobenzene | 5.208 | 146 | 148 | 111 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.223 | 0.99923 | TRUE | TRUE
Benzyl alcohol | 5.299 | 108 | 79 | 77 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.9347 | 0.99971 | TRUE | TRUE
1,2-Dichlorobenzene | 5.338 | 146 | 148 | 111 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 3.1136 | 0.99844 | TRUE | TRUE
2-Methylphenol (o-cresol) | 5.383 | 107 | 108 | 77 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.1231 | 0.99966 | TRUE | TRUE
Dichloroisopropyl ether | 5.414 | 45 | 77 | 121 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.7378 | 0.99568 | TRUE | TRUE
N-Nitroso-di-n-propylamine | 5.517 | 107 | 108 | 77 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.4043 | 0.99957 | TRUE | TRUE
3/4-Methylphenol (m/p-cresol) | 5.531 | 70 | 101 | 130 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.444 | 0.99954 | TRUE | TRUE
Hexachloroethane | 5.637 | 117 | 201 | 199 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.9047 | 0.9988 | TRUE | TRUE
Nitrobenzene | 5.691 | 77 | 123 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.1955 | 0.9992 | TRUE | TRUE
Isophorone | 5.899 | 82 | 54 | 138 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.5552 | 0.99983 | TRUE | TRUE
2-Nitrophenol | 5.975 | 139 | 109 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.2071 | 0.99965 | TRUE | TRUE
2,4-Dimethylphenol | 6.003 | 122 | 107 | 121 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.9808 | 0.9988 | TRUE | TRUE
Bis(2-chloroethoxy)methane | 6.085 | 93 | 95 | 123 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.9746 | 0.99968 | TRUE | TRUE
2,4-Dichlorophenol | 6.192 | 162 | 164 | 98 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.6378 | 0.996 | TRUE | TRUE
1,2,4-Trichlorobenzene | 6.276 | 180 | 182 | 145 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.2675 | 0.99527 | TRUE | TRUE
1,3-Dimethyl-2-nitrobenzene [SS1] | 6.287 | 79 | 77 | 134 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 9.9815 | 0.98836 | TRUE | FALSE
Naphthalene | 6.352 | 128 | 129 | 127 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 1.6098 | 0.99963 | TRUE | TRUE
4-Chloroaniline | 6.394 | 127 | 129 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.5348 | 0.99611 | TRUE | TRUE
Hexachlorobutadiene | 6.463 | 225 | 223 | 227 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.1452 | 0.9955 | TRUE | TRUE
4-Chloro-3-methylphenol | 6.845 | 107 | 144 | 142 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.3202 | 0.99988 | TRUE | TRUE
2-Methylnaphthalene | 7.004 | 142 | 141 |  | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 3.008 | 0.99874 | TRUE | TRUE
1-Methylnaphthalene | 7.101 | 142 | 141 |  | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.1292 | 0.99929 | TRUE | TRUE
Hexachlorocyclopentadiene | 7.162 | 237 | 235 | 272 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 7.2746 | 0.99362 | TRUE | TRUE
2,4,6-Trichlorophenol | 7.268 | 196 | 198 | 200 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 7.4392 | 0.99332 | TRUE | TRUE
2,4,5-Trichlorophenol | 7.313 | 196 | 198 | 97 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.652 | 0.99598 | TRUE | TRUE
2-Chloronaphthalene | 7.467 | 162 | 127 | 164 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.1896 | 0.99928 | TRUE | TRUE
2-Nitroaniline | 7.573 | 65 | 92 | 138 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.2911 | 0.99968 | TRUE | TRUE
1,4-Dinitrobenzene | 7.681 | 168 | 75 | 50 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.3331 | 0.99987 | TRUE | TRUE
Dimethylphthalate | 7.735 | 163 | 194 | 164 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.1636 | 0.99658 | TRUE | TRUE
1,3-Dinitrobenzene | 7.765 | 168 | 76 | 50 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 8.5454 | 0.99149 | TRUE | TRUE
2,6-Dinitrotoluene | 7.792 | 165 | 63 | 89 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 7.3211 | 0.99332 | TRUE | TRUE
1,2-Dinitrobenzene | 7.848 | 168 | 50 | 63 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 6.2041 | 0.99519 | TRUE | TRUE
Acenaphthylene | 7.874 | 152 | 151 | 153 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.1054 | 0.99938 | TRUE | TRUE
3-Nitroaniline | 7.962 | 138 | 108 | 92 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.2143 | 0.9999 | TRUE | TRUE
Acenaphthene-d10 [IS1] | 8.009 | 164 | 162 | 160 | ISTD Internal | 1 | 1 | 1 | 1 | ug/mL | Lin | 4 | n.a. | n.a. | TRUE | TRUE
Acenaphthene | 8.038 | 153 | 152 | 76 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 0.7095 | 0.99992 | TRUE | TRUE
2,4-Dinitrophenol | 8.058 | 184 | 107 | 91 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 0.4057 | 0.99999 | TRUE | TRUE
4-Nitrophenol | 8.118 | 139 | 109 | 65 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.8467 | 0.9995 | TRUE | TRUE
2,4-Dinitrotoluene | 8.184 | 165 | 63 | 89 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.0486 | 0.99969 | TRUE | TRUE
Dibenzofuran | 8.215 | 168 | 139 |  | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.2454 | 0.99922 | TRUE | TRUE
2,3,5,6-Tetrachlorophenol | 8.285 | 131 | 230 | 166 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.8918 | 0.99974 | TRUE | TRUE
2,3,4,6-Tetrachlorophenol | 8.328 | 131 | 230 | 166 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.3846 | 0.99959 | TRUE | TRUE
Diethyl phthalate | 8.42 | 149 | 177 | 150 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 7.2109 | 0.99359 | TRUE | TRUE
4-Chlorophenyl phenyl ether | 8.537 | 204 | 206 | 141 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.0066 | 0.99938 | TRUE | TRUE
Fluorene | 8.546 | 166 | 165 | 167 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.1578 | 0.9993 | TRUE | TRUE
4-Nitroaniline | 8.559 | 138 | 65 | 108 | Internal Acenaphthene-d10 [IS1] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.9127 | 0.99975 | TRUE | TRUE
4,6-Dinitro-2-methylphenol (dinitro-o-cresol) | 8.59 | 198 | 105 |  | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.3813 | 0.99937 | TRUE | TRUE
N-nitrosodiphenylamine/Diphenylamine | 8.649 | 169 | 168 | 167 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.4152 | 0.99624 | TRUE | TRUE
Azobenzene | 8.696 | 182 | 105 |  | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.3573 | 0.99919 | TRUE | TRUE
4-Bromophenyl phenyl ether | 9.011 | 248 | 250 | 141 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 3.9108 | 0.99787 | TRUE | TRUE
Hexachlorobenzene | 9.087 | 284 | 142 | 249 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.4798 | 0.99914 | TRUE | TRUE
Pentachlorophenol | 9.276 | 266 | 264 | 268 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.472 | 0.9996 | TRUE | TRUE
Phenanthrene-d10 [IS2] | 9.453 | 188 | 94 | 80 | ISTD Internal | 1 | 1 | 1 | 1 | ug/mL | Lin | 4 | n.a. | n.a. | TRUE | TRUE
Phenanthrene | 9.48 | 178 | 176 | 179 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.6487 | 0.99892 | TRUE | TRUE
Anthracene | 9.531 | 178 | 176 | 179 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.2591 | 0.99929 | TRUE | TRUE
Carbazole | 9.695 | 167 | 168 | 139 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 5.4201 | 0.99619 | TRUE | TRUE
Di-n-butyl phthalate | 10.011 | 149 | 150 | 104 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.7265 | 0.99978 | TRUE | TRUE
Fluoranthene | 10.653 | 202 | 101 | 203 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 1.7506 | 0.99957 | TRUE | TRUE
Pyrene-d10 [SS2] | 10.848 | 212 | 210 | 213 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 4.1111 | 0.99768 | TRUE | TRUE
Pyrene | 10.874 | 202 | 200 | 203 | Internal Phenanthrene-d10 [IS2] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 2.5356 | 0.99901 | TRUE | TRUE
Benzyl butyl phthalate | 11.527 | 149 | 91 | 206 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.8697 | 0.99979 | TRUE | TRUE
Bis(2-ethylhexyl)adipate | 11.618 | 129 | 112 | 147 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.7313 | 0.99955 | TRUE | TRUE
Triphenyl phosphate [SS3] | 11.774 | 326 | 325 | 94 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.5549 | 0.99963 | TRUE | TRUE
Benzo[a]anthracene | 12.227 | 228 | 229 | 226 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 8.5376 | 0.99142 | TRUE | TRUE
Chrysene d-12 [IS3] | 12.249 | 240 | 120 | 236 | ISTD Internal | 1 | 1 | 1 | 1 | ug/mL | Lin | 4 | n.a. | n.a. | TRUE | TRUE
Bis(2-ethylhexyl)phthalate | 12.253 | 149 | 167 | 279 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.1951 | 0.99933 | TRUE | TRUE
Chrysene | 12.288 | 228 | 227 | 229 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Lin | 4 | 4.9357 | 0.99667 | TRUE | TRUE
Di-n-octyl phthalate | 13.226 | 149 | 167 | 57 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.8001 | 0.99916 | TRUE | TRUE
Benzo[b]fluoranthene | 13.861 | 252 | 253 | 125 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 1.1768 | 0.99991 | TRUE | TRUE
Benzo[k]fluoranthene | 13.94 | 252 | 253 | 125 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.3298 | 0.99924 | TRUE | TRUE
Benzo[a]pyrene | 14.446 | 252 | 253 | 125 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 4.1644 | 0.99893 | TRUE | TRUE
Perylene-d12 [SS4] | 14.585 | 264 | 263 | 265 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 3.2225 | 0.99929 | TRUE | TRUE
Indeno[1,2,3-cd]pyrene | 16.629 | 276 | 138 | 277 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Cubic | 4 | 2.9654 | 0.99978 | TRUE | TRUE
Dibenz[a,h]anthracene | 16.724 | 278 | 138 | 279 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 2.254 | 0.99972 | TRUE | TRUE
Benzo[g,h,i]perylene | 17.194 | 276 | 138 | 277 | Internal Chrysene d-12 [IS3] | 0.5 | 1 | 2 | 4 | ug/mL | Quad | 4 | 4.0855 | 0.99904 | TRUE | TRUE

## Carryover
4 ug/mL |  |  | DCM
Peak Name | Ret. Time | Area | Peak Name | Ret. Time | Area
| min | counts*min |  | min | counts*min | Carryover (%)
Phenol | 4.86 | 749850 | Phenol | n.a. | n.a. | #VALUE!
Aniline | 4.91 | 1012195 | Aniline | n.a. | n.a. | #VALUE!
Bis(2-chloroethyl) ether | 4.95 | 995483 | Bis(2-chloroethyl) ether | n.a. | n.a. | #VALUE!
2-Chlorophenol | 5.01 | 498183 | 2-Chlorophenol | n.a. | n.a. | #VALUE!
1,3-Dichlorobenzene | 5.14 | 644815 | 1,3-Dichlorobenzene | 5.15 | 22 | 0.0034118
1,4-Dichlorobenzene | 5.2 | 671538 | 1,4-Dichlorobenzene | 5.22 | 13 | 0.0019359
Benzyl alcohol | 5.29 | 351208 | Benzyl alcohol | n.a. | n.a. | #VALUE!
1,2-Dichlorobenzene | 5.33 | 619283 | 1,2-Dichlorobenzene | 5.35 | 6 | 0.00096886
2-Methylphenol (o-cresol) | 5.38 | 504790 | 2-Methylphenol (o-cresol) | n.a. | n.a. | #VALUE!
Dichloroisopropyl ether | 5.41 | 968086 | Dichloroisopropyl ether | n.a. | n.a. | #VALUE!
N-Nitroso-di-n-propylamine | 5.51 | 588682 | N-Nitroso-di-n-propylamine | 5.51 | 60 | 0.010192
3/4-Methylphenol (m/p-cresol) | 5.52 | 652459 | 3/4-Methylphenol (m/p-cresol) | n.a. | n.a. | #VALUE!
Hexachloroethane | 5.63 | 296533 | Hexachloroethane | n.a. | n.a. | #VALUE!
Nitrobenzene | 5.68 | 878060 | Nitrobenzene | n.a. | n.a. | #VALUE!
Isophorone | 5.89 | 1749668 | Isophorone | n.a. | n.a. | #VALUE!
2-Nitrophenol | 5.97 | 167813 | 2-Nitrophenol | n.a. | n.a. | #VALUE!
2,4-Dimethylphenol | 5.99 | 436490 | 2,4-Dimethylphenol | n.a. | n.a. | #VALUE!
Bis(2-chloroethoxy)methane | 6.08 | 1031029 | Bis(2-chloroethoxy)methane | n.a. | n.a. | #VALUE!
2,4-Dichlorophenol | 6.18 | 271248 | 2,4-Dichlorophenol | n.a. | n.a. | #VALUE!
1,2,4-Trichlorobenzene | 6.27 | 372746 | 1,2,4-Trichlorobenzene | n.a. | n.a. | #VALUE!
1,3-Dimethyl-2-nitrobenzene [SS1] | 6.28 | 643129 | 1,3-Dimethyl-2-nitrobenzene [SS1] | n.a. | n.a. | #VALUE!
Naphthalene | 6.34 | 2015993 | Naphthalene | n.a. | n.a. | #VALUE!
4-Chloroaniline | 6.39 | 707152 | 4-Chloroaniline | 6.41 | 200 | 0.028282
Hexachlorobutadiene | 6.46 | 177928 | Hexachlorobutadiene | 6.44 | 2 | 0.0011241
4-Chloro-3-methylphenol | 6.83 | 542300 | 4-Chloro-3-methylphenol | 6.85 | 83 | 0.015305
2-Methylnaphthalene | 7 | 1208655 | 2-Methylnaphthalene | 7 | 33 | 0.0027303
1-Methylnaphthalene | 7.09 | 1083501 | 1-Methylnaphthalene | 7.1 | 25 | 0.0023073
Hexachlorocyclopentadiene | 7.15 | 129391 | Hexachlorocyclopentadiene | n.a. | n.a. | #VALUE!
2,4,6-Trichlorophenol | 7.26 | 135994 | 2,4,6-Trichlorophenol | n.a. | n.a. | #VALUE!
2,4,5-Trichlorophenol | 7.3 | 147001 | 2,4,5-Trichlorophenol | n.a. | n.a. | #VALUE!
2-Chloronaphthalene | 7.46 | 894323 | 2-Chloronaphthalene | n.a. | n.a. | #VALUE!
2-Nitroaniline | 7.56 | 369675 | 2-Nitroaniline | n.a. | n.a. | #VALUE!
1,4-Dinitrobenzene | 7.68 | 82800 | 1,4-Dinitrobenzene | n.a. | n.a. | #VALUE!
Dimethylphthalate | 7.73 | 746388 | Dimethylphthalate | n.a. | n.a. | #VALUE!
1,3-Dinitrobenzene | 7.76 | 98260 | 1,3-Dinitrobenzene | n.a. | n.a. | #VALUE!
2,6-Dinitrotoluene | 7.79 | 162456 | 2,6-Dinitrotoluene | n.a. | n.a. | #VALUE!
1,2-Dinitrobenzene | 7.84 | 63837 | 1,2-Dinitrobenzene | n.a. | n.a. | #VALUE!
Acenaphthylene | 7.87 | 1541499 | Acenaphthylene | 7.87 | 67 | 0.0043464
3-Nitroaniline | 7.96 | 195356 | 3-Nitroaniline | n.a. | n.a. | #VALUE!
Acenaphthene-d10 [IS1] | 8 | 190073 | Acenaphthene-d10 [IS1] | n.a. | n.a. | #VALUE!
Acenaphthene | 8.04 | 1078461 | Acenaphthene | 8.04 | 47 | 0.0043581
2,4-Dinitrophenol | 8.05 | 38061 | 2,4-Dinitrophenol | 8.07 | 0 | 0
4-Nitrophenol | 8.11 | 106449 | 4-Nitrophenol | n.a. | n.a. | #VALUE!
2,4-Dinitrotoluene | 8.18 | 220717 | 2,4-Dinitrotoluene | 8.18 | 6 | 0.0027184
Dibenzofuran | 8.2 | 1224387 | Dibenzofuran | 8.21 | 108 | 0.0088207
2,3,5,6-Tetrachlorophenol | 8.28 | 100994 | 2,3,5,6-Tetrachlorophenol | n.a. | n.a. | #VALUE!
2,3,4,6-Tetrachlorophenol | 8.32 | 126794 | 2,3,4,6-Tetrachlorophenol | n.a. | n.a. | #VALUE!
Diethyl phthalate | 8.42 | 878271 | Diethyl phthalate | 8.42 | 282 | 0.032109
4-Chlorophenyl phenyl ether | 8.53 | 346129 | 4-Chlorophenyl phenyl ether | 8.53 | 69 | 0.019935
Fluorene | 8.54 | 965115 | Fluorene | 8.54 | 54 | 0.0055952
4-Nitroaniline | 8.55 | 137610 | 4-Nitroaniline | n.a. | n.a. | #VALUE!
4,6-Dinitro-2-methylphenol (dinitro-o-cresol) | 8.58 | 58554 | 4,6-Dinitro-2-methylphenol (dinitro-o-cresol) | 8.58 | 2 | 0.0034157
N-nitrosodiphenylamine/Diphenylamine | 8.64 | 806375 | N-nitrosodiphenylamine/Diphenylamine | 8.65 | 44 | 0.0054565
Azobenzene | 8.68 | 220505 | Azobenzene | 8.69 | 18 | 0.0081631
4-Bromophenyl phenyl ether | 9.01 | 231467 | 4-Bromophenyl phenyl ether | 9.01 | 14 | 0.0060484
Hexachlorobenzene | 9.08 | 344221 | Hexachlorobenzene | 9.08 | 103 | 0.029923
Pentachlorophenol | 9.27 | 104984 | Pentachlorophenol | n.a. | n.a. | #VALUE!
Phenanthrene-d10 [IS2] | 9.45 | 260517 | Phenanthrene-d10 [IS2] | 9.45 | 29 | 0.011132
Phenanthrene | 9.47 | 1251193 | Phenanthrene | 9.48 | 94 | 0.0075128
Anthracene | 9.52 | 1327108 | Anthracene | 9.53 | 77 | 0.0058021
Carbazole | 9.68 | 1092873 | Carbazole | 9.69 | 79 | 0.0072287
Di-n-butyl phthalate | 10.01 | 1343308 | Di-n-butyl phthalate | 10.02 | 78 | 0.0058066
Fluoranthene | 10.64 | 1125090 | Fluoranthene | 10.64 | 53 | 0.0047107
Pyrene-d10 [SS2] | 10.84 | 833069 | Pyrene-d10 [SS2] | 10.85 | 70 | 0.0084027
Pyrene | 10.86 | 1160080 | Pyrene | 10.86 | 67 | 0.0057755
Benzyl butyl phthalate | 11.52 | 484818 | Benzyl butyl phthalate | n.a. | n.a. | #VALUE!
Bis(2-ethylhexyl)adipate | 11.61 | 558992 | Bis(2-ethylhexyl)adipate | n.a. | n.a. | #VALUE!
Triphenyl phosphate [SS3] | 11.76 | 231195 | Triphenyl phosphate [SS3] | n.a. | n.a. | #VALUE!
Benzo[a]anthracene | 12.23 | 885304 | Benzo[a]anthracene | n.a. | n.a. | #VALUE!
Chrysene d-12 [IS3] | 12.24 | 198297 | Chrysene d-12 [IS3] | 12.24 | 8 | 0.0040344
Bis(2-ethylhexyl)phthalate | 12.25 | 822654 | Bis(2-ethylhexyl)phthalate | 12.25 | 1687 | 0.20507
Chrysene | 12.27 | 1094799 | Chrysene | 12.28 | 30 | 0.0027402
Di-n-octyl phthalate | 13.21 | 1130538 | Di-n-octyl phthalate | n.a. | n.a. | #VALUE!
Benzo[b]fluoranthene | 13.87 | 805129 | Benzo[b]fluoranthene | n.a. | n.a. | #VALUE!
Benzo[k]fluoranthene | 13.92 | 1238886 | Benzo[k]fluoranthene | n.a. | n.a. | #VALUE!
Benzo[a]pyrene | 14.44 | 728451 | Benzo[a]pyrene | n.a. | n.a. | #VALUE!
Perylene-d12 [SS4] | 14.54 | 835440 | Perylene-d12 [SS4] | n.a. | n.a. | #VALUE!
Indeno[1,2,3-cd]pyrene | 16.62 | 582907 | Indeno[1,2,3-cd]pyrene | n.a. | n.a. | #VALUE!
Dibenz[a,h]anthracene | 16.67 | 914249 | Dibenz[a,h]anthracene | n.a. | n.a. | #VALUE!
Benzo[g,h,i]perylene | 17.16 | 1034339 | Benzo[g,h,i]perylene | n.a. | n.a. | #VALUE!

## CCV
Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio |  |  |  | RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area?
| min | counts*min | % | ppb | Confirmation | True Value | Pass? | Acenaphthene-d10 [IS1] | 8 | 206237 | 8.009 | 225677 | TRUE | TRUE
Phenol | 4.86 | 185942 | 0.95 | 1.148 | Confirmed | 1 | TRUE | Phenanthrene-d10 [IS2] | 9.45 | 272402 | 9.453 | 306073 | TRUE | TRUE
Aniline | 4.91 | 247116 | 1.26 | 0.925 | Confirmed | 1 | TRUE | Chrysene d-12 [IS3] | 12.23 | 197650 | 12.249 | 214873 | TRUE | TRUE
Bis(2-chloroethyl) ether | 4.94 | 275122 | 1.4 | 1.035 | Confirmed | 1 | TRUE
2-Chlorophenol | 5.01 | 122585 | 0.62 | 1.103 | Confirmed | 1 | TRUE
1,3-Dichlorobenzene | 5.14 | 179631 | 0.92 | 1.019 | Confirmed | 1 | TRUE
1,4-Dichlorobenzene | 5.2 | 186656 | 0.95 | 1.018 | Confirmed | 1 | TRUE | Total Analytes | 81
Benzyl alcohol | 5.29 | 85336 | 0.43 | 1.101 | Confirmed | 1 | TRUE | Failed | 3
1,2-Dichlorobenzene | 5.33 | 170685 | 0.87 | 1.006 | Confirmed | 1 | TRUE | Allowance | 16.2
2-Methylphenol (o-cresol) | 5.37 | 117128 | 0.6 | 1.081 | Confirmed | 1 | TRUE | Pass? | TRUE
Dichloroisopropyl ether | 5.4 | 271911 | 1.39 | 1.059 | Confirmed | 1 | TRUE
N-Nitroso-di-n-propylamine | 5.5 | 138405 | 0.71 | 1.121 | Confirmed | 1 | TRUE
3/4-Methylphenol (m/p-cresol) | 5.52 | 159119 | 0.81 | 1.126 | Confirmed | 1 | TRUE
Hexachloroethane | 5.63 | 77677 | 0.4 | 0.976 | Confirmed | 1 | TRUE
Nitrobenzene | 5.68 | 200320 | 1.02 | 1.045 | Confirmed | 1 | TRUE
Isophorone | 5.89 | 420298 | 2.14 | 1.181 | Confirmed | 1 | TRUE
2-Nitrophenol | 5.96 | 37039 | 0.19 | 1.094 | Confirmed | 1 | TRUE
2,4-Dimethylphenol | 5.99 | 120259 | 0.61 | 1.026 | Confirmed | 1 | TRUE
Bis(2-chloroethoxy)methane | 6.07 | 267258 | 1.36 | 1.119 | Confirmed | 1 | TRUE
2,4-Dichlorophenol | 6.18 | 63723 | 0.32 | 0.883 | Confirmed | 1 | TRUE
1,2,4-Trichlorobenzene | 6.26 | 108523 | 0.55 | 1.052 | Confirmed | 1 | TRUE
1,3-Dimethyl-2-nitrobenzene [SS1] | 6.28 | 147257 | 0.75 | 0.877 | Confirmed | 1 | TRUE
Naphthalene | 6.34 | 557726 | 2.84 | 1.026 | Confirmed | 1 | TRUE
4-Chloroaniline | 6.38 | 174210 | 0.89 | 0.926 | Confirmed | 1 | TRUE
Hexachlorobutadiene | 6.46 | 50586 | 0.26 | 1.028 | Confirmed | 1 | TRUE
4-Chloro-3-methylphenol | 6.83 | 114701 | 0.58 | 1.101 | Confirmed | 1 | TRUE
2-Methylnaphthalene | 6.99 | 321866 | 1.64 | 0.993 | Confirmed | 1 | TRUE
1-Methylnaphthalene | 7.09 | 312166 | 1.59 | 1.054 | Confirmed | 1 | TRUE
Hexachlorocyclopentadiene | 7.15 | 29726 | 0.15 | 0.869 | Confirmed | 1 | TRUE
2,4,6-Trichlorophenol | 7.26 | 31514 | 0.16 | 0.877 | Confirmed | 1 | TRUE
2,4,5-Trichlorophenol | 7.29 | 37283 | 0.19 | 0.944 | Confirmed | 1 | TRUE
2-Chloronaphthalene | 7.46 | 245587 | 1.25 | 1.015 | Confirmed | 1 | TRUE
2-Nitroaniline | 7.56 | 72212 | 0.37 | 1.116 | Confirmed | 1 | TRUE
1,4-Dinitrobenzene | 7.67 | 16567 | 0.08 | 0.986 | Confirmed | 1 | TRUE
Dimethylphthalate | 7.73 | 191208 | 0.97 | 0.962 | Confirmed | 1 | TRUE
1,3-Dinitrobenzene | 7.75 | 20163 | 0.1 | 0.778 | Confirmed | 1 | FALSE
2,6-Dinitrotoluene | 7.78 | 35857 | 0.18 | 0.837 | Confirmed | 1 | TRUE
1,2-Dinitrobenzene | 7.83 | 14771 | 0.08 | 0.872 | Confirmed | 1 | TRUE
Acenaphthylene | 7.86 | 395309 | 2.01 | 0.953 | Confirmed | 1 | TRUE
3-Nitroaniline | 7.95 | 37922 | 0.19 | 0.969 | Confirmed | 1 | TRUE
Acenaphthene-d10 [IS1] | 8 | 206237 | 1.05 | 1 | Confirmed | 1 | TRUE
Acenaphthene | 8.03 | 290271 | 1.48 | 0.992 | Confirmed | 1 | TRUE
2,4-Dinitrophenol | 8.05 | 4788 | 0.02 | 1.126 | Confirmed | 1 | TRUE
4-Nitrophenol | 8.11 | 20176 | 0.1 | 1.028 | Confirmed | 1 | TRUE
2,4-Dinitrotoluene | 8.17 | 46597 | 0.24 | 0.995 | Confirmed | 1 | TRUE
Dibenzofuran | 8.2 | 332985 | 1.7 | 0.999 | Confirmed | 1 | TRUE
2,3,5,6-Tetrachlorophenol | 8.28 | 21346 | 0.11 | 1.011 | Confirmed | 1 | TRUE
2,3,4,6-Tetrachlorophenol | 8.32 | 27746 | 0.14 | 1.044 | Confirmed | 1 | TRUE
Diethyl phthalate | 8.41 | 216021 | 1.1 | 0.932 | Confirmed | 1 | TRUE
4-Chlorophenyl phenyl ether | 8.52 | 99013 | 0.5 | 1.046 | Confirmed | 1 | TRUE
Fluorene | 8.53 | 270045 | 1.38 | 1.023 | Confirmed | 1 | TRUE
4-Nitroaniline | 8.55 | 33547 | 0.17 | 1.258 | Confirmed | 1 | FALSE
4,6-Dinitro-2-methylphenol (dinitro-o-cresol) | 8.57 | 9332 | 0.05 | 1.039 | Confirmed | 1 | TRUE
N-nitrosodiphenylamine/Diphenylamine | 8.64 | 213036 | 1.09 | 1.032 | Confirmed | 1 | TRUE
Azobenzene | 8.68 | 60555 | 0.31 | 1.057 | Confirmed | 1 | TRUE
4-Bromophenyl phenyl ether | 9 | 61008 | 0.31 | 1.023 | Confirmed | 1 | TRUE
Hexachlorobenzene | 9.08 | 95904 | 0.49 | 1.076 | Confirmed | 1 | TRUE
Pentachlorophenol | 9.26 | 21540 | 0.11 | 1.124 | Confirmed | 1 | TRUE
Phenanthrene-d10 [IS2] | 9.45 | 272402 | 1.39 | 1 | Confirmed | 1 | TRUE
Phenanthrene | 9.47 | 349029 | 1.78 | 1.066 | Confirmed | 1 | TRUE
Anthracene | 9.52 | 347484 | 1.77 | 1.009 | Confirmed | 1 | TRUE
Carbazole | 9.67 | 289916 | 1.48 | 1.036 | Confirmed | 1 | TRUE
Di-n-butyl phthalate | 10 | 314061 | 1.6 | 1.139 | Confirmed | 1 | TRUE
Fluoranthene | 10.63 | 301855 | 1.54 | 1.031 | Confirmed | 1 | TRUE
Pyrene-d10 [SS2] | 10.84 | 218932 | 1.12 | 1.018 | Confirmed | 1 | TRUE
Pyrene | 10.85 | 332339 | 1.69 | 1.091 | Confirmed | 1 | TRUE
Benzyl butyl phthalate | 11.52 | 94346 | 0.48 | 1.176 | Confirmed | 1 | TRUE
Bis(2-ethylhexyl)adipate | 11.6 | 102911 | 0.52 | 1.141 | Confirmed | 1 | TRUE
Triphenyl phosphate [SS3] | 11.75 | 45393 | 0.23 | 1.273 | Confirmed | 1 | FALSE
Benzo[a]anthracene | 12.22 | 197677 | 1.01 | 0.924 | Confirmed | 1 | TRUE
Chrysene d-12 [IS3] | 12.23 | 197650 | 1.01 | 1 | Confirmed | 1 | TRUE
Bis(2-ethylhexyl)phthalate | 12.24 | 164004 | 0.84 | 1.106 | Confirmed | 1 | TRUE
Chrysene | 12.26 | 260209 | 1.33 | 0.971 | Confirmed | 1 | TRUE
Di-n-octyl phthalate | 13.2 | 193318 | 0.99 | 1.091 | Confirmed | 1 | TRUE
Benzo[b]fluoranthene | 13.85 | 160248 | 0.82 | 1.12 | Confirmed | 1 | TRUE
Benzo[k]fluoranthene | 13.9 | 261981 | 1.34 | 1.175 | Confirmed | 1 | TRUE
Benzo[a]pyrene | 14.43 | 140503 | 0.72 | 1.132 | Confirmed | 1 | TRUE
Perylene-d12 [SS4] | 14.53 | 178664 | 0.91 | 1.148 | Confirmed | 1 | TRUE
Indeno[1,2,3-cd]pyrene | 16.61 | 88447 | 0.45 | 1.157 | Confirmed | 1 | TRUE
Dibenz[a,h]anthracene | 16.65 | 152006 | 0.77 | 1.15 | Confirmed | 1 | TRUE
Benzo[g,h,i]perylene | 17.14 | 185126 | 0.94 | 1.185 | Confirmed | 1 | TRUE

## Blank
Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio
| min | counts*min | % | ppb | Confirmation
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | LLOQ | <1/2LLOQ | Non-target | Fail?
Phenol | 4.87 | 230 | 0.03 | 0.002 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Aniline | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Bis(2-chloroethyl) ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Chlorophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,3-Dichlorobenzene | 5.15 | 17 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,4-Dichlorobenzene | 5.2 | 16 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzyl alcohol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,2-Dichlorobenzene | 5.33 | 14 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Methylphenol (o-cresol) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Dichloroisopropyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
N-Nitroso-di-n-propylamine | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
3/4-Methylphenol (m/p-cresol) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Isophorone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Nitrophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4-Dimethylphenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Bis(2-chloroethoxy)methane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4-Dichlorophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,2,4-Trichlorobenzene | 6.27 | 15 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,3-Dimethyl-2-nitrobenzene [SS1] | 6.28 | 70333 | 10.12 | 0.646 | Confirmed | 0.5
Naphthalene | 6.35 | 113 | 0.02 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Chloroaniline | 6.4 | 35 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Hexachlorobutadiene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Chloro-3-methylphenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Methylnaphthalene | 7 | 65 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1-Methylnaphthalene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Hexachlorocyclopentadiene | 7.16 | 1 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4,6-Trichlorophenol | 7.27 | 21 | 0 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4,5-Trichlorophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Chloronaphthalene | 7.46 | 59 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2-Nitroaniline | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,4-Dinitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Dimethylphthalate | 7.74 | 255 | 0.04 | 0.002 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,3-Dinitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,6-Dinitrotoluene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
1,2-Dinitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Acenaphthylene | 7.86 | 74 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
3-Nitroaniline | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Acenaphthene-d10 [IS1] | 8 | 133658 | 19.24 | 1 | Confirmed | 1
Acenaphthene | 8.03 | 89 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4-Dinitrophenol | 8.07 | 1 | 0 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Nitrophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,4-Dinitrotoluene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Dibenzofuran | 8.2 | 74 | 0.01 | 0 | Confirmed | 0.25 | TRUE | FALSE | FALSE
2,3,5,6-Tetrachlorophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
2,3,4,6-Tetrachlorophenol | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Diethyl phthalate | 8.42 | 358 | 0.05 | 0.002 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Chlorophenyl phenyl ether | 8.53 | 24 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Fluorene | 8.54 | 56 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Nitroaniline | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4,6-Dinitro-2-methylphenol (dinitro-o-cresol) | 8.59 | 8 | 0 | 0.002 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
N-nitrosodiphenylamine/Diphenylamine | 8.65 | 59 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Azobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
4-Bromophenyl phenyl ether | 9 | 15 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Hexachlorobenzene | 9.08 | 37 | 0.01 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Pentachlorophenol | 9.27 | 12 | 0 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Phenanthrene-d10 [IS2] | 9.45 | 191289 | 27.54 | 1 | Confirmed | 1
Phenanthrene | 9.47 | 115 | 0.02 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Anthracene | 9.52 | 99 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Carbazole | 9.68 | 61 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Di-n-butyl phthalate | 10.01 | 5322 | 0.77 | 0.031 | Confirmed | 0.25 | TRUE | FALSE | FALSE
Fluoranthene | 10.66 | 8 | 0 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Pyrene-d10 [SS2] | 10.84 | 101647 | 14.63 | 0.673 | Confirmed | 0.25
Pyrene | 10.86 | 81 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzyl butyl phthalate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Bis(2-ethylhexyl)adipate | 11.61 | 251 | 0.04 | 0.006 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Triphenyl phosphate [SS3] | 11.76 | 17738 | 2.55 | 0.893 | Confirmed | 0.25
Chrysene d-12 [IS3] | 12.24 | 120243 | 17.31 | 1 | Confirmed | 1
Benzo[a]anthracene | 12.24 | 193 | 0.03 | 0.001 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Bis(2-ethylhexyl)phthalate | 12.24 | 3035 | 0.44 | 0.039 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Chrysene | 12.27 | 58 | 0.01 | 0 | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Di-n-octyl phthalate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzo[b]fluoranthene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzo[k]fluoranthene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzo[a]pyrene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Perylene-d12 [SS4] | 14.54 | 37734 | 5.43 | 0.435 | Confirmed | 0.25
Indeno[1,2,3-cd]pyrene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Dibenz[a,h]anthracene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE
Benzo[g,h,i]perylene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 0.25 | TRUE | TRUE | FALSE