## IS-RSD
[8 empty] | Peak Name | Peak Area
[9 empty] | counts*min | 1 | 2 | 3 | 4 | 5
Total Ion Count | [7 empty] | First Injection | Startup | H2O-R1 | H2O-R2 | H2O-R3 | H2O-R4
RSD | RSD (1-3) | RSD (2-4) | RSD (3-5) | RSD (4-6) | RSD (5-7) | RSD (6-8) | RSD (7-9) | TIC | TIC | TIC | TIC | TIC | TIC
2.6496 | 3.0548 | 2.7489 | 1.1454 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Pentafluorobenzene [IS1] | n.a. | 622895.47 | 621091.567 | 655457.034 | 644924.718
3.6278 | 3.9586 | 4.0752 | 0.68175 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1,4-Difluorobenzene [IS2] | n.a. | 813468.18 | 797505.602 | 860248.215 | 851993.95
5.3068 | 5.2463 | 6.4162 | 1.0702 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Chlorobenzene-d5 [IS3] | n.a. | 1077668.057 | 1012964.7 | 1124982.243 | 1142138.029
8.1794 | 8.1873 | 8.2231 | 0.44203 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1,4-Dichlorobenzene-d4 [IS4] | n.a. | 1145658.096 | 1135830.816 | 1310313.115 | 1318529.884

## BFB
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2023/231121.seq/693.smp/MSDevice.xraw
Channel | TIC
File Path | chrom://wrrcgcms/ChromeleonLocal/Instrument Data/GC_MS_PT/2023/231205.seq/336.smp/MSDevice.xraw |  |  | Mass (m/z) | Value (counts) | Criteria | Check | Actual (%) | Min (%) | Max (%) | Pass?
Channel | TIC |  |  | 95 | 10618347 | 50-200% of mass 174 | 95/174 | 135.6852 | 50 | 200 | TRUE
|  |  |  | 96 | 724106.6875 | 5 to 9% of m/z 95 | 96/95 | 6.8194 | 5 | 9 | TRUE
Injection Information: |  |  |  | 173 | #N/A | <2% of m/z 174 | 173/174 | #N/A | 0 | 2 | #N/A
Data Vault | ChromeleonLocal |  |  | 174 | 7825724.5 | 50-200% of mass 95 | 174/95 | 73.7 | 50 | 200 | TRUE
Injection | H2O-R5 |  |  | 175 | 540060.0625 | 5 to 9% of m/z 174 | 175/174 | 6.9011 | 5 | 9 | TRUE
Injection Number | 7 |  |  | 176 | 7792182 | 95 to 105% of m/z 174 | 176/174 | 99.5714 | 95 | 105 | TRUE
Position | 0 |  |  | 177 | 477481.875 | 5 to 10% of m/z 176 | 177/176 | 6.1277 | 5 | 10 | TRUE
Comment
Processing Method | VOC-10
Instrument Method | VOC-26
Type | Unknown
Status | Finished | *N/A for P&T
Injection Date | 2023-12-05
Injection Time | 13:38
Injection Volume (µl) | 20
Dilution Factor | 1
Weight | 1
Raw Data Information:
Data Points | 235
Mass Min. (m/z) | 35
Mass Max. (m/z) | 259.9807
Retention Time (min) | 9.8387
Mass Resolution | 0.5
Tolerance Unit | Amu
Scan Number | 3287
Scan Type | + c EI Full ms [35.00-260.00]
Is Centroid Scan | TRUE
Base Peak Mass | 95.0202
Base Peak Intensity | 10618347 | Rounded:
Total Ion Count | 42344168.686 | Mass (m/z) | Value (counts)
|  | 0 | 0
Raw Data: |  | #VALUE! | 0
Mass (m/z) | Value (counts) | #VALUE! | Value (counts)
36.1 | 47733.4141 | 36 | 47733.4141
37 | 274753.2188 | 37 | 274753.2188
38.1 | 243540.1562 | 38 | 243540.1562
39.1 | 102896.5859 | 39 | 102896.5859
40 | 60194.2969 | 40 | 60194.2969
41 | 3823.3406 | 41 | 3823.3406
41.8 | 985.1888 | 42 | 985.1888
43 | 2868.4478 | 43 | 2868.4478
44 | 142457.4375 | 44 | 142457.4375
45 | 59624.8594 | 45 | 59624.8594
46 | 4359.2065 | 46 | 4359.2065
47.1 | 122205.0391 | 47 | 122205.0391
48 | 42992.5 | 48 | 42992.5
49 | 243613.7969 | 49 | 243613.7969
50 | 1284194 | 50 | 1284194
51 | 369523.375 | 51 | 369523.375
52 | 13022.7217 | 52 | 13022.7217
53 | 768.8026 | 53 | 768.8026
54 | 1495.158 | 54 | 1495.158
55.1 | 11345.502 | 55 | 11345.502
56 | 67208.4453 | 56 | 67208.4453
57 | 142450.8906 | 57 | 142450.8906
58 | 6332.9717 | 58 | 6332.9717
59 | 1456.0457 | 59 | 1456.0457
60 | 51690.0234 | 60 | 51690.0234
61 | 267751.5312 | 61 | 267751.5312
62 | 253498.9688 | 62 | 253498.9688
63 | 179837.2031 | 63 | 179837.2031
64 | 22812.8242 | 64 | 22812.8242
65.1 | 9832.8984 | 65 | 9832.8984
66.1 | 1072.6223 | 66 | 1072.6223
67 | 21036.2598 | 67 | 21036.2598
68 | 725109.1875 | 68 | 725109.1875
69 | 661110.6875 | 69 | 661110.6875
70 | 59393.3516 | 70 | 59393.3516
71 | 1430.7936 | 71 | 1430.7936
72 | 28168.2656 | 72 | 28168.2656
73 | 279810.625 | 73 | 279810.625
74 | 1094287.75 | 74 | 1094287.75
75 | 4098431.75 | 75 | 4098431.75
76.1 | 338869.1562 | 76 | 338869.1562
77 | 56447.1406 | 77 | 56447.1406
78 | 45275.6211 | 78 | 45275.6211
78.9 | 66631.5469 | 79 | 66631.5469
80 | 21208.2949 | 80 | 21208.2949
80.9 | 68712.1328 | 81 | 68712.1328
82 | 12978.7617 | 82 | 12978.7617
82.9 | 2749.3621 | 83 | 2749.3621
84.1 | 694.6611 | 84 | 694.6611
85.2 | 46.3331 | 85 | 46.3331
86 | 10210.3818 | 86 | 10210.3818
87 | 424154.5 | 87 | 424154.5
88 | 444519.5312 | 88 | 444519.5312
89.5 | 83.5332 | 90 | 83.5332
90.9 | 17165.1484 | 91 | 17165.1484
92 | 179325.8281 | 92 | 179325.8281
93 | 297831.4375 | 93 | 297831.4375
94 | 842324.4375 | 94 | 842324.4375
95 | 10618347 | 95 | 10618347
96 | 724106.6875 | 96 | 724106.6875
97.1 | 20622.5098 | 97 | 20622.5098
98.1 | 726.8544 | 98 | 726.8544
98.8 | 172.246 | 99 | 172.246
99.5 | 92.6932 | 100 | 92.6932
100.1 | 0.3625 | 100 | 0.3625
101.4 | 130.1849 | 101 | 130.1849
101.9 | 65.6123 | 102 | 65.6123
102.8 | 2607.2102 | 103 | 2607.2102
103.9 | 19979.4512 | 104 | 19979.4512
105 | 6919.8496 | 105 | 6919.8496
105.9 | 19262.8359 | 106 | 19262.8359
106.9 | 6818.7959 | 107 | 6818.7959
107.8 | 434.8232 | 108 | 434.8232
108.6 | 18.0934 | 109 | 18.0934
109.7 | 1004.5674 | 110 | 1004.5674
110.8 | 3163.2925 | 111 | 3163.2925
112.1 | 1494.8875 | 112 | 1494.8875
113 | 2187.7019 | 113 | 2187.7019
113.7 | 65.1575 | 114 | 65.1575
114.9 | 6468.4668 | 115 | 6468.4668
115.9 | 14938.0107 | 116 | 14938.0107
116.9 | 31286.1113 | 117 | 31286.1113
118 | 17044.3555 | 118 | 17044.3555
118.9 | 28079.1094 | 119 | 28079.1094
120 | 1122.394 | 120 | 1122.394
121.3 | 1035.5314 | 121 | 1035.5314
121.9 | 652.0935 | 122 | 652.0935
123.1 | 1860.717 | 123 | 1860.717
124 | 3998.1338 | 124 | 3998.1338
125 | 2421.3159 | 125 | 2421.3159
126 | 2379.8687 | 126 | 2379.8687
127 | 1799.9745 | 127 | 1799.9745
127.9 | 20487.623 | 128 | 20487.623
129 | 9934.2295 | 129 | 9934.2295
130 | 24385.752 | 130 | 24385.752
131 | 7081.1865 | 131 | 7081.1865
131.7 | 361.816 | 132 | 361.816
132.2 | 510.0121 | 132 | 510.0121
133.2 | 757.2488 | 133 | 757.2488
133.9 | 1920.3834 | 134 | 1920.3834
134.9 | 8163.1406 | 135 | 8163.1406
136.1 | 886.7258 | 136 | 886.7258
136.9 | 6015.8032 | 137 | 6015.8032
137.7 | 174.767 | 138 | 174.767
138.7 | 566.1059 | 139 | 566.1059
139.5 | 1685.9594 | 140 | 1685.9594
140.1 | 1383.9956 | 140 | 1383.9956
140.9 | 41849.3633 | 141 | 41849.3633
142 | 6588.1792 | 142 | 6588.1792
142.9 | 47444.543 | 143 | 47444.543
143.8 | 2154.7632 | 144 | 2154.7632
145 | 5601.9653 | 145 | 5601.9653
146 | 9231.5 | 146 | 9231.5
146.9 | 3952.9766 | 147 | 3952.9766
147.9 | 17283.377 | 148 | 17283.377
148.9 | 2887.8503 | 149 | 2887.8503
149.9 | 4560.0391 | 150 | 4560.0391
151.1 | 399.439 | 151 | 399.439
151.9 | 2547.1826 | 152 | 2547.1826
153 | 4902.9502 | 153 | 4902.9502
153.9 | 4359.9619 | 154 | 4359.9619
155 | 21371.7285 | 155 | 21371.7285
155.9 | 2893.3848 | 156 | 2893.3848
156.9 | 11596.6748 | 157 | 11596.6748
157.9 | 971.5923 | 158 | 971.5923
158.5 | 542.1011 | 159 | 542.1011
159 | 4852.5566 | 159 | 4852.5566
159.8 | 795.005 | 160 | 795.005
160.9 | 5254.9487 | 161 | 5254.9487
161.8 | 593.0021 | 162 | 593.0021
162.9 | 115.4734 | 163 | 115.4734
163.9 | 106.4305 | 164 | 106.4305
164.5 | 99.1225 | 165 | 99.1225
165.2 | 2.6993 | 165 | 2.6993
166.2 | 165.2263 | 166 | 165.2263
167.3 | 374.5231 | 167 | 374.5231
167.8 | 87.3384 | 168 | 87.3384
168.6 | 0.035758 | 169 | 0.035758
169.7 | 763.7057 | 170 | 763.7057
170.3 | 292.6858 | 170 | 292.6858
170.9 | 2737.8577 | 171 | 2737.8577
172 | 18027.4414 | 172 | 18027.4414
173.9 | 7825724.5 | 174 | 7825724.5
175 | 540060.0625 | 175 | 540060.0625
175.9 | 7792182 | 176 | 7792182
177 | 477481.875 | 177 | 477481.875
178 | 16727.9473 | 178 | 16727.9473
178.9 | 176.9718 | 179 | 176.9718
179.8 | 314.774 | 180 | 314.774
180.7 | 0 | 181 | 0
181.3 | 0.008035 | 181 | 0.008035
182.5 | 191.1416 | 183 | 191.1416
183.6 | 0.010539 | 184 | 0.010539
184.6 | 27.941 | 185 | 27.941
186.5 | 0.005888 | 187 | 0.005888
187.2 | 2.8442 | 187 | 2.8442
187.9 | 2.7184 | 188 | 2.7184
188.5 | 0.000136 | 189 | 0.000136
189.3 | 97.4566 | 189 | 97.4566
189.9 | 405.3826 | 190 | 405.3826
190.6 | 34.972 | 191 | 34.972
191.8 | 259.7259 | 192 | 259.7259
192.9 | 445.7775 | 193 | 445.7775
194.1 | 74.4343 | 194 | 74.4343
194.9 | 47.4831 | 195 | 47.4831
195.5 | 62.8253 | 196 | 62.8253
196.7 | 2.3435 | 197 | 2.3435
197.3 | 0.024759 | 197 | 0.024759
199.2 | 88.1559 | 199 | 88.1559
199.9 | 0 | 200 | 0
200.8 | 0.086102 | 201 | 0.086102
201.5 | 74.9775 | 202 | 74.9775
202.5 | 3.1e-05 | 203 | 3.1e-05
203.1 | 45.4034 | 203 | 45.4034
204.4 | 178.1125 | 204 | 178.1125
205.1 | 3.793 | 205 | 3.793
206.3 | 0.016503 | 206 | 0.016503
207 | 2050.5977 | 207 | 2050.5977
207.8 | 138.532 | 208 | 138.532
208.7 | 244.0094 | 209 | 244.0094
209.2 | 35.1804 | 209 | 35.1804
210.7 | 51.915 | 211 | 51.915
212 | 0.000368 | 212 | 0.000368
212.7 | 1.8e-05 | 213 | 1.8e-05
213.7 | 0.008898 | 214 | 0.008898
214.7 | 76.8212 | 215 | 76.8212
215.6 | 0.000103 | 216 | 0.000103
216.1 | 1.5e-05 | 216 | 1.5e-05
217.1 | 0.016515 | 217 | 0.016515
218.4 | 24.5339 | 218 | 24.5339
219.1 | 20.6162 | 219 | 20.6162
219.7 | 0 | 220 | 0
221.2 | 106.6795 | 221 | 106.6795
221.8 | 142.3619 | 222 | 142.3619
223.1 | 0.015925 | 223 | 0.015925
224 | 31.6807 | 224 | 31.6807
225.1 | 0.23525 | 225 | 0.23525
225.7 | 0 | 226 | 0
226.2 | 47.0709 | 226 | 47.0709
227.6 | 0.11204 | 228 | 0.11204
229 | 0 | 229 | 0
229.8 | 2.0515 | 230 | 2.0515
230.5 | 0 | 231 | 0
231 | 0.2803 | 231 | 0.2803
232 | 341.5435 | 232 | 341.5435
233 | 4.8282 | 233 | 4.8282
233.6 | 0.2232 | 234 | 0.2232
234.4 | 6e-06 | 234 | 6e-06
235.1 | 5.4e-05 | 235 | 5.4e-05
236.4 | 36.3577 | 236 | 36.3577
237.4 | 36.0084 | 237 | 36.0084
238.5 | 0.16808 | 239 | 0.16808
239.3 | 0.1694 | 239 | 0.1694
239.8 | 0.23997 | 240 | 0.23997
240.6 | 0 | 241 | 0
241.7 | 26.9246 | 242 | 26.9246
242.4 | 0.081595 | 242 | 0.081595
243.1 | 26.5981 | 243 | 26.5981
244.2 | 0 | 244 | 0
245.3 | 7e-06 | 245 | 7e-06
246.1 | 0.00827 | 246 | 0.00827
247.1 | 0.000951 | 247 | 0.000951
248.5 | 0.25737 | 249 | 0.25737
249.1 | 41.6064 | 249 | 41.6064
250 | 108.0888 | 250 | 108.0888
251.2 | 0.11817 | 251 | 0.11817
252.3 | 57.6673 | 252 | 57.6673
253.7 | 2.6718 | 254 | 2.6718
254.6 | 0.016504 | 255 | 0.016504
255.4 | 0.000366 | 255 | 0.000366
256.2 | 0.01611 | 256 | 0.01611
256.9 | 0.003291 | 257 | 0.003291
257.6 | 82.9635 | 258 | 82.9635
258.8 | 63.283 | 259 | 63.283
259.9 | 275.1269 | 260 | 275.1269
257.5 | 0.006207 | 258 | 0.006207
258.3 | 0
259.2 | 0.023469
259.8 | 2.9e-05
257.3 | 0.004411
258.3 | 181.4352
259 | 0.043369
259.9 | 274.5332
259.9 | 122.1577

## ICAL
[12 empty] | TV | [6 empty] | Measured | [5 empty] | %TV | [6 empty] | Ret.Time | Cal.Type | Number | Rel.Std.Dev. | Coeff.of | R^2
Name | RT | Window | Quant | Conf1 | Conf2 | Window |  | Channel | Eval. Type | IS | Cal.Type | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Unit | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Level 1 | Level 2 | Level 3 | Level 4 | Level 5 | Level 6 | Name | min |  | of Points | % | Determination | Rounded | RSD<20% | R^2>0.99 | Total Analytes | 81
Chloromethane (methyl chloride) | 1.47 | 0.050 AN | 50 | 52 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0596 | 2.0495 | 5.3017 | 10.5574 | 26.3892 | 99.5803 | 105.96 | 102.475 | 106.034 | 105.574 | 105.5568 | 99.5803 | Chloromethane (methyl chloride) | 1.47 | Lin | 6 | 2.9353 | 0.99965 | 1 | TRUE | TRUE | RSD Failed | 0
Chloroethene (vinyl chloride) | 1.56 | 0.050 AN | 62 | 64 | 61 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1057 | 2.2464 | 6.0122 | 11.2445 | 27.9545 | 99.0803 | 110.57 | 112.32 | 120.244 | 112.445 | 111.818 | 99.0803 | Chloroethene (vinyl chloride) | 1.55 | Lin | 6 | 6.3532 | 0.99829 | 0.998 | TRUE | TRUE | R^2 Failed | 0
Bromomethane (methyl bromide) | 1.832 | 0.058 AN | 94 | 96 | 93 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset, 1/A | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7957 | 2.1303 | 5.958 | 9.9578 | 23.6456 | 97.2695 | 79.57 | 106.515 | 119.16 | 99.578 | 94.5824 | 97.2695 | Bromomethane (methyl bromide) | 1.85 | Quad, WithOffset, 1/A | 6 | 14.8861 | 0.99393 | 0.994 | TRUE | TRUE
Chloroethane (ethyl chloride) | 1.944 | 0.050 AN | 64 | 66 | 49 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9608 | 2.1951 | 5.6965 | 10.9028 | 26.9702 | 99.3788 | 96.08 | 109.755 | 113.93 | 109.028 | 107.8808 | 99.3788 | Chloroethane (ethyl chloride) | n.a. | Lin | 6 | 4.3489 | 0.99922 | 0.999 | TRUE | TRUE
Trichlorofluoromethane | 2.2 | 0.050 AN | 101 | 103 | 105 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0349 | 2.129 | 5.8383 | 10.7957 | 27.0158 | 99.3716 | 103.49 | 106.45 | 116.766 | 107.957 | 108.0632 | 99.3716 | Trichlorofluoromethane | 2.19 | Lin | 6 | 4.4251 | 0.99919 | 0.999 | TRUE | TRUE
Diethyl ether | 2.52 | 0.080 AN | 59 | 74 | 45 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9767 | 2.0043 | 4.7686 | 10.0343 | 25.277 | 99.939 | 97.67 | 100.215 | 95.372 | 100.343 | 101.108 | 99.939 | Diethyl ether | n.a. | Lin, WithOffset | 6 | 0.7754 | 0.99998 | 1 | TRUE | TRUE
1,1-Dichloroethene | 2.742 | 0.070 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.031 | 2.1141 | 5.5389 | 10.8106 | 27.2342 | 99.3309 | 103.1 | 105.705 | 110.778 | 108.106 | 108.9368 | 99.3309 | 1,1-Dichloroethene | 2.74 | Lin | 6 | 4.6478 | 0.99911 | 0.999 | TRUE | TRUE
Acetone | 2.87 | 0.088 AN | 43 | 58 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1.8 | 3.6 | 9 | 18 | 45 | 180 | ppb | 2.2116 | 4.206 | 9.0083 | 17.2888 | 44.5078 | 180.1775 | 122.8667 | 116.8333 | 100.0922 | 96.0489 | 98.9062 | 100.0986 | Acetone | n.a. | Lin, WithOffset | 6 | 1.3743 | 0.99994 | 1 | TRUE | TRUE
Iodomethane | 2.9 | 0.050 AN | 142 | 127 | 141 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | P-P | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1 | 2 | 5 | 10 | 25 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Iodomethane | 2.9 | P-P | 6 | 0 | 1 | 1 | TRUE | TRUE
Carbon disulfide | 2.97 | 0.050 AN | 76 | 78 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0157 | 2.1222 | 5.6852 | 10.8898 | 26.6563 | 99.4601 | 101.57 | 106.11 | 113.704 | 108.898 | 106.6252 | 99.4601 | Carbon disulfide | 2.96 | Lin | 6 | 3.8205 | 0.9994 | 0.999 | TRUE | TRUE
3-Chloropropene (allyl chloride) | 3.215 | 0.100 AN | 41 | 39 | 76 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2648 | 2.1547 | 5.03 | 9.5244 | 24.981 | 100.0451 | 126.48 | 107.735 | 100.6 | 95.244 | 99.924 | 100.0451 | 3-Chloropropene (allyl chloride) | n.a. | Lin, WithOffset | 6 | 1.2185 | 0.99996 | 1 | TRUE | TRUE
Methylene chloride (DCM) | 3.363 | 0.100 AN | 49 | 84 | 86 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0892 | 2.1848 | 5.4101 | 10.5943 | 26.2288 | 99.6083 | 108.92 | 109.24 | 108.202 | 105.943 | 104.9152 | 99.6083 | Methylene chloride (DCM) | 3.37 | Lin | 6 | 2.7593 | 0.99969 | 1 | TRUE | TRUE
trans-1,2-Dichloroethene | 3.68 | 0.100 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9389 | 1.9598 | 5.0533 | 9.8993 | 25.6065 | 99.8572 | 93.89 | 97.99 | 101.066 | 98.993 | 102.426 | 99.8572 | trans-1,2-Dichloroethene | 3.69 | Lin | 6 | 1.1938 | 0.99994 | 1 | TRUE | TRUE
Methyl tert-butyl ether (MTBE) | 3.702 | 0.100 AN | 73 | 41 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.766 | 1.624 | 4.056 | 8.4145 | 22.8568 | 100.7514 | 76.6 | 81.2 | 81.12 | 84.145 | 91.4272 | 100.7514 | Methyl tert-butyl ether (MTBE) | n.a. | Lin | 6 | 5.735 | 0.99884 | 0.999 | TRUE | TRUE
1,1-Dichloroethane | 4.192 | 0.080 AN | 63 | 65 | 83 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8119 | 1.8597 | 4.97 | 9.9814 | 25.492 | 99.885 | 81.19 | 92.985 | 99.4 | 99.814 | 101.968 | 99.885 | 1,1-Dichloroethane | 4.19 | Lin, WithOffset | 6 | 1.1648 | 0.99996 | 1 | TRUE | TRUE
2,2-Dichloropropane | 4.822 | 0.060 AN | 77 | 41 | 79 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7841 | 1.668 | 4.4873 | 8.9799 | 23.2186 | 100.5818 | 78.41 | 83.4 | 89.746 | 89.799 | 92.8744 | 100.5818 | 2,2-Dichloropropane | n.a. | Lin | 6 | 4.2823 | 0.99934 | 0.999 | TRUE | TRUE
cis-1,2-Dichloroethene | 4.83 | 0.060 AN | 61 | 96 | 98 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1214 | 2.0976 | 5.0366 | 9.8278 | 24.8726 | 100.0441 | 112.14 | 104.88 | 100.732 | 98.278 | 99.4904 | 100.0441 | cis-1,2-Dichloroethene | 4.83 | Lin, WithOffset | 6 | 0.5738 | 0.99999 | 1 | TRUE | TRUE
2-Butanone (MEK) | 4.86 | 0.060 AN | 43 | 72 | 57 | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1.8 | 3.6 | 9 | 18 | 45 | 180 | ppb | 2.5534 | 3.9834 | 9.0973 | 17.4796 | 44.0057 | 180.2805 | 141.8556 | 110.65 | 101.0811 | 97.1089 | 97.7904 | 100.1558 | 2-Butanone (MEK) | n.a. | Lin, WithOffset | 6 | 1.7332 | 0.99991 | 1 | TRUE | TRUE
Methyl acrylate | 4.953 | 0.060 AN | 55 | 85 |  | 0.2 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset, 1/A² | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0246 | 1.97 | 4.4214 | 10.252 | 27.2408 | 98.6367 | 102.46 | 98.5 | 88.428 | 102.52 | 108.9632 | 98.6367 | Methyl acrylate | n.a. | Quad, WithOffset, 1/A² | 6 | 18.9632 | 0.99382 | 0.994 | TRUE | TRUE
Bromochloromethane | 5.073 | 0.060 AN | 49 | 130 | 128 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2036 | 2.1483 | 5.0031 | 9.7529 | 24.8301 | 100.062 | 120.36 | 107.415 | 100.062 | 97.529 | 99.3204 | 100.062 | Bromochloromethane | 5.07 | Lin, WithOffset | 6 | 0.8409 | 0.99998 | 1 | TRUE | TRUE
Methacrylonitrile | 5.076 | 0.060 AN | 67 | 52 | 40 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2198 | 2.1649 | 4.7221 | 9.7369 | 25.1621 | 99.994 | 121.98 | 108.245 | 94.442 | 97.369 | 100.6484 | 99.994 | Methacrylonitrile | n.a. | Quad, WithOffset | 6 | 1.1589 | 0.99997 | 1 | TRUE | TRUE
Tetrahydrofuran | 5.1 | 0.060 AN | 42 | 72 | 71 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2408 | 2.1324 | 4.7974 | 9.7928 | 25.0145 | 100.0222 | 124.08 | 106.62 | 95.948 | 97.928 | 100.058 | 100.0222 | Tetrahydrofuran | n.a. | Lin, WithOffset | 6 | 0.8606 | 0.99998 | 1 | TRUE | TRUE
Trichloromethane (chloroform) | 5.211 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9766 | 2.0739 | 5.3353 | 10.4726 | 26.2077 | 99.6328 | 97.66 | 103.695 | 106.706 | 104.726 | 104.8308 | 99.6328 | Trichloromethane (chloroform) | 5.2 | Lin | 6 | 2.5796 | 0.99973 | 1 | TRUE | TRUE
1,1,1-Trichloroethane | 5.341 | 0.050 AN | 97 | 99 | 61 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8696 | 1.6962 | 4.9851 | 9.8108 | 26.4642 | 99.661 | 86.96 | 84.81 | 99.702 | 98.108 | 105.8568 | 99.661 | 1,1,1-Trichloroethane | 5.34 | Lin | 6 | 2.8995 | 0.99967 | 1 | TRUE | TRUE
Dibromofluoromethane [SS1] | 5.366 | 0.060 AN | 113 | 111 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 19.0753 | 19.8213 | 20.2388 | 19.8867 | 20.476 | 20.5018 | 95.3765 | 99.1065 | 101.194 | 99.4335 | 102.38 | 102.509 | Dibromofluoromethane [SS1] | 5.36 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
Pentafluorobenzene [IS1] | 5.431 | 0.060 AN | 168 | 99 |  | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Pentafluorobenzene [IS1] | 5.42 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
1-Chlorobutane (butyl chloride) | 5.486 | 0.040 AN | 56 | 41 | 43 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9915 | 1.9309 | 5.1058 | 9.955 | 25.0216 | 99.9953 | 99.15 | 96.545 | 102.116 | 99.55 | 100.0864 | 99.9953 | 1-Chlorobutane (butyl chloride) | 5.48 | Lin, WithOffset | 6 | 0.2871 | 1 | 1 | TRUE | TRUE
Carbon tetrachloride | 5.501 | 0.040 AN | 119 | 121 |  | 0.15 | RT | All Channels | Area | Internal Pentafluorobenzene [IS1] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1966 | 2.1538 | 5.0614 | 9.6497 | 24.8821 | 100.0564 | 119.66 | 107.69 | 101.228 | 96.497 | 99.5284 | 100.0564 | Carbon tetrachloride | n.a. | Lin, WithOffset | 6 | 0.9686 | 0.99997 | 1 | TRUE | TRUE
1,1-Dichloropropene | 5.519 | 0.040 AN | 75 | 77 | 110 | 0.15 | RT | All Channels | Area | Internal 1,4-Difluorobenzene [IS2] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9223 | 1.9309 | 5.3023 | 10.3965 | 26.1995 | 99.6475 | 92.23 | 96.545 | 106.046 | 103.965 | 104.798 | 99.6475 | 1,1-Dichloropropene | 5.51 | Lin | 6 | 2.5086 | 0.99975 | 1 | TRUE | TRUE
Benzene | 5.702 | 0.060 AN | 78 | 77 | 52 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9267 | 1.7856 | 4.9319 | 9.8051 | 27.7739 | 99.3344 | 92.67 | 89.28 | 98.638 | 98.051 | 111.0956 | 99.3344 | Benzene | 5.71 | Lin | 6 | 5.3255 | 0.99887 | 0.999 | TRUE | TRUE
1,2-Dichloroethane | 5.788 | 0.060 AN | 62 | 64 | 49 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2398 | 2.0864 | 4.9287 | 9.5133 | 25.2478 | 99.9835 | 123.98 | 104.32 | 98.574 | 95.133 | 100.9912 | 99.9835 | 1,2-Dichloroethane | 5.78 | Quad, WithOffset | 6 | 1.7208 | 0.99992 | 1 | TRUE | TRUE
1,4-Difluorobenzene [IS2] | 6.181 | 0.080 AN | 114 | 88 | 63 | 0.15 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Difluorobenzene [IS2] | 6.17 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
Trichloroethene | 6.399 | 0.080 AN | 130 | 132 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0533 | 2.0009 | 5.4266 | 10.2963 | 27.7385 | 99.2639 | 105.33 | 100.045 | 108.532 | 102.963 | 110.954 | 99.2639 | Trichloroethene | 6.39 | Lin | 6 | 5.3073 | 0.99885 | 0.999 | TRUE | TRUE
1,2-Dichloropropane | 6.652 | 0.060 AN | 63 | 62 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9164 | 1.7753 | 4.8596 | 9.8486 | 26.7342 | 99.5939 | 91.64 | 88.765 | 97.192 | 98.486 | 106.9368 | 99.5939 | 1,2-Dichloropropane | n.a. | Lin | 6 | 3.3772 | 0.99955 | 1 | TRUE | TRUE
Dibromomethane | 6.729 | 0.060 AN | 174 | 93 | 95 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0225 | 2.0018 | 5.1761 | 10.238 | 27.6494 | 99.3048 | 102.25 | 100.09 | 103.522 | 102.38 | 110.5976 | 99.3048 | Dibromomethane | 6.73 | Lin | 6 | 5.0848 | 0.99895 | 0.999 | TRUE | TRUE
Methyl methacrylate | 6.744 | 0.060 AN | 41 | 69 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.4923 | 2.2466 | 4.7104 | 9.1443 | 25.4207 | 99.9846 | 149.23 | 112.33 | 94.208 | 91.443 | 101.6828 | 99.9846 | Methyl methacrylate | n.a. | Quad, WithOffset | 6 | 2.7055 | 0.99985 | 1 | TRUE | TRUE
Bromodichloromethane | 6.925 | 0.060 AN | 83 | 85 | 47 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9149 | 1.7453 | 4.6534 | 9.7285 | 26.6907 | 99.6277 | 91.49 | 87.265 | 93.068 | 97.285 | 106.7628 | 99.6277 | Bromodichloromethane | 6.92 | Lin | 6 | 3.381 | 0.99956 | 1 | TRUE | TRUE
2-Nitropropane | 7.144 | 0.080 AN | 43 | 41 | 39 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8504 | 1.6871 | 4.0587 | 8.7339 | 23.4579 | 100.567 | 85.04 | 84.355 | 81.174 | 87.339 | 93.8316 | 100.567 | 2-Nitropropane | n.a. | Lin | 6 | 4.4366 | 0.99929 | 0.999 | TRUE | TRUE
cis-1,3-Dichloropropene | 7.361 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.258 | 2.0644 | 4.9471 | 9.532 | 25.2045 | 99.9936 | 125.8 | 103.22 | 98.942 | 95.32 | 100.818 | 99.9936 | cis-1,3-Dichloropropene | 7.36 | Quad, WithOffset | 6 | 1.2697 | 0.99997 | 1 | TRUE | TRUE
4-Methyl-2-pentanone (MIBK) | 7.532 | 0.049 AN | 43 | 58 | 41 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1.8 | 3.6 | 9 | 18 | 45 | 180 | ppb | 1.4347 | 2.8548 | 7.5167 | 15.9172 | 42.7902 | 180.8534 | 79.7056 | 79.3 | 83.5189 | 88.4289 | 95.0893 | 100.4741 | 4-Methyl-2-pentanone (MIBK) | n.a. | Lin | 6 | 3.8249 | 0.99947 | 0.999 | TRUE | TRUE
Toluene-d8 [SS2] | 7.62 | 0.058 AN | 98 | 100 | 70 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20.0076 | 19.3667 | 20.0338 | 20.1539 | 20.4641 | 19.9738 | 100.038 | 96.8335 | 100.169 | 100.7695 | 102.3205 | 99.869 | Toluene-d8 [SS2] | 7.61 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
Toluene | 7.684 | 0.047 AN | 91 | 92 | 65 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9243 | 1.7822 | 5.0056 | 9.6344 | 26.2194 | 99.7365 | 92.43 | 89.11 | 100.112 | 96.344 | 104.8776 | 99.7365 | Toluene | 7.68 | Lin | 6 | 2.4722 | 0.99976 | 1 | TRUE | TRUE
trans-1,3-Dichloropropene | 7.931 | 0.060 AN | 75 | 39 | 77 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.2861 | 2.0581 | 4.9413 | 9.5063 | 25.2142 | 99.9936 | 128.61 | 102.905 | 98.826 | 95.063 | 100.8568 | 99.9936 | trans-1,3-Dichloropropene | 7.92 | Quad, WithOffset | 6 | 1.3244 | 0.99997 | 1 | TRUE | TRUE
Ethyl methacrylate | 8 | 0.060 AN | 69 | 41 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1895 | 1.991 | 4.841 | 9.9118 | 25.0691 | 99.9975 | 118.95 | 99.55 | 96.82 | 99.118 | 100.2764 | 99.9975 | Ethyl methacrylate | 8 | Quad, WithOffset | 6 | 0.5914 | 0.99999 | 1 | TRUE | TRUE
1,1,2-Trichloroethane | 8.11 | 0.060 AN | 97 | 83 | 99 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0776 | 1.9255 | 4.7935 | 9.6766 | 25.6444 | 99.8823 | 107.76 | 96.275 | 95.87 | 96.766 | 102.5776 | 99.8823 | 1,1,2-Trichloroethane | 8.1 | Lin, WithOffset | 6 | 1.6182 | 0.99992 | 1 | TRUE | TRUE
Tetrachloroethene | 8.164 | 0.060 AN | 166 | 164 | 129 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9128 | 1.8985 | 5.2574 | 9.965 | 24.9637 | 100.003 | 91.28 | 94.925 | 105.148 | 99.65 | 99.8548 | 100.003 | Tetrachloroethene | 8.15 | Quad, WithOffset | 6 | 0.819 | 0.99998 | 1 | TRUE | TRUE
1,3-Dichloropropane | 8.25 | 0.060 AN | 76 | 41 | 78 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.3685 | 2.1834 | 5.0956 | 9.6713 | 24.5473 | 100.1339 | 136.85 | 109.17 | 101.912 | 96.713 | 98.1892 | 100.1339 | 1,3-Dichloropropane | 8.25 | Lin, WithOffset | 6 | 1.5294 | 0.99993 | 1 | TRUE | TRUE
2-Hexanone | 8.319 | 0.050 AN | 43 | 58 | 57 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1.8 | 3.6 | 9 | 18 | 45 | 180 | ppb | 1.3669 | 2.6273 | 7.3582 | 16.0337 | 42.3872 | 180.9557 | 75.9389 | 72.9806 | 81.7578 | 89.0761 | 94.1938 | 100.5309 | 2-Hexanone | n.a. | Lin | 6 | 4.2046 | 0.99937 | 0.999 | TRUE | TRUE
Dibromochloromethane | 8.435 | 0.050 AN | 129 | 127 | 131 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8256 | 1.5703 | 4.2579 | 8.6786 | 23.8149 | 100.4759 | 82.56 | 78.515 | 85.158 | 86.786 | 95.2596 | 100.4759 | Dibromochloromethane | 8.42 | Lin | 6 | 3.9117 | 0.99945 | 0.999 | TRUE | TRUE
1,2-Dibromoethane (EDB) | 8.52 | 0.049 AN | 107 | 109 | 93 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.957 | 1.7831 | 4.7996 | 9.6913 | 25.9644 | 99.8045 | 95.7 | 89.155 | 95.992 | 96.913 | 103.8576 | 99.8045 | 1,2-Dibromoethane (EDB) | 8.52 | Lin, WithOffset | 6 | 2.2573 | 0.99984 | 1 | TRUE | TRUE
Chlorobenzene-d5 [IS3] | 8.92 | 0.040 AF | 117 | 82 | 52 | 0.1 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | Chlorobenzene-d5 [IS3] | 8.91 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
Chlorobenzene | 8.94 | 0.040 AN | 112 | 77 | 114 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8364 | 1.7311 | 4.8454 | 9.6657 | 26.1643 | 99.7571 | 83.64 | 86.555 | 96.908 | 96.657 | 104.6572 | 99.7571 | Chlorobenzene | 8.94 | Lin, WithOffset | 6 | 2.6791 | 0.99978 | 1 | TRUE | TRUE
1,1,1,2-Tetrachloroethane | 9.02 | 0.040 AN | 131 | 133 | 117 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.3983 | 2.1825 | 4.8917 | 9.3146 | 25.1956 | 100.0174 | 139.83 | 109.125 | 97.834 | 93.146 | 100.7824 | 100.0174 | 1,1,1,2-Tetrachloroethane | 9.02 | Lin, WithOffset | 6 | 1.8155 | 0.9999 | 1 | TRUE | TRUE
Ethylbenzene | 9.03 | 0.040 AN | 91 | 106 | 51 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.978 | 1.9338 | 5.3095 | 10.4683 | 27.6817 | 99.2688 | 97.8 | 96.69 | 106.19 | 104.683 | 110.7268 | 99.2688 | Ethylbenzene | 9.02 | Lin | 6 | 5.2261 | 0.99889 | 0.999 | TRUE | TRUE
m/p-Xylene | 9.138 | 0.060 AN | 91 | 106 | 105 | 0.1 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset, 1/R | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0799 | 1.882 | 4.9725 | 9.5943 | 25.6507 | 99.8627 | 107.99 | 94.1 | 99.45 | 95.943 | 102.6028 | 99.8627 | m/p-Xylene | 9.14 | Quad, WithOffset, 1/R | 6 | 7.6799 | 0.99955 | 1 | TRUE | TRUE
o-Xylene | 9.435 | 0.060 AN | 91 | 106 | 105 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9514 | 1.8813 | 5.2944 | 10.2477 | 27.4759 | 99.3444 | 95.14 | 94.065 | 105.888 | 102.477 | 109.9036 | 99.3444 | o-Xylene | 9.43 | Lin | 6 | 4.7923 | 0.99907 | 0.999 | TRUE | TRUE
Styrene | 9.451 | 0.060 AN | 104 | 78 | 103 | 0.15 | RT | All Channels | Area | Internal Chlorobenzene-d5 [IS3] | Quad, WithOffset, 1/R² | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.0312 | 1.8764 | 4.9332 | 9.9362 | 27.441 | 98.2587 | 103.12 | 93.82 | 98.664 | 99.362 | 109.764 | 98.2587 | Styrene | 9.45 | Quad, WithOffset, 1/R² | 6 | 10.8325 | 0.99573 | 0.996 | TRUE | TRUE
Bromoform | 9.583 | 0.060 AN | 173 | 171 | 175 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8504 | 1.6782 | 4.405 | 9.159 | 25.5636 | 99.9791 | 85.04 | 83.91 | 88.1 | 91.59 | 102.2544 | 99.9791 | Bromoform | 9.57 | Quad | 6 | 2.3912 | 0.99984 | 1 | TRUE | TRUE
Isopropylbenzene (cumene) | 9.72 | 0.060 AN | 105 | 120 | 79 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8791 | 1.7384 | 4.8843 | 9.747 | 26.7889 | 99.5903 | 87.91 | 86.92 | 97.686 | 97.47 | 107.1556 | 99.5903 | Isopropylbenzene (cumene) | 9.71 | Lin | 6 | 3.5093 | 0.99952 | 1 | TRUE | TRUE
1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.843 | 0.060 AN | 95 | 174 | 176 | 0.15 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20.094 | 19.726 | 20.1035 | 19.5009 | 20.2717 | 20.3039 | 100.47 | 98.63 | 100.5175 | 97.5045 | 101.3585 | 101.5195 | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
Bromobenzene | 9.943 | 0.040 AN | 77 | 156 | 158 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9233 | 1.8404 | 4.6043 | 9.1711 | 24.1484 | 100.3195 | 92.33 | 92.02 | 92.086 | 91.711 | 96.5936 | 100.3195 | Bromobenzene | 9.94 | Lin | 6 | 2.4829 | 0.99977 | 1 | TRUE | TRUE
1,1,2,2-Tetrachloroethane | 9.954 | 0.040 AN | 83 | 85 | 95 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.3685 | 2.1721 | 4.6741 | 9.5305 | 25.2646 | 99.9898 | 136.85 | 108.605 | 93.482 | 95.305 | 101.0584 | 99.9898 | 1,1,2,2-Tetrachloroethane | 9.95 | Quad, WithOffset | 6 | 1.7866 | 0.99993 | 1 | TRUE | TRUE
1,2,3-Trichloropropane (TCP) | 9.988 | 0.040 AN | 77 | 110 | 61 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9063 | 1.7782 | 4.4759 | 9.3563 | 24.3009 | 100.2707 | 90.63 | 88.91 | 89.518 | 93.563 | 97.2036 | 100.2707 | trans-1,4-Dichloro-2-butene | 9.98 | Lin | 6 | 3.8251 | 0.99947 | 0.999 | TRUE | TRUE
trans-1,4-Dichloro-2-butene | 9.988 | 0.040 AN | 75 | 53 | 89 | 0.05 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8646 | 1.7886 | 4.3838 | 9.2912 | 23.3344 | 100.5237 | 86.46 | 89.43 | 87.676 | 92.912 | 93.3376 | 100.5237 | 1,2,3-Trichloropropane (TCP) | 9.98 | Lin | 6 | 2.1759 | 0.99982 | 1 | TRUE | TRUE
n-Propylbenzene | 10.023 | 0.026 AN | 91 | 120 | 65 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.903 | 1.835 | 5.0899 | 10.0986 | 26.9989 | 99.4902 | 90.3 | 91.75 | 101.798 | 100.986 | 107.9956 | 99.4902 | n-Propylbenzene | 10.02 | Lin | 6 | 3.8574 | 0.99941 | 0.999 | TRUE | TRUE
2-Chlorotoluene | 10.08 | 0.031 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9502 | 1.8937 | 5.0228 | 10.1616 | 26.232 | 99.6773 | 95.02 | 94.685 | 100.456 | 101.616 | 104.928 | 99.6773 | 2-Chlorotoluene | 10.08 | Lin | 6 | 2.4036 | 0.99977 | 1 | TRUE | TRUE
1,3,5-Trimethylbenzene | 10.159 | 0.040 AN | 105 | 120 | 119 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8098 | 1.6842 | 4.7123 | 9.6507 | 27.0997 | 99.5326 | 80.98 | 84.21 | 94.246 | 96.507 | 108.3988 | 99.5326 | 1,3,5-Trimethylbenzene | 10.15 | Lin | 6 | 4.1681 | 0.99932 | 0.999 | TRUE | TRUE
4-Chlorotoluene | 10.176 | 0.040 AN | 91 | 126 | 89 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9027 | 1.8571 | 5.0399 | 10.2445 | 28.1153 | 99.1986 | 90.27 | 92.855 | 100.798 | 102.445 | 112.4612 | 99.1986 | 4-Chlorotoluene | 10.17 | Lin | 6 | 5.9642 | 0.99856 | 0.999 | TRUE | TRUE
tert-Butylbenzene | 10.38 | 0.040 AN | 119 | 91 | 134 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.1229 | 1.9443 | 4.8407 | 9.5391 | 25.6654 | 99.8876 | 112.29 | 97.215 | 96.814 | 95.391 | 102.6616 | 99.8876 | tert-Butylbenzene | 10.37 | Lin, WithOffset | 6 | 1.7954 | 0.9999 | 1 | TRUE | TRUE
Pentachloroethane | 10.398 | 0.030 AN | 167 | 130 | 132 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | P-P | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1 | 2 | 5 | 10 | 25 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | Pentachloroethane | 10.39 | P-P | 6 | 0 | 1 | 1 | TRUE | TRUE
1,2,4-Trimethylbenzene | 10.422 | 0.030 AN | 105 | 120 | 77 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8427 | 1.7493 | 4.6722 | 9.9366 | 27.3264 | 99.4477 | 84.27 | 87.465 | 93.444 | 99.366 | 109.3056 | 99.4477 | 1,2,4-Trimethylbenzene | 10.42 | Lin | 6 | 4.5331 | 0.99919 | 0.999 | TRUE | TRUE
1-Methylpropylbenzene (sec-butylbenzene) | 10.535 | 0.037 AN | 105 | 134 | 91 | 0.08 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8573 | 1.7502 | 4.9226 | 10.2691 | 27.8779 | 99.2639 | 85.73 | 87.51 | 98.452 | 102.691 | 111.5116 | 99.2639 | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | Lin | 6 | 5.5494 | 0.99876 | 0.999 | TRUE | TRUE
1,3-Dichlorobenzene | 10.61 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9182 | 1.8754 | 4.8104 | 9.9099 | 26.4803 | 99.6517 | 91.82 | 93.77 | 96.208 | 99.099 | 105.9212 | 99.6517 | 1,3-Dichlorobenzene | 10.61 | Lin | 6 | 2.8812 | 0.99967 | 1 | TRUE | TRUE
4-Isopropyltoluene (p-cymene) | 10.631 | 0.040 AN | 119 | 91 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8486 | 1.7541 | 4.9002 | 10.2467 | 27.9722 | 99.2437 | 84.86 | 87.705 | 98.004 | 102.467 | 111.8888 | 99.2437 | 4-Isopropyltoluene (p-cymene) | 10.64 | Lin | 6 | 5.7232 | 0.99868 | 0.999 | TRUE | TRUE
1,4-Dichlorobenzene-d4 [IS4] | 10.67 | 0.036 AN | 152 | 115 |  | 0.07 | RT | All Channels | Area | ISTD Internal | Lin, Avg | 20 | 20 | 20 | 20 | 20 | 20 | ppb | 20 | 20 | 20 | 20 | 20 | 20 | 100 | 100 | 100 | 100 | 100 | 100 | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | Lin, Avg | 1 | n.a. | n.a. | n.a. | TRUE | TRUE
1,4-Dichlorobenzene | 10.687 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9774 | 1.9695 | 5.0125 | 10.1537 | 26.4326 | 99.6267 | 97.74 | 98.475 | 100.25 | 101.537 | 105.7304 | 99.6267 | 1,4-Dichlorobenzene | 10.68 | Lin | 6 | 2.7712 | 0.9997 | 1 | TRUE | TRUE
n-Butylbenzene | 10.923 | 0.040 AN | 91 | 92 | 134 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7999 | 1.6216 | 4.5577 | 9.5019 | 26.5797 | 99.6866 | 79.99 | 81.08 | 91.154 | 95.019 | 106.3188 | 99.6866 | n-Butylbenzene | 10.91 | Lin | 6 | 3.3734 | 0.99956 | 1 | TRUE | TRUE
1,2-Dichlorobenzene | 10.93 | 0.040 AN | 146 | 148 | 111 | 0.07 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.9652 | 1.8435 | 4.815 | 9.7729 | 27.2243 | 99.4794 | 96.52 | 92.175 | 96.3 | 97.729 | 108.8972 | 99.4794 | 1,2-Dichlorobenzene | 10.92 | Lin | 6 | 4.299 | 0.99927 | 0.999 | TRUE | TRUE
Hexachloroethane | 11.11 | 0.040 AN | 117 | 119 | 201 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.7797 | 1.5838 | 4.3998 | 9.2416 | 25.5386 | 99.9799 | 77.97 | 79.19 | 87.996 | 92.416 | 102.1544 | 99.9799 | Hexachloroethane | n.a. | Quad | 6 | 2.3333 | 0.99985 | 1 | TRUE | TRUE
1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 0.040 AN | 157 | 155 | 75 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad, WithOffset | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.4278 | 2.167 | 4.642 | 9.4901 | 25.2821 | 99.9901 | 142.78 | 108.35 | 92.84 | 94.901 | 101.1284 | 99.9901 | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | Quad, WithOffset | 6 | 1.8839 | 0.99993 | 1 | TRUE | TRUE
Nitrobenzene | 11.58 | 0.040 AN | 77 | 51 | 123 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8314 | 1.7797 | 3.9966 | 9.792 | 25.3579 | 99.9859 | 83.14 | 88.985 | 79.932 | 97.92 | 101.4316 | 99.9859 | Nitrobenzene | n.a. | Quad | 6 | 2.1082 | 0.99988 | 1 | TRUE | TRUE
1,2,4-Trichlorobenzene | 11.98 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8729 | 1.7774 | 4.5163 | 9.3506 | 25.4421 | 99.9833 | 87.29 | 88.87 | 90.326 | 93.506 | 101.7684 | 99.9833 | 1,2,4-Trichlorobenzene | 11.98 | Quad | 6 | 1.8814 | 0.9999 | 1 | TRUE | TRUE
Hexachloro-1,3-butadiene | 12.07 | 0.040 AN | 225 | 227 | 223 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8303 | 1.6838 | 4.5447 | 8.8821 | 23.9285 | 100.4105 | 83.03 | 84.19 | 90.894 | 88.821 | 95.714 | 100.4105 | Hexachloro-1,3-butadiene | 12.07 | Lin | 6 | 3.2588 | 0.99961 | 1 | TRUE | TRUE
Naphthalene | 12.153 | 0.040 AN | 128 | 127 | 129 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Quad, WithOffset, 1/R² | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 1.024 | 1.936 | 4.6082 | 10.4252 | 27.7079 | 97.9225 | 102.4 | 96.8 | 92.164 | 104.252 | 110.8316 | 97.9225 | Naphthalene | 12.15 | Quad, WithOffset, 1/R² | 6 | 13.8202 | 0.99317 | 0.993 | TRUE | TRUE
1,2,3-Trichlorobenzene | 12.29 | 0.040 AN | 180 | 182 | 145 | 0.1 | RT | All Channels | Area | Internal 1,4-Dichlorobenzene-d4 [IS4] | Lin | 1 | 2 | 5 | 10 | 25 | 100 | ppb | 0.8126 | 1.5888 | 4.0594 | 8.5662 | 23.2424 | 100.6399 | 81.26 | 79.44 | 81.188 | 85.662 | 92.9696 | 100.6399 | 1,2,3-Trichlorobenzene | 12.29 | Lin | 6 | 4.9787 | 0.99912 | 0.999 | TRUE | TRUE

## MRL
Name | MRL
Chloromethane (methyl chloride) | 1
Chloroethene (vinyl chloride) | 1
Bromomethane (methyl bromide) | 1
Chloroethane (ethyl chloride) | 1
Trichlorofluoromethane | 1
Diethyl ether | 1
1,1-Dichloroethene | 1
Acetone | 1.8
Iodomethane | 1
Carbon disulfide | 1
3-Chloropropene (allyl chloride) | 1
Methylene chloride (DCM) | 1
trans-1,2-Dichloroethene | 1
Methyl tert-butyl ether (MTBE) | 1
1,1-Dichloroethane | 1
2,2-Dichloropropane | 1
cis-1,2-Dichloroethene | 1
2-Butanone (MEK) | 1.8
Methyl acrylate | 1
Bromochloromethane | 1
Methacrylonitrile | 1
Tetrahydrofuran | 1
Trichloromethane (chloroform) | 1
1,1,1-Trichloroethane | 1
Dibromofluoromethane [SS1] | 20
Pentafluorobenzene [IS1] | 20
1-Chlorobutane (butyl chloride) | 1
Carbon tetrachloride | 1
1,1-Dichloropropene | 1
Benzene | 1
1,2-Dichloroethane | 1
1,4-Difluorobenzene [IS2] | 20
Trichloroethene | 1
1,2-Dichloropropane | 1
Dibromomethane | 1
Methyl methacrylate | 1
Bromodichloromethane | 1
2-Nitropropane | 1
cis-1,3-Dichloropropene | 1
4-Methyl-2-pentanone (MIBK) | 1.8
Toluene-d8 [SS2] | 20
Toluene | 1
trans-1,3-Dichloropropene | 1
Ethyl methacrylate | 1
1,1,2-Trichloroethane | 1
Tetrachloroethene | 1
1,3-Dichloropropane | 1
2-Hexanone | 1.8
Dibromochloromethane | 1
1,2-Dibromoethane (EDB) | 1
Chlorobenzene-d5 [IS3] | 20
Chlorobenzene | 1
1,1,1,2-Tetrachloroethane | 1
Ethylbenzene | 1
m/p-Xylene | 1
o-Xylene | 1
Styrene | 1
Bromoform | 1
Isopropylbenzene (cumene) | 1
1-Bromo-4-fluorobenzene (BFB) [SS3] | 20
Bromobenzene | 1
1,1,2,2-Tetrachloroethane | 1
trans-1,4-Dichloro-2-butene | 1
1,2,3-Trichloropropane (TCP) | 1
n-Propylbenzene | 1
2-Chlorotoluene | 1
1,3,5-Trimethylbenzene | 1
4-Chlorotoluene | 1
tert-Butylbenzene | 1
Pentachloroethane | 1
1,2,4-Trimethylbenzene | 1
1-Methylpropylbenzene (sec-butylbenzene) | 1
1,3-Dichlorobenzene | 1
4-Isopropyltoluene (p-cymene) | 1
1,4-Dichlorobenzene-d4 [IS4] | 20
1,4-Dichlorobenzene | 1
n-Butylbenzene | 1
1,2-Dichlorobenzene | 1
Hexachloroethane | 1
1,2-Dibromo-3-chloropropane (DBCP) | 1
Nitrobenzene | 1
1,2,4-Trichlorobenzene | 1
Hexachloro-1,3-butadiene | 1
Naphthalene | 1
1,2,3-Trichlorobenzene | 1

## Blank
|  |  |  | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
[5 empty] | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
MRL | <1/2LLOQ | Non-target | Fail? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1 | TRUE | TRUE | FALSE | Chloromethane (methyl chloride) | 1.47 | 76 | 0 | 0.036 | Not confirmed | 50 | 52 | 32.87 | n.a. | Not confirmed | 49 | 9.87 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Chloroethene (vinyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 31.53 | n.a. | Not confirmed | 61 | 7.81 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Bromomethane (methyl bromide) | 1.85 | 858 | 0.04 | n.a. | Confirmed | 94 | 96 | 97.36 | 92.42 | Confirmed | 93 | 20.34 | 19.65 | Confirmed
1 | TRUE | TRUE | FALSE | Chloroethane (ethyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 64 | 66 | 32.12 | n.a. | Not confirmed | 49 | 23.08 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichlorofluoromethane | 2.2 | 14 | 0 | 0.002 | Not confirmed | 101 | 103 | 64.84 | n.a. | Not confirmed | 105 | 10.5 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Diethyl ether | n.a. | n.a. | n.a. | n.a. | Not confirmed | 59 | 74 | 75.46 | n.a. | Not confirmed | 45 | 72.36 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethene | 2.76 | 19 | 0 | 0.003 | Not confirmed | 61 | 96 | 72.45 | n.a. | Not confirmed | 98 | 46.37 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | Acetone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 40 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | FALSE | FALSE | Iodomethane | 2.9 | 745 | 0.04 | 0.244 | Confirmed | 142 | 127 | 31 | 26.3 | Confirmed | 141 | 13.37 | 9.42 | Confirmed
1 | TRUE | TRUE | FALSE | Carbon disulfide | 2.95 | 482 | 0.02 | 0.028 | Not confirmed | 76 | 78 | 8.96 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 3-Chloropropene (allyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 39 | 52.2 | n.a. | Not confirmed | 76 | 36.97 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Methylene chloride (DCM) | 3.36 | 361 | 0.02 | 0.044 | Confirmed | 49 | 84 | 87.31 | 100.22 | Confirmed | 86 | 55.97 | 72.74 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,2-Dichloroethene | 3.68 | 54 | 0 | 0.006 | Not confirmed | 61 | 96 | 75.67 | 125.63 | Not confirmed | 98 | 47.1 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl tert-butyl ether (MTBE) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 73 | 41 | 33.25 | n.a. | Not confirmed | 57 | 23.18 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 65 | 31.98 | n.a. | Not confirmed | 83 | 12.43 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,2-Dichloroethene | 4.82 | 28 | 0 | 0.222 | Not confirmed | 61 | 96 | 80.14 | 166.61 | Not confirmed | 98 | 51.39 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 41 | 101.06 | n.a. | Not confirmed | 79 | 32.22 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Butanone (MEK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 72 | 28.44 | n.a. | Not confirmed | 57 | 7.2 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl acrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 55 | 85 | 16.75 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | Bromochloromethane | 5.05 | 24 | 0 | 0.262 | Not confirmed | 49 | 130 | 82.37 | 84.27 | Not confirmed | 128 | 64.27 | 75.12 | Not confirmed
1 | TRUE | TRUE | FALSE | Methacrylonitrile | n.a. | n.a. | n.a. | n.a. | Not confirmed | 67 | 52 | 30.53 | n.a. | Not confirmed | 40 | 35.3 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Tetrahydrofuran | n.a. | n.a. | n.a. | n.a. | Not confirmed | 42 | 72 | 42.66 | n.a. | Not confirmed | 71 | 40.38 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Trichloromethane (chloroform) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 63.95 | n.a. | Not confirmed | 47 | 17.3 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,1-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 99 | 63.95 | n.a. | Not confirmed | 61 | 48.19 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Dibromofluoromethane [SS1] | 5.36 | 101846 | 4.82 | 19.744 | Confirmed | 113 | 111 | 104.17 | 102.79 | Confirmed | n.a. | n.a. | n.a. | n.a.
20 | TRUE | FALSE | FALSE | Pentafluorobenzene [IS1] | 5.43 | 220572 | 10.44 | 20 | Confirmed | 168 | 99 | 48.98 | 48.06 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1-Chlorobutane (butyl chloride) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 56 | 41 | 59.56 | n.a. | Not confirmed | 43 | 26.98 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Carbon tetrachloride | n.a. | n.a. | n.a. | n.a. | Not confirmed | 119 | 121 | 31.25 | n.a. | Not confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,1-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 77 | 31.46 | n.a. | Not confirmed | 110 | 41.64 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Benzene | 5.71 | 189 | 0.01 | 0.007 | Not confirmed | 78 | 77 | 24.33 | n.a. | Not confirmed | 52 | 14.88 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 62 | 64 | 31.81 | n.a. | Not confirmed | 49 | 32.62 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | 1,4-Difluorobenzene [IS2] | 6.17 | 346693 | 16.42 | 20 | Confirmed | 114 | 88 | 18.42 | 18.45 | Confirmed | 63 | 18.27 | 18.04 | Confirmed
1 | TRUE | TRUE | FALSE | Trichloroethene | 6.39 | 70 | 0 | 0.01 | Not confirmed | 130 | 132 | 96.07 | 108.12 | Confirmed | 95 | 97.35 | 165.68 | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dichloropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 63 | 62 | 69.44 | n.a. | Not confirmed | 41 | 41.97 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromomethane | 6.73 | 33 | 0 | 0.007 | Not confirmed | 174 | 93 | 91.77 | n.a. | Not confirmed | 95 | 78.67 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Methyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 41 | 69 | 88.3 | n.a. | Not confirmed | 39 | 41.9 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Bromodichloromethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 62.95 | n.a. | Not confirmed | 47 | 15.49 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 2-Nitropropane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 41 | 87.28 | n.a. | Not confirmed | 39 | 28.76 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | cis-1,3-Dichloropropene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 39 | 39.82 | n.a. | Not confirmed | 77 | 31.39 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 4-Methyl-2-pentanone (MIBK) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 39.86 | n.a. | Not confirmed | 41 | 22.34 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Toluene-d8 [SS2] | 7.61 | 482048 | 22.82 | 19.405 | Confirmed | 98 | 100 | 62 | 63.05 | Confirmed | 70 | 10.79 | 10.85 | Confirmed
1 | TRUE | FALSE | FALSE | Toluene | 7.67 | 479 | 0.02 | 0.015 | Confirmed | 91 | 92 | 55.73 | 56 | Confirmed | 65 | 11.33 | 8.68 | Confirmed
1 | TRUE | TRUE | FALSE | trans-1,3-Dichloropropene | 7.94 | 83 | 0 | 0.479 | Not confirmed | 75 | 39 | 41.41 | n.a. | Not confirmed | 77 | 31.34 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethyl methacrylate | n.a. | n.a. | n.a. | n.a. | Not confirmed | 69 | 41 | 61.12 | n.a. | Not confirmed | 99 | 24.12 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,1,2-Trichloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 97 | 83 | 86.1 | n.a. | Not confirmed | 99 | 61.03 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | Tetrachloroethene | 8.16 | 149 | 0.01 | n.a. | Confirmed | 166 | 164 | 78.43 | 85.26 | Confirmed | 129 | 72.49 | 70.24 | Confirmed
1 | FALSE | TRUE | FALSE | 1,3-Dichloropropane | 8.24 | 68 | 0 | 0.503 | Not confirmed | 76 | 41 | 62.02 | n.a. | Not confirmed | 78 | 32.22 | n.a. | Not confirmed
1.8 | TRUE | TRUE | FALSE | 2-Hexanone | n.a. | n.a. | n.a. | n.a. | Not confirmed | 43 | 58 | 55.88 | n.a. | Not confirmed | 57 | 19.46 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Dibromochloromethane | 8.42 | 21 | 0 | 0.003 | Not confirmed | 129 | 127 | 77.94 | 145.15 | Not confirmed | 131 | 24.25 | 60.18 | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromoethane (EDB) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 107 | 109 | 94.8 | n.a. | Not confirmed | 93 | 4.53 | n.a. | Not confirmed
20 | TRUE | FALSE | FALSE | Chlorobenzene-d5 [IS3] | 8.91 | 354069 | 16.76 | 20 | Confirmed | 117 | 82 | 58.94 | 56.46 | Confirmed | 52 | 13.5 | 12.53 | Confirmed
1 | TRUE | TRUE | FALSE | Chlorobenzene | 8.93 | 408 | 0.02 | n.a. | Not confirmed | 112 | 77 | 62.04 | 537.89 | Not confirmed | 114 | 31.57 | 29.09 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,1,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 131 | 133 | 95.48 | n.a. | Not confirmed | 117 | 79 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Ethylbenzene | 9.03 | 858 | 0.04 | 0.029 | Not confirmed | 91 | 106 | 35.7 | 32.36 | Confirmed | 51 | 8.99 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | m/p-Xylene | 9.13 | 1713 | 0.08 | 0.294 | Confirmed | 91 | 106 | 50.83 | 53.33 | Confirmed | 105 | 21.8 | 20.54 | Confirmed
1 | TRUE | FALSE | FALSE | o-Xylene | 9.43 | 598 | 0.03 | 0.022 | Confirmed | 91 | 106 | 50.02 | 49.65 | Confirmed | 105 | 25.86 | 31.09 | Confirmed
1 | TRUE | FALSE | FALSE | Styrene | 9.44 | 598 | 0.03 | 0.214 | Confirmed | 104 | 78 | 53.26 | 58.54 | Confirmed | 103 | 54.16 | 45.85 | Confirmed
1 | TRUE | FALSE | FALSE | Bromoform | 9.57 | 19 | 0 | 0.005 | Confirmed | 173 | 171 | 51.73 | 47.82 | Confirmed | 175 | 48.95 | 47.33 | Confirmed
1 | TRUE | FALSE | FALSE | Isopropylbenzene (cumene) | 9.71 | 604 | 0.03 | 0.02 | Confirmed | 105 | 120 | 29.67 | 24.93 | Confirmed | 79 | 15.83 | 14.37 | Confirmed
20 | TRUE | FALSE | FALSE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 179129 | 8.48 | 20.238 | Confirmed | 95 | 174 | 77.35 | 77.55 | Confirmed | 176 | 75.23 | 75.25 | Confirmed
1 | TRUE | FALSE | FALSE | Bromobenzene | 9.94 | 395 | 0.02 | 0.029 | Confirmed | 77 | 156 | 64.2 | 50.34 | Confirmed | 158 | 60.88 | 65.08 | Confirmed
1 | TRUE | TRUE | FALSE | 1,1,2,2-Tetrachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 83 | 85 | 64.65 | n.a. | Not confirmed | 95 | 12.92 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2,3-Trichloropropane (TCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 110 | 77.35 | n.a. | Not confirmed | 61 | 59.19 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | trans-1,4-Dichloro-2-butene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 75 | 53 | 19.71 | n.a. | Not confirmed | 89 | 9.73 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | n-Propylbenzene | 10.02 | 1346 | 0.06 | 0.038 | Confirmed | 91 | 120 | 26.51 | 24.21 | Confirmed | 65 | 9.84 | 11.69 | Confirmed
1 | TRUE | FALSE | FALSE | 2-Chlorotoluene | 10.08 | 622 | 0.03 | 0.028 | Confirmed | 91 | 126 | 35.91 | 35.32 | Confirmed | 89 | 17.14 | 15.91 | Confirmed
1 | TRUE | TRUE | FALSE | 1,3,5-Trimethylbenzene | 10.15 | 574 | 0.03 | 0.02 | Not confirmed | 105 | 120 | 49.52 | 54.49 | Confirmed | 119 | 11.62 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 4-Chlorotoluene | 10.17 | 981 | 0.05 | 0.039 | Confirmed | 91 | 126 | 32.83 | 36.93 | Confirmed | 89 | 11.29 | 11.44 | Confirmed
1 | TRUE | FALSE | FALSE | tert-Butylbenzene | 10.38 | 370 | 0.02 | 0.358 | Confirmed | 119 | 91 | 66.88 | 76.87 | Confirmed | 134 | 24.28 | 26.96 | Confirmed
1 | TRUE | TRUE | FALSE | Pentachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 167 | 130 | 65.82 | n.a. | Not confirmed | 132 | 73.17 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trimethylbenzene | 10.42 | 702 | 0.03 | 0.025 | Confirmed | 105 | 120 | 47 | 49.4 | Confirmed | 77 | 11.06 | 14.09 | Confirmed
1 | TRUE | TRUE | FALSE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 1097 | 0.05 | 0.032 | Not confirmed | 105 | 134 | 21.16 | 20.77 | Confirmed | 91 | 14.82 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,3-Dichlorobenzene | 10.61 | 860 | 0.04 | 0.055 | Confirmed | 146 | 148 | 63.78 | 58.7 | Confirmed | 111 | 42.28 | 35.58 | Confirmed
1 | TRUE | FALSE | FALSE | 4-Isopropyltoluene (p-cymene) | 10.64 | 849 | 0.04 | 0.031 | Confirmed | 119 | 91 | 27.27 | 32.21 | Confirmed | 134 | 29.5 | 32.34 | Confirmed
20 | TRUE | FALSE | FALSE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 191020 | 9.04 | 20 | Confirmed | 152 | 115 | 52.82 | 62.19 | Confirmed | n.a. | n.a. | n.a. | n.a.
1 | TRUE | TRUE | FALSE | 1,4-Dichlorobenzene | 10.68 | 966 | 0.05 | 0.062 | Not confirmed | 146 | 148 | 62.86 | 99.76 | Not confirmed | 111 | 42.75 | 228.62 | Not confirmed
1 | TRUE | FALSE | FALSE | n-Butylbenzene | 10.92 | 1376 | 0.07 | 0.05 | Confirmed | 91 | 92 | 52.73 | 57.82 | Confirmed | 134 | 28.18 | 29.23 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2-Dichlorobenzene | 10.93 | 743 | 0.04 | 0.046 | Confirmed | 146 | 148 | 64.22 | 61.75 | Confirmed | 111 | 43.87 | 44.13 | Confirmed
1 | TRUE | TRUE | FALSE | Hexachloroethane | n.a. | n.a. | n.a. | n.a. | Not confirmed | 117 | 119 | 97.1 | n.a. | Not confirmed | 201 | 88.55 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | n.a. | n.a. | n.a. | n.a. | Not confirmed | 157 | 155 | 76.87 | n.a. | Not confirmed | 75 | 85.26 | n.a. | Not confirmed
1 | TRUE | TRUE | FALSE | Nitrobenzene | n.a. | n.a. | n.a. | n.a. | Not confirmed | 77 | 51 | 45.42 | n.a. | Not confirmed | 123 | 45.76 | n.a. | Not confirmed
1 | TRUE | FALSE | FALSE | 1,2,4-Trichlorobenzene | 11.98 | 845 | 0.04 | 0.08 | Confirmed | 180 | 182 | 96.02 | 93.33 | Confirmed | 145 | 32.33 | 28.97 | Confirmed
1 | TRUE | FALSE | FALSE | Hexachloro-1,3-butadiene | 12.07 | 248 | 0.01 | 0.05 | Confirmed | 225 | 227 | 63.52 | 64.16 | Confirmed | 223 | 64.42 | 48.95 | Confirmed
1 | TRUE | FALSE | FALSE | Naphthalene | 12.15 | 1226 | 0.06 | 0.219 | Confirmed | 128 | 127 | 12.82 | 13.12 | Confirmed | 129 | 10.56 | 10.03 | Confirmed
1 | TRUE | FALSE | FALSE | 1,2,3-Trichlorobenzene | 12.29 | 662 | 0.03 | 0.055 | Confirmed | 180 | 182 | 95.7 | 93.14 | Confirmed | 145 | 34.15 | 31.14 | Confirmed

## CCV
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? |  |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 207938 | 5.43 | 229548 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.17 | 333742 | 6.18 | 368257 | TRUE | TRUE | 112.79 | 10 | TRUE | Chloromethane (methyl chloride) | 1.47 | 22302 | 0.1 | 11.279 | Confirmed | 50 | 52 | 32.87 | 32.63 | Confirmed | 49 | 9.87 | 9.16 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 320785 | 8.92 | 355665 | TRUE | TRUE | 116.67 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.56 | 46142 | 0.22 | 11.667 | Confirmed | 62 | 64 | 31.53 | 31.67 | Confirmed | 61 | 7.81 | 8.01 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 177359 | 10.67 | 204454 | TRUE | TRUE | 149.51 | 10 | FALSE | Bromomethane (methyl bromide) | 1.84 | 75100 | 0.35 | 14.951 | Confirmed | 94 | 96 | 97.36 | 93.08 | Confirmed | 93 | 20.34 | 19.56 | Confirmed
[7 empty] | 118.13 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.95 | 41662 | 0.2 | 11.813 | Confirmed | 64 | 66 | 32.12 | 32.45 | Confirmed | 49 | 23.08 | 23.44 | Confirmed
[7 empty] | 109.2 | 10 | TRUE | Trichlorofluoromethane | 2.19 | 64688 | 0.3 | 10.92 | Confirmed | 101 | 103 | 64.84 | 65.02 | Confirmed | 105 | 10.5 | 10.7 | Confirmed
Total Analytes | 81 | [5 empty] | 99.37 | 10 | TRUE | Diethyl ether | 2.5 | 54690 | 0.26 | 9.937 | Confirmed | 59 | 74 | 75.46 | 77.39 | Confirmed | 45 | 72.36 | 75.26 | Confirmed
Failed | 4 | [5 empty] | 106.7 | 10 | TRUE | 1,1-Dichloroethene | 2.74 | 73260 | 0.34 | 10.67 | Confirmed | 61 | 96 | 72.45 | 75.53 | Confirmed | 98 | 46.37 | 47.45 | Confirmed
[7 empty] | 87.3833 | 18 | TRUE | Acetone | 2.83 | 29033 | 0.14 | 15.729 | Confirmed | 43 | 58 | 40 | 37.88 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 76.48 | 10 | FALSE | Iodomethane | 2.89 | 47547 | 0.22 | 7.648 | Confirmed | 142 | 127 | 31 | 30.03 | Confirmed | 141 | 13.37 | 12.71 | Confirmed
[7 empty] | 111.79 | 10 | TRUE | Carbon disulfide | 2.95 | 182556 | 0.86 | 11.179 | Confirmed | 76 | 78 | 8.96 | 8.74 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 104.05 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.19 | 85693 | 0.4 | 10.405 | Confirmed | 41 | 39 | 52.2 | 49.94 | Confirmed | 76 | 36.97 | 35.53 | Confirmed
[7 empty] | 114.91 | 10 | TRUE | Methylene chloride (DCM) | 3.36 | 89424 | 0.42 | 11.491 | Confirmed | 49 | 84 | 87.31 | 88.83 | Confirmed | 86 | 55.97 | 57.31 | Confirmed
[7 empty] | 104.04 | 10 | TRUE | trans-1,2-Dichloroethene | 3.68 | 83201 | 0.39 | 10.404 | Confirmed | 61 | 96 | 75.67 | 76.42 | Confirmed | 98 | 47.1 | 48.1 | Confirmed
[7 empty] | 81.59 | 10 | TRUE | Methyl tert-butyl ether (MTBE) | 3.69 | 128788 | 0.61 | 8.159 | Confirmed | 73 | 41 | 33.25 | 33.44 | Confirmed | 57 | 23.18 | 24.06 | Confirmed
[7 empty] | 107.09 | 10 | TRUE | 1,1-Dichloroethane | 4.19 | 106441 | 0.5 | 10.709 | Confirmed | 63 | 65 | 31.98 | 31.78 | Confirmed | 83 | 12.43 | 12.4 | Confirmed
[7 empty] | 91.92 | 10 | TRUE | 2,2-Dichloropropane | 4.81 | 43605 | 0.21 | 9.192 | Confirmed | 77 | 41 | 101.06 | 100.31 | Confirmed | 79 | 32.22 | 31.76 | Confirmed
[7 empty] | 106.05 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 93584 | 0.44 | 10.605 | Confirmed | 61 | 96 | 80.14 | 81.55 | Confirmed | 98 | 51.39 | 51.6 | Confirmed
[7 empty] | 83.6444 | 18 | TRUE | 2-Butanone (MEK) | 4.84 | 43783 | 0.21 | 15.056 | Confirmed | 43 | 72 | 28.44 | 27.79 | Confirmed | 57 | 7.2 | 7.82 | Confirmed
[7 empty] | 91.39 | 10 | TRUE | Methyl acrylate | 4.94 | 36335 | 0.17 | 9.139 | Confirmed | 55 | 85 | 16.75 | 16.65 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 90.51 | 10 | TRUE | Methacrylonitrile | 5.06 | 28077 | 0.13 | 9.051 | Confirmed | 67 | 52 | 30.53 | 31.2 | Confirmed | 40 | 35.3 | 33.34 | Confirmed
[7 empty] | 107.86 | 10 | TRUE | Bromochloromethane | 5.07 | 59866 | 0.28 | 10.786 | Confirmed | 49 | 130 | 82.37 | 80.44 | Confirmed | 128 | 64.27 | 61.87 | Confirmed
[7 empty] | 83.8 | 10 | TRUE | Tetrahydrofuran | 5.08 | 20010 | 0.09 | 8.38 | Confirmed | 42 | 72 | 42.66 | 40.59 | Confirmed | 71 | 40.38 | 42.23 | Confirmed
[7 empty] | 114.53 | 10 | TRUE | Trichloromethane (chloroform) | 5.2 | 105002 | 0.49 | 11.453 | Confirmed | 83 | 85 | 63.95 | 65.72 | Confirmed | 47 | 17.3 | 17.92 | Confirmed
[7 empty] | 100.63 | 10 | TRUE | 1,1,1-Trichloroethane | 5.33 | 61550 | 0.29 | 10.063 | Confirmed | 97 | 99 | 63.95 | 63.94 | Confirmed | 61 | 48.19 | 47.48 | Confirmed
[7 empty] | 112.46 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.36 | 109373 | 0.51 | 22.492 | Confirmed | 113 | 111 | 104.17 | 103.51 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 207938 | 0.98 | 20 | Confirmed | 168 | 99 | 48.98 | 49.36 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 97.99 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 95412 | 0.45 | 9.799 | Confirmed | 56 | 41 | 59.56 | 59.08 | Confirmed | 43 | 26.98 | 27.41 | Confirmed
[7 empty] | 95.15 | 10 | TRUE | Carbon tetrachloride | 5.49 | 46460 | 0.22 | 9.515 | Confirmed | 119 | 121 | 31.25 | 32.47 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 101.66 | 10 | TRUE | 1,1-Dichloropropene | 5.51 | 68605 | 0.32 | 10.166 | Confirmed | 75 | 77 | 31.46 | 31.77 | Confirmed | 110 | 41.64 | 43 | Confirmed
[7 empty] | 101.99 | 10 | TRUE | Benzene | 5.7 | 267039 | 1.26 | 10.199 | Confirmed | 78 | 77 | 24.33 | 24.56 | Confirmed | 52 | 14.88 | 14.77 | Confirmed
[7 empty] | 98.51 | 10 | TRUE | 1,2-Dichloroethane | 5.77 | 66914 | 0.32 | 9.851 | Confirmed | 62 | 64 | 31.81 | 32.05 | Confirmed | 49 | 32.62 | 33.21 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.17 | 333742 | 1.57 | 20 | Confirmed | 114 | 88 | 18.42 | 18.01 | Confirmed | 63 | 18.27 | 17.87 | Confirmed
[7 empty] | 103.56 | 10 | TRUE | Trichloroethene | 6.39 | 67377 | 0.32 | 10.356 | Confirmed | 130 | 132 | 96.07 | 100.14 | Confirmed | 95 | 97.35 | 95.03 | Confirmed
[7 empty] | 101.12 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 64821 | 0.31 | 10.112 | Confirmed | 63 | 62 | 69.44 | 68.96 | Confirmed | 41 | 41.97 | 41.63 | Confirmed
[7 empty] | 111.6 | 10 | TRUE | Dibromomethane | 6.72 | 46382 | 0.22 | 11.16 | Confirmed | 174 | 93 | 91.77 | 91.89 | Confirmed | 95 | 78.67 | 78.96 | Confirmed
[7 empty] | 87.66 | 10 | TRUE | Methyl methacrylate | 6.74 | 34166 | 0.16 | 8.766 | Confirmed | 41 | 69 | 88.3 | 89.65 | Confirmed | 39 | 41.9 | 41.43 | Confirmed
[7 empty] | 103.42 | 10 | TRUE | Bromodichloromethane | 6.92 | 69094 | 0.33 | 10.342 | Confirmed | 83 | 85 | 62.95 | 62.82 | Confirmed | 47 | 15.49 | 15.36 | Confirmed
[7 empty] | 79.36 | 10 | FALSE | 2-Nitropropane | 7.15 | 8917 | 0.04 | 7.936 | Confirmed | 43 | 41 | 87.28 | 91.02 | Confirmed | 39 | 28.76 | 27.81 | Confirmed
[7 empty] | 99.68 | 10 | TRUE | cis-1,3-Dichloropropene | 7.36 | 68026 | 0.32 | 9.968 | Confirmed | 75 | 39 | 39.82 | 39.57 | Confirmed | 77 | 31.39 | 31.2 | Confirmed
[7 empty] | 81.2722 | 18 | TRUE | 4-Methyl-2-pentanone (MIBK) | 7.52 | 99130 | 0.47 | 14.629 | Confirmed | 43 | 58 | 39.86 | 39.65 | Confirmed | 41 | 22.34 | 22.15 | Confirmed
[7 empty] | 101.515 | 20 | TRUE | Toluene-d8 [SS2] | 7.61 | 456940 | 2.15 | 20.303 | Confirmed | 98 | 100 | 62 | 63.19 | Confirmed | 70 | 10.79 | 10.85 | Confirmed
[7 empty] | 104.42 | 10 | TRUE | Toluene | 7.67 | 295453 | 1.39 | 10.442 | Confirmed | 91 | 92 | 55.73 | 55.97 | Confirmed | 65 | 11.33 | 11.15 | Confirmed
[7 empty] | 97.11 | 10 | TRUE | trans-1,3-Dichloropropene | 7.92 | 49019 | 0.23 | 9.711 | Confirmed | 75 | 39 | 41.41 | 41.21 | Confirmed | 77 | 31.34 | 31.29 | Confirmed
[7 empty] | 96.79 | 10 | TRUE | Ethyl methacrylate | 7.99 | 57067 | 0.27 | 9.679 | Confirmed | 69 | 41 | 61.12 | 62.05 | Confirmed | 99 | 24.12 | 25.12 | Confirmed
[7 empty] | 97.3 | 10 | TRUE | 1,1,2-Trichloroethane | 8.1 | 58370 | 0.27 | 9.73 | Confirmed | 97 | 83 | 86.1 | 87.63 | Confirmed | 99 | 61.03 | 62.23 | Confirmed
[7 empty] | 101.56 | 10 | TRUE | Tetrachloroethene | 8.15 | 97980 | 0.46 | 10.156 | Confirmed | 166 | 164 | 78.43 | 77.61 | Confirmed | 129 | 72.49 | 73.56 | Confirmed
[7 empty] | 98.17 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 93212 | 0.44 | 9.817 | Confirmed | 76 | 41 | 62.02 | 62.26 | Confirmed | 78 | 32.22 | 32.17 | Confirmed
[7 empty] | 81.2833 | 18 | TRUE | 2-Hexanone | 8.31 | 65250 | 0.31 | 14.631 | Confirmed | 43 | 58 | 55.88 | 54.42 | Confirmed | 57 | 19.46 | 18.67 | Confirmed
[7 empty] | 90.28 | 10 | TRUE | Dibromochloromethane | 8.43 | 50778 | 0.24 | 9.028 | Confirmed | 129 | 127 | 77.94 | 79.06 | Confirmed | 131 | 24.25 | 24.76 | Confirmed
[7 empty] | 99.49 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.52 | 53574 | 0.25 | 9.949 | Confirmed | 107 | 109 | 94.8 | 95.23 | Confirmed | 93 | 4.53 | 4.43 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 320785 | 1.51 | 20 | Confirmed | 117 | 82 | 58.94 | 58.61 | Confirmed | 52 | 13.5 | 13.5 | Confirmed
[7 empty] | 105.77 | 10 | TRUE | Chlorobenzene | 8.93 | 191280 | 0.9 | 10.577 | Confirmed | 112 | 77 | 62.04 | 61.48 | Confirmed | 114 | 31.57 | 31.81 | Confirmed
[7 empty] | 101.43 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 48679 | 0.23 | 10.143 | Confirmed | 131 | 133 | 95.48 | 93.81 | Confirmed | 117 | 79 | 76.84 | Confirmed
[7 empty] | 112.23 | 10 | TRUE | Ethylbenzene | 9.02 | 304622 | 1.43 | 11.223 | Confirmed | 91 | 106 | 35.7 | 35.59 | Confirmed | 51 | 8.99 | 8.77 | Confirmed
[7 empty] | 102 | 10 | TRUE | m/p-Xylene | 9.13 | 531185 | 2.5 | 10.2 | Confirmed | 91 | 106 | 50.83 | 52.28 | Confirmed | 105 | 21.8 | 21.95 | Confirmed
[7 empty] | 111.81 | 10 | TRUE | o-Xylene | 9.43 | 275159 | 1.3 | 11.181 | Confirmed | 91 | 106 | 50.02 | 49.43 | Confirmed | 105 | 25.86 | 25.51 | Confirmed
[7 empty] | 109.53 | 10 | TRUE | Styrene | 9.44 | 228807 | 1.08 | 10.953 | Confirmed | 104 | 78 | 53.26 | 53.53 | Confirmed | 103 | 54.16 | 54.52 | Confirmed
[7 empty] | 99.69 | 10 | TRUE | Bromoform | 9.57 | 35411 | 0.17 | 9.969 | Confirmed | 173 | 171 | 51.73 | 51.35 | Confirmed | 175 | 48.95 | 48.16 | Confirmed
[7 empty] | 108.75 | 10 | TRUE | Isopropylbenzene (cumene) | 9.71 | 308245 | 1.45 | 10.875 | Confirmed | 105 | 120 | 29.67 | 29.57 | Confirmed | 79 | 15.83 | 15.53 | Confirmed
[7 empty] | 101.3 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 166502 | 0.78 | 20.26 | Confirmed | 95 | 174 | 77.35 | 78.81 | Confirmed | 176 | 75.23 | 75.22 | Confirmed
[7 empty] | 105.68 | 10 | TRUE | Bromobenzene | 9.93 | 132978 | 0.63 | 10.568 | Confirmed | 77 | 156 | 64.2 | 64.24 | Confirmed | 158 | 60.88 | 61.92 | Confirmed
[7 empty] | 100.44 | 10 | TRUE | 1,1,2,2-Tetrachloroethane | 9.95 | 73780 | 0.35 | 10.044 | Confirmed | 83 | 85 | 64.65 | 65.24 | Confirmed | 95 | 12.92 | 13.36 | Confirmed
[7 empty] | 95.78 | 10 | TRUE | 1,2,3-Trichloropropane (TCP) | 9.98 | 27089 | 0.13 | 9.578 | Confirmed | 77 | 110 | 77.35 | 78.41 | Confirmed | 61 | 59.19 | 59.49 | Confirmed
[7 empty] | 94.9 | 10 | TRUE | trans-1,4-Dichloro-2-butene | 9.98 | 81251 | 0.38 | 9.49 | Confirmed | 75 | 53 | 19.71 | 19.6 | Confirmed | 89 | 9.73 | 9.88 | Confirmed
[7 empty] | 114.05 | 10 | TRUE | n-Propylbenzene | 10.02 | 376139 | 1.77 | 11.405 | Confirmed | 91 | 120 | 26.51 | 25.58 | Confirmed | 65 | 9.84 | 9.78 | Confirmed
[7 empty] | 115.25 | 10 | TRUE | 2-Chlorotoluene | 10.08 | 233533 | 1.1 | 11.525 | Confirmed | 91 | 126 | 35.91 | 36.52 | Confirmed | 89 | 17.14 | 17.37 | Confirmed
[7 empty] | 107.34 | 10 | TRUE | 1,3,5-Trimethylbenzene | 10.15 | 290037 | 1.37 | 10.734 | Confirmed | 105 | 120 | 49.52 | 49.99 | Confirmed | 119 | 11.62 | 11.81 | Confirmed
[7 empty] | 117.34 | 10 | TRUE | 4-Chlorotoluene | 10.17 | 275738 | 1.3 | 11.734 | Confirmed | 91 | 126 | 32.83 | 33 | Confirmed | 89 | 11.29 | 11.29 | Confirmed
[7 empty] | 109.25 | 10 | TRUE | tert-Butylbenzene | 10.37 | 259758 | 1.22 | 10.925 | Confirmed | 119 | 91 | 66.88 | 63.66 | Confirmed | 134 | 24.28 | 24.06 | Confirmed
[7 empty] | 120.24 | 10 | FALSE | Pentachloroethane | 10.39 | 19872 | 0.09 | 12.024 | Confirmed | 167 | 130 | 65.82 | 64.48 | Confirmed | 132 | 73.17 | 67.98 | Confirmed
[7 empty] | 112.48 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.41 | 294874 | 1.39 | 11.248 | Confirmed | 105 | 120 | 47 | 47.87 | Confirmed | 77 | 11.06 | 10.88 | Confirmed
[7 empty] | 111.63 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 351685 | 1.66 | 11.163 | Confirmed | 105 | 134 | 21.16 | 21.61 | Confirmed | 91 | 14.82 | 14.99 | Confirmed
[7 empty] | 115.02 | 10 | TRUE | 1,3-Dichlorobenzene | 10.61 | 166072 | 0.78 | 11.502 | Confirmed | 146 | 148 | 63.78 | 63.64 | Confirmed | 111 | 42.28 | 42.94 | Confirmed
[7 empty] | 113.45 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.63 | 292463 | 1.38 | 11.345 | Confirmed | 119 | 91 | 27.27 | 26.99 | Confirmed | 134 | 29.5 | 30.05 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 177359 | 0.83 | 20 | Confirmed | 152 | 115 | 52.82 | 51.89 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 118.01 | 10 | TRUE | 1,4-Dichlorobenzene | 10.68 | 170688 | 0.8 | 11.801 | Confirmed | 146 | 148 | 62.86 | 64.14 | Confirmed | 111 | 42.75 | 42.86 | Confirmed
[7 empty] | 107.66 | 10 | TRUE | n-Butylbenzene | 10.91 | 276671 | 1.3 | 10.766 | Confirmed | 91 | 92 | 52.73 | 51.15 | Confirmed | 134 | 28.18 | 28.24 | Confirmed
[7 empty] | 117.36 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 174478 | 0.82 | 11.736 | Confirmed | 146 | 148 | 64.22 | 63.39 | Confirmed | 111 | 43.87 | 43.26 | Confirmed
[7 empty] | 107.31 | 10 | TRUE | Hexachloroethane | 11.1 | 31720 | 0.15 | 10.731 | Confirmed | 117 | 119 | 97.1 | 96.12 | Confirmed | 201 | 88.55 | 87.21 | Confirmed
[7 empty] | 95.24 | 10 | TRUE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 14889 | 0.07 | 9.524 | Confirmed | 157 | 155 | 76.87 | 76.72 | Confirmed | 75 | 85.26 | 86.58 | Confirmed
[7 empty] | 87.96 | 10 | TRUE | Nitrobenzene | 11.58 | 3670 | 0.02 | 8.796 | Confirmed | 77 | 51 | 45.42 | 45.89 | Confirmed | 123 | 45.76 | 48.3 | Confirmed
[7 empty] | 107.01 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.98 | 107193 | 0.5 | 10.701 | Confirmed | 180 | 182 | 96.02 | 95.08 | Confirmed | 145 | 32.33 | 32.24 | Confirmed
[7 empty] | 98.56 | 10 | TRUE | Hexachloro-1,3-butadiene | 12.07 | 45769 | 0.22 | 9.856 | Confirmed | 225 | 227 | 63.52 | 64.46 | Confirmed | 223 | 64.42 | 63.71 | Confirmed
[7 empty] | 101.76 | 10 | TRUE | Naphthalene | 12.15 | 270303 | 1.27 | 10.176 | Confirmed | 128 | 127 | 12.82 | 12.75 | Confirmed | 129 | 10.56 | 10.34 | Confirmed
[7 empty] | 96.18 | 10 | TRUE | 1,2,3-Trichlorobenzene | 12.29 | 106821 | 0.5 | 9.618 | Confirmed | 180 | 182 | 95.7 | 95.59 | Confirmed | 145 | 34.15 | 33.64 | Confirmed

## CCV2
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? |  |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 246250 | 5.43 | 229548 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.17 | 386577 | 6.18 | 368257 | TRUE | TRUE | 90.64 | 10 | TRUE | Chloromethane (methyl chloride) | 1.47 | 21225 | 0.1 | 9.064 | Confirmed | 50 | 52 | 32.87 | 32.74 | Confirmed | 49 | 9.87 | 10.13 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 390416 | 8.92 | 355665 | TRUE | TRUE | 87.16 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.56 | 40825 | 0.2 | 8.716 | Confirmed | 62 | 64 | 31.53 | 31.61 | Confirmed | 61 | 7.81 | 7.57 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 230139 | 10.67 | 204454 | TRUE | TRUE | 102.63 | 10 | TRUE | Bromomethane (methyl bromide) | 1.84 | 63986 | 0.31 | 10.263 | Confirmed | 94 | 96 | 97.36 | 93.57 | Confirmed | 93 | 20.34 | 20.42 | Confirmed
[7 empty] | 94.52 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.95 | 39479 | 0.19 | 9.452 | Confirmed | 64 | 66 | 32.12 | 32.24 | Confirmed | 49 | 23.08 | 23.23 | Confirmed
[7 empty] | 83.82 | 10 | TRUE | Trichlorofluoromethane | 2.19 | 58804 | 0.28 | 8.382 | Confirmed | 101 | 103 | 64.84 | 66.85 | Confirmed | 105 | 10.5 | 10.69 | Confirmed
Total Analytes | 81 | [5 empty] | 84.78 | 10 | TRUE | Diethyl ether | 2.5 | 55181 | 0.27 | 8.478 | Confirmed | 59 | 74 | 75.46 | 75.7 | Confirmed | 45 | 72.36 | 74.4 | Confirmed
Failed | 34 | [5 empty] | 85.13 | 10 | TRUE | 1,1-Dichloroethene | 2.74 | 69219 | 0.33 | 8.513 | Confirmed | 61 | 96 | 72.45 | 75.35 | Confirmed | 98 | 46.37 | 47.5 | Confirmed
[7 empty] | 79.6944 | 18 | FALSE | Acetone | 2.83 | 31117 | 0.15 | 14.345 | Confirmed | 43 | 58 | 40 | 35.5 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 66.88 | 10 | FALSE | Iodomethane | 2.89 | 46734 | 0.23 | 6.688 | Confirmed | 142 | 127 | 31 | 30.6 | Confirmed | 141 | 13.37 | 13.31 | Confirmed
[7 empty] | 88.51 | 10 | TRUE | Carbon disulfide | 2.96 | 171167 | 0.83 | 8.851 | Confirmed | 76 | 78 | 8.96 | 8.88 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 79 | 10 | FALSE | 3-Chloropropene (allyl chloride) | 3.2 | 75838 | 0.37 | 7.9 | Confirmed | 41 | 39 | 52.2 | 51.56 | Confirmed | 76 | 36.97 | 36.79 | Confirmed
[7 empty] | 93.64 | 10 | TRUE | Methylene chloride (DCM) | 3.36 | 86298 | 0.42 | 9.364 | Confirmed | 49 | 84 | 87.31 | 89.9 | Confirmed | 86 | 55.97 | 57.33 | Confirmed
[7 empty] | 86.15 | 10 | TRUE | trans-1,2-Dichloroethene | 3.68 | 81589 | 0.39 | 8.615 | Confirmed | 61 | 96 | 75.67 | 73.21 | Confirmed | 98 | 47.1 | 47.78 | Confirmed
[7 empty] | 70.05 | 10 | FALSE | Methyl tert-butyl ether (MTBE) | 3.69 | 130958 | 0.63 | 7.005 | Confirmed | 73 | 41 | 33.25 | 33.91 | Confirmed | 57 | 23.18 | 23.81 | Confirmed
[7 empty] | 87.17 | 10 | TRUE | 1,1-Dichloroethane | 4.19 | 102882 | 0.5 | 8.717 | Confirmed | 63 | 65 | 31.98 | 31.48 | Confirmed | 83 | 12.43 | 12.26 | Confirmed
[7 empty] | 51.15 | 10 | FALSE | 2,2-Dichloropropane | 4.81 | 28732 | 0.14 | 5.115 | Confirmed | 77 | 41 | 101.06 | 104.5 | Confirmed | 79 | 32.22 | 32.22 | Confirmed
[7 empty] | 83.49 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 86761 | 0.42 | 8.349 | Confirmed | 61 | 96 | 80.14 | 81.82 | Confirmed | 98 | 51.39 | 52.6 | Confirmed
[7 empty] | 73.3167 | 18 | FALSE | 2-Butanone (MEK) | 4.84 | 44754 | 0.22 | 13.197 | Confirmed | 43 | 72 | 28.44 | 24.59 | Confirmed | 57 | 7.2 | 6.83 | Confirmed
[7 empty] | 80.17 | 10 | TRUE | Methyl acrylate | 4.93 | 37393 | 0.18 | 8.017 | Confirmed | 55 | 85 | 16.75 | 16.36 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 77.28 | 10 | FALSE | Methacrylonitrile | 5.06 | 27992 | 0.14 | 7.728 | Confirmed | 67 | 52 | 30.53 | 32.11 | Confirmed | 40 | 35.3 | 39.6 | Confirmed
[7 empty] | 87.16 | 10 | TRUE | Bromochloromethane | 5.07 | 56959 | 0.28 | 8.716 | Confirmed | 49 | 130 | 82.37 | 84.54 | Confirmed | 128 | 64.27 | 64.88 | Confirmed
[7 empty] | 75.85 | 10 | FALSE | Tetrahydrofuran | 5.08 | 21282 | 0.1 | 7.585 | Confirmed | 42 | 72 | 42.66 | 36.85 | Confirmed | 71 | 40.38 | 38.46 | Confirmed
[7 empty] | 92.19 | 10 | TRUE | Trichloromethane (chloroform) | 5.2 | 100095 | 0.48 | 9.219 | Confirmed | 83 | 85 | 63.95 | 65.42 | Confirmed | 47 | 17.3 | 17.56 | Confirmed
[7 empty] | 77.15 | 10 | FALSE | 1,1,1-Trichloroethane | 5.34 | 57430 | 0.28 | 7.715 | Confirmed | 97 | 99 | 63.95 | 63.71 | Confirmed | 61 | 48.19 | 48.31 | Confirmed
[7 empty] | 102.325 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.36 | 117854 | 0.57 | 20.465 | Confirmed | 113 | 111 | 104.17 | 103.4 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 246250 | 1.19 | 20 | Confirmed | 168 | 99 | 48.98 | 47.82 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 82.33 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 92653 | 0.45 | 8.233 | Confirmed | 56 | 41 | 59.56 | 59.2 | Confirmed | 43 | 26.98 | 27.63 | Confirmed
[7 empty] | 78.83 | 10 | FALSE | Carbon tetrachloride | 5.49 | 45159 | 0.22 | 7.883 | Confirmed | 119 | 121 | 31.25 | 32.31 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 85.53 | 10 | TRUE | 1,1-Dichloropropene | 5.51 | 66854 | 0.32 | 8.553 | Confirmed | 75 | 77 | 31.46 | 31.35 | Confirmed | 110 | 41.64 | 42.46 | Confirmed
[7 empty] | 80.75 | 10 | TRUE | Benzene | 5.7 | 257328 | 1.24 | 8.075 | Confirmed | 78 | 77 | 24.33 | 24.63 | Confirmed | 52 | 14.88 | 15.28 | Confirmed
[7 empty] | 79.99 | 10 | FALSE | 1,2-Dichloroethane | 5.77 | 65861 | 0.32 | 7.999 | Confirmed | 62 | 64 | 31.81 | 32.13 | Confirmed | 49 | 32.62 | 33.69 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.17 | 386577 | 1.87 | 20 | Confirmed | 114 | 88 | 18.42 | 18.09 | Confirmed | 63 | 18.27 | 17.96 | Confirmed
[7 empty] | 85.39 | 10 | TRUE | Trichloroethene | 6.38 | 67619 | 0.33 | 8.539 | Confirmed | 130 | 132 | 96.07 | 94.94 | Confirmed | 95 | 97.35 | 93.11 | Confirmed
[7 empty] | 82.1 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 64057 | 0.31 | 8.21 | Confirmed | 63 | 62 | 69.44 | 69.97 | Confirmed | 41 | 41.97 | 41.98 | Confirmed
[7 empty] | 92.5 | 10 | TRUE | Dibromomethane | 6.72 | 46787 | 0.23 | 9.25 | Confirmed | 174 | 93 | 91.77 | 87.93 | Confirmed | 95 | 78.67 | 74.71 | Confirmed
[7 empty] | 71.2 | 10 | FALSE | Methyl methacrylate | 6.74 | 33008 | 0.16 | 7.12 | Confirmed | 41 | 69 | 88.3 | 89.39 | Confirmed | 39 | 41.9 | 43.1 | Confirmed
[7 empty] | 82.01 | 10 | TRUE | Bromodichloromethane | 6.92 | 66681 | 0.32 | 8.201 | Confirmed | 83 | 85 | 62.95 | 63 | Confirmed | 47 | 15.49 | 15.27 | Confirmed
[7 empty] | 67.77 | 10 | FALSE | 2-Nitropropane | 7.14 | 9267 | 0.04 | 6.777 | Confirmed | 43 | 41 | 87.28 | 80.86 | Confirmed | 39 | 28.76 | 25.81 | Confirmed
[7 empty] | 78.2 | 10 | FALSE | cis-1,3-Dichloropropene | 7.36 | 63846 | 0.31 | 7.82 | Confirmed | 75 | 39 | 39.82 | 40.54 | Confirmed | 77 | 31.39 | 31.09 | Confirmed
[7 empty] | 65.7556 | 18 | FALSE | 4-Methyl-2-pentanone (MIBK) | 7.52 | 97615 | 0.47 | 11.836 | Confirmed | 43 | 58 | 39.86 | 40.06 | Confirmed | 41 | 22.34 | 22.98 | Confirmed
[7 empty] | 95.975 | 20 | TRUE | Toluene-d8 [SS2] | 7.6 | 525790 | 2.54 | 19.195 | Confirmed | 98 | 100 | 62 | 64.06 | Confirmed | 70 | 10.79 | 10.61 | Confirmed
[7 empty] | 80.33 | 10 | TRUE | Toluene | 7.67 | 276612 | 1.34 | 8.033 | Confirmed | 91 | 92 | 55.73 | 55.6 | Confirmed | 65 | 11.33 | 11.42 | Confirmed
[7 empty] | 77.01 | 10 | FALSE | trans-1,3-Dichloropropene | 7.92 | 46446 | 0.22 | 7.701 | Confirmed | 75 | 39 | 41.41 | 39.77 | Confirmed | 77 | 31.34 | 31.83 | Confirmed
[7 empty] | 80.24 | 10 | TRUE | Ethyl methacrylate | 7.99 | 56970 | 0.28 | 8.024 | Confirmed | 69 | 41 | 61.12 | 61.49 | Confirmed | 99 | 24.12 | 24.4 | Confirmed
[7 empty] | 80.04 | 10 | TRUE | 1,1,2-Trichloroethane | 8.1 | 58253 | 0.28 | 8.004 | Confirmed | 97 | 83 | 86.1 | 86.41 | Confirmed | 99 | 61.03 | 62.54 | Confirmed
[7 empty] | 79.36 | 10 | FALSE | Tetrachloroethene | 8.15 | 94148 | 0.45 | 7.936 | Confirmed | 166 | 164 | 78.43 | 77.93 | Confirmed | 129 | 72.49 | 72.42 | Confirmed
[7 empty] | 81.84 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 93563 | 0.45 | 8.184 | Confirmed | 76 | 41 | 62.02 | 61.34 | Confirmed | 78 | 32.22 | 32.08 | Confirmed
[7 empty] | 65.6778 | 18 | FALSE | 2-Hexanone | 8.31 | 64170 | 0.31 | 11.822 | Confirmed | 43 | 58 | 55.88 | 55.49 | Confirmed | 57 | 19.46 | 18.76 | Confirmed
[7 empty] | 75.36 | 10 | FALSE | Dibromochloromethane | 8.42 | 51587 | 0.25 | 7.536 | Confirmed | 129 | 127 | 77.94 | 76.4 | Confirmed | 131 | 24.25 | 24.05 | Confirmed
[7 empty] | 82.19 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.51 | 53803 | 0.26 | 8.219 | Confirmed | 107 | 109 | 94.8 | 96.23 | Confirmed | 93 | 4.53 | 4.32 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 390416 | 1.89 | 20 | Confirmed | 117 | 82 | 58.94 | 55.51 | Confirmed | 52 | 13.5 | 13.47 | Confirmed
[7 empty] | 82.24 | 10 | TRUE | Chlorobenzene | 8.93 | 181656 | 0.88 | 8.224 | Confirmed | 112 | 77 | 62.04 | 61.38 | Confirmed | 114 | 31.57 | 31.97 | Confirmed
[7 empty] | 80.59 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 46316 | 0.22 | 8.059 | Confirmed | 131 | 133 | 95.48 | 95.99 | Confirmed | 117 | 79 | 78.51 | Confirmed
[7 empty] | 88.36 | 10 | TRUE | Ethylbenzene | 9.02 | 291883 | 1.41 | 8.836 | Confirmed | 91 | 106 | 35.7 | 34.98 | Confirmed | 51 | 8.99 | 8.64 | Confirmed
[7 empty] | 79.51 | 10 | FALSE | m/p-Xylene | 9.13 | 502273 | 2.43 | 7.951 | Confirmed | 91 | 106 | 50.83 | 51.29 | Confirmed | 105 | 21.8 | 21.78 | Confirmed
[7 empty] | 87.79 | 10 | TRUE | o-Xylene | 9.43 | 262964 | 1.27 | 8.779 | Confirmed | 91 | 106 | 50.02 | 49.71 | Confirmed | 105 | 25.86 | 25.24 | Confirmed
[7 empty] | 84.48 | 10 | TRUE | Styrene | 9.44 | 214173 | 1.03 | 8.448 | Confirmed | 104 | 78 | 53.26 | 52.74 | Confirmed | 103 | 54.16 | 54.72 | Confirmed
[7 empty] | 77.93 | 10 | FALSE | Bromoform | 9.57 | 35789 | 0.17 | 7.793 | Confirmed | 173 | 171 | 51.73 | 50.43 | Confirmed | 175 | 48.95 | 47.65 | Confirmed
[7 empty] | 79.14 | 10 | FALSE | Isopropylbenzene (cumene) | 9.71 | 291085 | 1.41 | 7.914 | Confirmed | 105 | 120 | 29.67 | 29.81 | Confirmed | 79 | 15.83 | 15.46 | Confirmed
[7 empty] | 97.16 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 207221 | 1 | 19.432 | Confirmed | 95 | 174 | 77.35 | 77.23 | Confirmed | 176 | 75.23 | 74.61 | Confirmed
[7 empty] | 78.77 | 10 | FALSE | Bromobenzene | 9.93 | 128616 | 0.62 | 7.877 | Confirmed | 77 | 156 | 64.2 | 65.6 | Confirmed | 158 | 60.88 | 63.7 | Confirmed
[7 empty] | 76.28 | 10 | FALSE | 1,1,2,2-Tetrachloroethane | 9.95 | 71128 | 0.34 | 7.628 | Confirmed | 83 | 85 | 64.65 | 63.89 | Confirmed | 95 | 12.92 | 13.47 | Confirmed
[7 empty] | 73.27 | 10 | FALSE | 1,2,3-Trichloropropane (TCP) | 9.98 | 26889 | 0.13 | 7.327 | Confirmed | 77 | 110 | 77.35 | 79.85 | Confirmed | 61 | 59.19 | 59.05 | Confirmed
[7 empty] | 69.95 | 10 | FALSE | trans-1,4-Dichloro-2-butene | 9.98 | 77715 | 0.38 | 6.995 | Confirmed | 75 | 53 | 19.71 | 18.15 | Confirmed | 89 | 9.73 | 9.08 | Confirmed
[7 empty] | 83.47 | 10 | TRUE | n-Propylbenzene | 10.02 | 357192 | 1.73 | 8.347 | Confirmed | 91 | 120 | 26.51 | 25.95 | Confirmed | 65 | 9.84 | 9.83 | Confirmed
[7 empty] | 82.72 | 10 | TRUE | 2-Chlorotoluene | 10.07 | 217516 | 1.05 | 8.272 | Confirmed | 91 | 126 | 35.91 | 37.42 | Confirmed | 89 | 17.14 | 17.56 | Confirmed
[7 empty] | 78.94 | 10 | FALSE | 1,3,5-Trimethylbenzene | 10.15 | 276768 | 1.34 | 7.894 | Confirmed | 105 | 120 | 49.52 | 49.62 | Confirmed | 119 | 11.62 | 11.57 | Confirmed
[7 empty] | 86.5 | 10 | TRUE | 4-Chlorotoluene | 10.16 | 263774 | 1.27 | 8.65 | Confirmed | 91 | 126 | 32.83 | 33.14 | Confirmed | 89 | 11.29 | 11.16 | Confirmed
[7 empty] | 80.54 | 10 | TRUE | tert-Butylbenzene | 10.37 | 245597 | 1.19 | 8.054 | Confirmed | 119 | 91 | 66.88 | 64.19 | Confirmed | 134 | 24.28 | 24.39 | Confirmed
[7 empty] | 93.21 | 10 | TRUE | Pentachloroethane | 10.39 | 17745 | 0.09 | 9.321 | Confirmed | 167 | 130 | 65.82 | 64.15 | Confirmed | 132 | 73.17 | 68.2 | Confirmed
[7 empty] | 82.47 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.41 | 280565 | 1.36 | 8.247 | Confirmed | 105 | 120 | 47 | 48.12 | Confirmed | 77 | 11.06 | 10.72 | Confirmed
[7 empty] | 81.04 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 331318 | 1.6 | 8.104 | Confirmed | 105 | 134 | 21.16 | 21.38 | Confirmed | 91 | 14.82 | 14.82 | Confirmed
[7 empty] | 84.54 | 10 | TRUE | 1,3-Dichlorobenzene | 10.61 | 158386 | 0.77 | 8.454 | Confirmed | 146 | 148 | 63.78 | 62.6 | Confirmed | 111 | 42.28 | 41.59 | Confirmed
[7 empty] | 80.87 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.63 | 270514 | 1.31 | 8.087 | Confirmed | 119 | 91 | 27.27 | 27.37 | Confirmed | 134 | 29.5 | 30.3 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 230139 | 1.11 | 20 | Confirmed | 152 | 115 | 52.82 | 53.35 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 88.02 | 10 | TRUE | 1,4-Dichlorobenzene | 10.68 | 165203 | 0.8 | 8.802 | Confirmed | 146 | 148 | 62.86 | 63.13 | Confirmed | 111 | 42.75 | 42.61 | Confirmed
[7 empty] | 73.53 | 10 | FALSE | n-Butylbenzene | 10.91 | 245206 | 1.18 | 7.353 | Confirmed | 91 | 92 | 52.73 | 51.82 | Confirmed | 134 | 28.18 | 28.95 | Confirmed
[7 empty] | 84.97 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 163911 | 0.79 | 8.497 | Confirmed | 146 | 148 | 64.22 | 62.8 | Confirmed | 111 | 43.87 | 42.01 | Confirmed
[7 empty] | 79.15 | 10 | FALSE | Hexachloroethane | 11.1 | 30209 | 0.15 | 7.915 | Confirmed | 117 | 119 | 97.1 | 96.41 | Confirmed | 201 | 88.55 | 86.64 | Confirmed
[7 empty] | 75.73 | 10 | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 15004 | 0.07 | 7.573 | Confirmed | 157 | 155 | 76.87 | 75.77 | Confirmed | 75 | 85.26 | 83.58 | Confirmed
[7 empty] | 64.7 | 10 | FALSE | Nitrobenzene | 11.57 | 3486 | 0.02 | 6.47 | Confirmed | 77 | 51 | 45.42 | 48.9 | Confirmed | 123 | 45.76 | 52.45 | Confirmed
[7 empty] | 79.61 | 10 | FALSE | 1,2,4-Trichlorobenzene | 11.97 | 103067 | 0.5 | 7.961 | Confirmed | 180 | 182 | 96.02 | 97.11 | Confirmed | 145 | 32.33 | 31.96 | Confirmed
[7 empty] | 70.09 | 10 | FALSE | Hexachloro-1,3-butadiene | 12.06 | 42234 | 0.2 | 7.009 | Confirmed | 225 | 227 | 63.52 | 65.92 | Confirmed | 223 | 64.42 | 63.04 | Confirmed
[7 empty] | 80.95 | 10 | TRUE | Naphthalene | 12.15 | 275927 | 1.33 | 8.095 | Confirmed | 128 | 127 | 12.82 | 12.59 | Confirmed | 129 | 10.56 | 10.15 | Confirmed
[7 empty] | 71.28 | 10 | FALSE | 1,2,3-Trichlorobenzene | 12.29 | 102724 | 0.5 | 7.128 | Confirmed | 180 | 182 | 95.7 | 96.55 | Confirmed | 145 | 34.15 | 33.87 | Confirmed

## CCV3
[10 empty] | Peak Name | Ret. Time | Area | Rel Area | Amount | Overall Ion Ratio | Quant. Ion | Conf. Ion #1 | Ion Ratio #1 | Ion Ratio #1 | Ion Ratio #1 | Conf.Ion #2 | Ion Ratio #2 | Ion Ratio #2 | Ion Ratio #2
| RT | Area | ICAL Rt | ICAL Area | Pass_RT? | Pass_Area? |  |  |  |  | min | counts*min | % | ppb | Confirmation | m/z | m/z | (Expected) | (Observed) | Within Window | m/z | (Expected) | (Observed) | Within Window
Pentafluorobenzene [IS1] | 5.42 | 245625 | 5.43 | 229548 | TRUE | TRUE | % of TV | True Value | Pass? | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
1,4-Difluorobenzene [IS2] | 6.17 | 384052 | 6.18 | 368257 | TRUE | TRUE | 96.25 | 10 | TRUE | Chloromethane (methyl chloride) | 1.47 | 22482 | 0.11 | 9.625 | Confirmed | 50 | 52 | 32.87 | 33.01 | Confirmed | 49 | 9.87 | 8.42 | Confirmed
Chlorobenzene-d5 [IS3] | 8.91 | 367848 | 8.92 | 355665 | TRUE | TRUE | 95.69 | 10 | TRUE | Chloroethene (vinyl chloride) | 1.57 | 44703 | 0.21 | 9.569 | Confirmed | 62 | 64 | 31.53 | 31.69 | Confirmed | 61 | 7.81 | 7.66 | Confirmed
1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 220342 | 10.67 | 204454 | TRUE | TRUE | 99.47 | 10 | TRUE | Bromomethane (methyl bromide) | 1.84 | 62097 | 0.29 | 9.947 | Confirmed | 94 | 96 | 97.36 | 92.27 | Confirmed | 93 | 20.34 | 19.68 | Confirmed
[7 empty] | 99.18 | 10 | TRUE | Chloroethane (ethyl chloride) | 1.96 | 41317 | 0.2 | 9.918 | Confirmed | 64 | 66 | 32.12 | 32.5 | Confirmed | 49 | 23.08 | 22.98 | Confirmed
[7 empty] | 88.73 | 10 | TRUE | Trichlorofluoromethane | 2.2 | 62087 | 0.29 | 8.873 | Confirmed | 101 | 103 | 64.84 | 65.44 | Confirmed | 105 | 10.5 | 10.58 | Confirmed
Total Analytes | 81 | [5 empty] | 87.45 | 10 | TRUE | Diethyl ether | 2.5 | 56789 | 0.27 | 8.745 | Confirmed | 59 | 74 | 75.46 | 76.15 | Confirmed | 45 | 72.36 | 73.9 | Confirmed
Failed | 16 | [5 empty] | 88.36 | 10 | TRUE | 1,1-Dichloroethene | 2.74 | 71657 | 0.34 | 8.836 | Confirmed | 61 | 96 | 72.45 | 74.65 | Confirmed | 98 | 46.37 | 48.33 | Confirmed
[7 empty] | 76.6556 | 18 | FALSE | Acetone | 2.83 | 29751 | 0.14 | 13.798 | Confirmed | 43 | 58 | 40 | 33.79 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 92.83 | 10 | TRUE | Iodomethane | 2.89 | 72415 | 0.34 | 9.283 | Confirmed | 142 | 127 | 31 | 30.39 | Confirmed | 141 | 13.37 | 13.26 | Confirmed
[7 empty] | 92.29 | 10 | TRUE | Carbon disulfide | 2.96 | 178029 | 0.84 | 9.229 | Confirmed | 76 | 78 | 8.96 | 8.94 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 89.08 | 10 | TRUE | 3-Chloropropene (allyl chloride) | 3.2 | 85939 | 0.41 | 8.908 | Confirmed | 41 | 39 | 52.2 | 50.88 | Confirmed | 76 | 36.97 | 37.19 | Confirmed
[7 empty] | 98.75 | 10 | TRUE | Methylene chloride (DCM) | 3.36 | 90776 | 0.43 | 9.875 | Confirmed | 49 | 84 | 87.31 | 88.36 | Confirmed | 86 | 55.97 | 56.85 | Confirmed
[7 empty] | 86.95 | 10 | TRUE | trans-1,2-Dichloroethene | 3.68 | 82142 | 0.39 | 8.695 | Confirmed | 61 | 96 | 75.67 | 76.08 | Confirmed | 98 | 47.1 | 48.95 | Confirmed
[7 empty] | 72.66 | 10 | FALSE | Methyl tert-butyl ether (MTBE) | 3.69 | 135490 | 0.64 | 7.266 | Confirmed | 73 | 41 | 33.25 | 32.26 | Confirmed | 57 | 23.18 | 23.35 | Confirmed
[7 empty] | 92.03 | 10 | TRUE | 1,1-Dichloroethane | 4.19 | 108264 | 0.51 | 9.203 | Confirmed | 63 | 65 | 31.98 | 32.68 | Confirmed | 83 | 12.43 | 12.52 | Confirmed
[7 empty] | 78.98 | 10 | FALSE | 2,2-Dichloropropane | 4.82 | 44254 | 0.21 | 7.898 | Confirmed | 77 | 41 | 101.06 | 98.7 | Confirmed | 79 | 32.22 | 31.74 | Confirmed
[7 empty] | 91.4 | 10 | TRUE | cis-1,2-Dichloroethene | 4.82 | 94953 | 0.45 | 9.14 | Confirmed | 61 | 96 | 80.14 | 79.35 | Confirmed | 98 | 51.39 | 50.92 | Confirmed
[7 empty] | 75.0056 | 18 | FALSE | 2-Butanone (MEK) | 4.84 | 45797 | 0.22 | 13.501 | Confirmed | 43 | 72 | 28.44 | 25.33 | Confirmed | 57 | 7.2 | 7.67 | Confirmed
[7 empty] | 84.59 | 10 | TRUE | Methyl acrylate | 4.94 | 39508 | 0.19 | 8.459 | Confirmed | 55 | 85 | 16.75 | 16.8 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 78.17 | 10 | FALSE | Methacrylonitrile | 5.06 | 28271 | 0.13 | 7.817 | Confirmed | 67 | 52 | 30.53 | 32.95 | Confirmed | 40 | 35.3 | 39.46 | Confirmed
[7 empty] | 89.55 | 10 | TRUE | Bromochloromethane | 5.07 | 58420 | 0.28 | 8.955 | Confirmed | 49 | 130 | 82.37 | 84.66 | Confirmed | 128 | 64.27 | 66.56 | Confirmed
[7 empty] | 73.33 | 10 | FALSE | Tetrahydrofuran | 5.08 | 20464 | 0.1 | 7.333 | Confirmed | 42 | 72 | 42.66 | 41.12 | Confirmed | 71 | 40.38 | 41.47 | Confirmed
[7 empty] | 97.2 | 10 | TRUE | Trichloromethane (chloroform) | 5.2 | 105272 | 0.5 | 9.72 | Confirmed | 83 | 85 | 63.95 | 64.73 | Confirmed | 47 | 17.3 | 17.94 | Confirmed
[7 empty] | 87.16 | 10 | TRUE | 1,1,1-Trichloroethane | 5.34 | 61133 | 0.29 | 8.716 | Confirmed | 97 | 99 | 63.95 | 64.14 | Confirmed | 61 | 48.19 | 47 | Confirmed
[7 empty] | 103.415 | 20 | TRUE | Dibromofluoromethane [SS1] | 5.36 | 118803 | 0.56 | 20.683 | Confirmed | 113 | 111 | 104.17 | 103.45 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 100 | 20 | TRUE | Pentafluorobenzene [IS1] | 5.42 | 245625 | 1.17 | 20 | Confirmed | 168 | 99 | 48.98 | 47.75 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 85.98 | 10 | TRUE | 1-Chlorobutane (butyl chloride) | 5.48 | 96189 | 0.46 | 8.598 | Confirmed | 56 | 41 | 59.56 | 58.98 | Confirmed | 43 | 26.98 | 26.64 | Confirmed
[7 empty] | 82.66 | 10 | TRUE | Carbon tetrachloride | 5.49 | 47356 | 0.22 | 8.266 | Confirmed | 119 | 121 | 31.25 | 32.17 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 89.14 | 10 | TRUE | 1,1-Dichloropropene | 5.51 | 69224 | 0.33 | 8.914 | Confirmed | 75 | 77 | 31.46 | 31.52 | Confirmed | 110 | 41.64 | 41.34 | Confirmed
[7 empty] | 89.56 | 10 | TRUE | Benzene | 5.7 | 268899 | 1.28 | 8.956 | Confirmed | 78 | 77 | 24.33 | 23.94 | Confirmed | 52 | 14.88 | 14.56 | Confirmed
[7 empty] | 88.61 | 10 | TRUE | 1,2-Dichloroethane | 5.77 | 68899 | 0.33 | 8.861 | Confirmed | 62 | 64 | 31.81 | 31.6 | Confirmed | 49 | 32.62 | 31.78 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Difluorobenzene [IS2] | 6.17 | 384052 | 1.82 | 20 | Confirmed | 114 | 88 | 18.42 | 18.19 | Confirmed | 63 | 18.27 | 17.68 | Confirmed
[7 empty] | 92.83 | 10 | TRUE | Trichloroethene | 6.38 | 69258 | 0.33 | 9.283 | Confirmed | 130 | 132 | 96.07 | 97.26 | Confirmed | 95 | 97.35 | 94.2 | Confirmed
[7 empty] | 85.16 | 10 | TRUE | 1,2-Dichloropropane | 6.64 | 62602 | 0.3 | 8.516 | Confirmed | 63 | 62 | 69.44 | 70.1 | Confirmed | 41 | 41.97 | 41.45 | Confirmed
[7 empty] | 98.47 | 10 | TRUE | Dibromomethane | 6.72 | 46928 | 0.22 | 9.847 | Confirmed | 174 | 93 | 91.77 | 85.76 | Confirmed | 95 | 78.67 | 73.37 | Confirmed
[7 empty] | 75.47 | 10 | FALSE | Methyl methacrylate | 6.74 | 33192 | 0.16 | 7.547 | Confirmed | 41 | 69 | 88.3 | 89.59 | Confirmed | 39 | 41.9 | 41.53 | Confirmed
[7 empty] | 85.77 | 10 | TRUE | Bromodichloromethane | 6.92 | 65707 | 0.31 | 8.577 | Confirmed | 83 | 85 | 62.95 | 64.97 | Confirmed | 47 | 15.49 | 15.54 | Confirmed
[7 empty] | 67.55 | 10 | FALSE | 2-Nitropropane | 7.15 | 8704 | 0.04 | 6.755 | Confirmed | 43 | 41 | 87.28 | 87.72 | Confirmed | 39 | 28.76 | 26.37 | Confirmed
[7 empty] | 87.63 | 10 | TRUE | cis-1,3-Dichloropropene | 7.36 | 67972 | 0.32 | 8.763 | Confirmed | 75 | 39 | 39.82 | 39.29 | Confirmed | 77 | 31.39 | 31.52 | Confirmed
[7 empty] | 67.0667 | 18 | FALSE | 4-Methyl-2-pentanone (MIBK) | 7.52 | 93801 | 0.44 | 12.072 | Confirmed | 43 | 58 | 39.86 | 40.22 | Confirmed | 41 | 22.34 | 21.63 | Confirmed
[7 empty] | 98.105 | 20 | TRUE | Toluene-d8 [SS2] | 7.61 | 506391 | 2.4 | 19.621 | Confirmed | 98 | 100 | 62 | 64.14 | Confirmed | 70 | 10.79 | 10.62 | Confirmed
[7 empty] | 85.61 | 10 | TRUE | Toluene | 7.67 | 277753 | 1.32 | 8.561 | Confirmed | 91 | 92 | 55.73 | 56.84 | Confirmed | 65 | 11.33 | 11.11 | Confirmed
[7 empty] | 84.64 | 10 | TRUE | trans-1,3-Dichloropropene | 7.93 | 48469 | 0.23 | 8.464 | Confirmed | 75 | 39 | 41.41 | 40.66 | Confirmed | 77 | 31.34 | 32.09 | Confirmed
[7 empty] | 82.27 | 10 | TRUE | Ethyl methacrylate | 7.99 | 55117 | 0.26 | 8.227 | Confirmed | 69 | 41 | 61.12 | 62.15 | Confirmed | 99 | 24.12 | 25.71 | Confirmed
[7 empty] | 82.8 | 10 | TRUE | 1,1,2-Trichloroethane | 8.1 | 56815 | 0.27 | 8.28 | Confirmed | 97 | 83 | 86.1 | 87.98 | Confirmed | 99 | 61.03 | 62.78 | Confirmed
[7 empty] | 85.52 | 10 | TRUE | Tetrachloroethene | 8.16 | 95293 | 0.45 | 8.552 | Confirmed | 166 | 164 | 78.43 | 78.29 | Confirmed | 129 | 72.49 | 71.98 | Confirmed
[7 empty] | 82.3 | 10 | TRUE | 1,3-Dichloropropane | 8.24 | 88682 | 0.42 | 8.23 | Confirmed | 76 | 41 | 62.02 | 62.53 | Confirmed | 78 | 32.22 | 32.83 | Confirmed
[7 empty] | 65.2389 | 18 | FALSE | 2-Hexanone | 8.31 | 60055 | 0.28 | 11.743 | Confirmed | 43 | 58 | 55.88 | 56.19 | Confirmed | 57 | 19.46 | 19.34 | Confirmed
[7 empty] | 81.45 | 10 | TRUE | Dibromochloromethane | 8.43 | 52533 | 0.25 | 8.145 | Confirmed | 129 | 127 | 77.94 | 76.54 | Confirmed | 131 | 24.25 | 23.85 | Confirmed
[7 empty] | 88.91 | 10 | TRUE | 1,2-Dibromoethane (EDB) | 8.52 | 54866 | 0.26 | 8.891 | Confirmed | 107 | 109 | 94.8 | 93.37 | Confirmed | 93 | 4.53 | 4.24 | Confirmed
[7 empty] | 100 | 20 | TRUE | Chlorobenzene-d5 [IS3] | 8.91 | 367848 | 1.75 | 20 | Confirmed | 117 | 82 | 58.94 | 56.6 | Confirmed | 52 | 13.5 | 12.86 | Confirmed
[7 empty] | 88.43 | 10 | TRUE | Chlorobenzene | 8.93 | 183847 | 0.87 | 8.843 | Confirmed | 112 | 77 | 62.04 | 60.17 | Confirmed | 114 | 31.57 | 32 | Confirmed
[7 empty] | 89.84 | 10 | TRUE | 1,1,1,2-Tetrachloroethane | 9.01 | 49048 | 0.23 | 8.984 | Confirmed | 131 | 133 | 95.48 | 93.55 | Confirmed | 117 | 79 | 77.46 | Confirmed
[7 empty] | 94.47 | 10 | TRUE | Ethylbenzene | 9.02 | 294025 | 1.39 | 9.447 | Confirmed | 91 | 106 | 35.7 | 35.85 | Confirmed | 51 | 8.99 | 8.63 | Confirmed
[7 empty] | 86.56 | 10 | TRUE | m/p-Xylene | 9.13 | 515940 | 2.45 | 8.656 | Confirmed | 91 | 106 | 50.83 | 52.74 | Confirmed | 105 | 21.8 | 22.39 | Confirmed
[7 empty] | 92.76 | 10 | TRUE | o-Xylene | 9.43 | 261779 | 1.24 | 9.276 | Confirmed | 91 | 106 | 50.02 | 50.62 | Confirmed | 105 | 25.86 | 26.03 | Confirmed
[7 empty] | 92.2 | 10 | TRUE | Styrene | 9.45 | 220514 | 1.05 | 9.22 | Confirmed | 104 | 78 | 53.26 | 52.37 | Confirmed | 103 | 54.16 | 54.48 | Confirmed
[7 empty] | 83.9 | 10 | TRUE | Bromoform | 9.58 | 36924 | 0.18 | 8.39 | Confirmed | 173 | 171 | 51.73 | 50.55 | Confirmed | 175 | 48.95 | 48.57 | Confirmed
[7 empty] | 85.64 | 10 | TRUE | Isopropylbenzene (cumene) | 9.71 | 301566 | 1.43 | 8.564 | Confirmed | 105 | 120 | 29.67 | 29.02 | Confirmed | 79 | 15.83 | 14.89 | Confirmed
[7 empty] | 95.28 | 20 | TRUE | 1-Bromo-4-fluorobenzene (BFB) [SS3] | 9.84 | 194564 | 0.92 | 19.056 | Confirmed | 95 | 174 | 77.35 | 79.03 | Confirmed | 176 | 75.23 | 75.78 | Confirmed
[7 empty] | 80.65 | 10 | TRUE | Bromobenzene | 9.94 | 126070 | 0.6 | 8.065 | Confirmed | 77 | 156 | 64.2 | 67.37 | Confirmed | 158 | 60.88 | 63.92 | Confirmed
[7 empty] | 77.77 | 10 | FALSE | 1,1,2,2-Tetrachloroethane | 9.95 | 69551 | 0.33 | 7.777 | Confirmed | 83 | 85 | 64.65 | 66.11 | Confirmed | 95 | 12.92 | 13.42 | Confirmed
[7 empty] | 75.34 | 10 | FALSE | 1,2,3-Trichloropropane (TCP) | 9.98 | 26470 | 0.13 | 7.534 | Confirmed | 77 | 110 | 77.35 | 79.78 | Confirmed | 61 | 59.19 | 57.88 | Confirmed
[7 empty] | 73.06 | 10 | FALSE | trans-1,4-Dichloro-2-butene | 9.98 | 77712 | 0.37 | 7.306 | Confirmed | 75 | 53 | 19.71 | 19.87 | Confirmed | 89 | 9.73 | 10.72 | Confirmed
[7 empty] | 87.5 | 10 | TRUE | n-Propylbenzene | 10.02 | 358495 | 1.7 | 8.75 | Confirmed | 91 | 120 | 26.51 | 26.28 | Confirmed | 65 | 9.84 | 9.52 | Confirmed
[7 empty] | 88.06 | 10 | TRUE | 2-Chlorotoluene | 10.08 | 221691 | 1.05 | 8.806 | Confirmed | 91 | 126 | 35.91 | 37.51 | Confirmed | 89 | 17.14 | 17.03 | Confirmed
[7 empty] | 85.18 | 10 | TRUE | 1,3,5-Trimethylbenzene | 10.15 | 285946 | 1.36 | 8.518 | Confirmed | 105 | 120 | 49.52 | 48.73 | Confirmed | 119 | 11.62 | 11.73 | Confirmed
[7 empty] | 92.36 | 10 | TRUE | 4-Chlorotoluene | 10.17 | 269653 | 1.28 | 9.236 | Confirmed | 91 | 126 | 32.83 | 33 | Confirmed | 89 | 11.29 | 10.78 | Confirmed
[7 empty] | 83.87 | 10 | TRUE | tert-Butylbenzene | 10.37 | 245288 | 1.16 | 8.387 | Confirmed | 119 | 91 | 66.88 | 63.77 | Confirmed | 134 | 24.28 | 24.68 | Confirmed
[7 empty] | 101.66 | 10 | TRUE | Pentachloroethane | 10.39 | 19049 | 0.09 | 10.166 | Confirmed | 167 | 130 | 65.82 | 61.14 | Confirmed | 132 | 73.17 | 64.23 | Confirmed
[7 empty] | 88.94 | 10 | TRUE | 1,2,4-Trimethylbenzene | 10.42 | 289664 | 1.37 | 8.894 | Confirmed | 105 | 120 | 47 | 47.35 | Confirmed | 77 | 11.06 | 10.37 | Confirmed
[7 empty] | 86.91 | 10 | TRUE | 1-Methylpropylbenzene (sec-butylbenzene) | 10.53 | 340161 | 1.61 | 8.691 | Confirmed | 105 | 134 | 21.16 | 21.66 | Confirmed | 91 | 14.82 | 14.67 | Confirmed
[7 empty] | 91.29 | 10 | TRUE | 1,3-Dichlorobenzene | 10.61 | 163745 | 0.78 | 9.129 | Confirmed | 146 | 148 | 63.78 | 62.72 | Confirmed | 111 | 42.28 | 41.35 | Confirmed
[7 empty] | 92.45 | 10 | TRUE | 4-Isopropyltoluene (p-cymene) | 10.64 | 296068 | 1.4 | 9.245 | Confirmed | 119 | 91 | 27.27 | 25.58 | Confirmed | 134 | 29.5 | 29.12 | Confirmed
[7 empty] | 100 | 20 | TRUE | 1,4-Dichlorobenzene-d4 [IS4] | 10.66 | 220342 | 1.05 | 20 | Confirmed | 152 | 115 | 52.82 | 51.11 | Confirmed | n.a. | n.a. | n.a. | n.a.
[7 empty] | 93.43 | 10 | TRUE | 1,4-Dichlorobenzene | 10.68 | 167888 | 0.8 | 9.343 | Confirmed | 146 | 148 | 62.86 | 64.31 | Confirmed | 111 | 42.75 | 41.95 | Confirmed
[7 empty] | 83.23 | 10 | TRUE | n-Butylbenzene | 10.91 | 265709 | 1.26 | 8.323 | Confirmed | 91 | 92 | 52.73 | 52.03 | Confirmed | 134 | 28.18 | 29.44 | Confirmed
[7 empty] | 92.74 | 10 | TRUE | 1,2-Dichlorobenzene | 10.92 | 171293 | 0.81 | 9.274 | Confirmed | 146 | 148 | 64.22 | 62.61 | Confirmed | 111 | 43.87 | 42.8 | Confirmed
[7 empty] | 84.02 | 10 | TRUE | Hexachloroethane | 11.1 | 30730 | 0.15 | 8.402 | Confirmed | 117 | 119 | 97.1 | 96.91 | Confirmed | 201 | 88.55 | 88.32 | Confirmed
[7 empty] | 78.28 | 10 | FALSE | 1,2-Dibromo-3-chloropropane (DBCP) | 11.45 | 14905 | 0.07 | 7.828 | Confirmed | 157 | 155 | 76.87 | 79.17 | Confirmed | 75 | 85.26 | 81.05 | Confirmed
[7 empty] | 66.15 | 10 | FALSE | Nitrobenzene | 11.58 | 3414 | 0.02 | 6.615 | Confirmed | 77 | 51 | 45.42 | 44.56 | Confirmed | 123 | 45.76 | 48.68 | Confirmed
[7 empty] | 86.6 | 10 | TRUE | 1,2,4-Trichlorobenzene | 11.98 | 107448 | 0.51 | 8.66 | Confirmed | 180 | 182 | 96.02 | 94.52 | Confirmed | 145 | 32.33 | 32.57 | Confirmed
[7 empty] | 80.15 | 10 | TRUE | Hexachloro-1,3-butadiene | 12.07 | 46240 | 0.22 | 8.015 | Confirmed | 225 | 227 | 63.52 | 62.44 | Confirmed | 223 | 64.42 | 61.88 | Confirmed
[7 empty] | 84.61 | 10 | TRUE | Naphthalene | 12.15 | 276724 | 1.31 | 8.461 | Confirmed | 128 | 127 | 12.82 | 12.46 | Confirmed | 129 | 10.56 | 10.44 | Confirmed
[7 empty] | 76.88 | 10 | FALSE | 1,2,3-Trichlorobenzene | 12.29 | 106079 | 0.5 | 7.688 | Confirmed | 180 | 182 | 95.7 | 94.18 | Confirmed | 145 | 34.15 | 33.64 | Confirmed

## Samples
Instrument Data\GC_MS_PT\2023
231214
Peak Name | Amount
ppb
First Injection | H2O-R1 | H2O-R2 | H2O-R3 | H2O-V1 | H2O-R4 | CCV1 | H2O-V2 | H2O-V3 | MB | 12-4-KDP-1 | 12-6-KDP-1 | 12-8-KDP-1 | 12-8-KDP-2 | 12-8-KDP-3 | 12-8-KDP-4 | 12-8-KDP-5 | 12-8-KDP-6 | 12-8-KDP-7 | 12-8-KDP-8 | 12-8-KDP-9 | 231214-Holmes | H2O-V4 | H2O-R5 | H2O-R6 | CCV2 (next day) | CCV3 (next day)
MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak | MS Quantitation Peak
Chloromethane (methyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.2791 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0641 | 9.6253
Chloroethene (vinyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.6668 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7164 | 9.5687
Bromomethane (methyl bromide) | n.a. | n.a. | n.a. | n.a. | n.a. | 14.9508 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | 10.2627 | 9.9474
Chloroethane (ethyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.8131 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4524 | 9.9176
Trichlorofluoromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.9198 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3821 | 8.8726
Diethyl ether | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9366 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4779 | 8.7447
1,1-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.6704 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5133 | 8.8356
Acetone | 5.2801 | 3.0085 | 2.8259 | n.a./n.r. | 1.6514 | 15.7286 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 4.9246 | 1.9807 | 6.4859 | 5.5471 | n.a./n.r. | 5.8998 | 1.7767 | n.a./n.r. | 5.5208 | 7.6464 | 1.6013 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 14.3446 | 13.7978
Iodomethane | 0.2971 | 0.2528 | 0.2073 | 0.2287 | 0.186 | 7.648 | 0.2543 | 0.2442 | 0.1995 | 0.1889 | 0.2041 | 0.2058 | 0.2153 | 0.165 | 0.1769 | 0.1047 | 0.1831 | 0.1416 | 0.1407 | 0.1337 | 0.16 | 0.1251 | 0.1238 | 0.1595 | 6.6875 | 9.2826
Carbon disulfide | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.1794 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8512 | 9.2295
3-Chloropropene (allyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.4054 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9003 | 8.9084
Methylene chloride (DCM) | 0.1387 | 0.066 | 0.0648 | 0.05 | 0.0543 | 11.4913 | 0.0529 | 0.0437 | 0.0542 | 0.0579 | 0.05 | 0.0599 | 0.0561 | 0.0544 | 0.0486 | 0.0554 | n.a./n.r. | 0.0464 | 0.0543 | 0.058 | 0.0505 | n.a./n.r. | 0.0725 | 0.0816 | 9.3642 | 9.8753
trans-1,2-Dichloroethene | 0.0362 | n.a./n.r. | 0.0143 | n.a./n.r. | n.a./n.r. | 10.4038 | 0.0243 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0245 | n.a./n.r. | 8.6149 | 8.6954
Methyl tert-butyl ether (MTBE) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1585 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.0053 | 7.2661
1,1-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.7088 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7166 | 9.2031
2,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1925 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.1148 | 7.898
cis-1,2-Dichloroethene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.6046 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3492 | 9.1396
2-Butanone (MEK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 15.0557 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.964 | 2.0961 | n.a./n.r. | 2.4563 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 13.1966 | 13.5005
Methyl acrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.1393 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0173 | 8.4593
Bromochloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.7859 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.2638 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7162 | 8.9553
Methacrylonitrile | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.0506 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7283 | 7.8166
Tetrahydrofuran | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.3802 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.9797 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.5852 | 7.3331
Trichloromethane (chloroform) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.4526 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 31.0698 | 18.8322 | 2.716 | 2.193 | n.a./n.r. | 2.5564 | 3.8722 | n.a./n.r. | 2.64 | 3.0004 | 0.2024 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2188 | 9.7203
1,1,1-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0631 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7149 | 8.7162
Dibromofluoromethane [SS1] | 21.2567 | 20.7379 | 19.938 | 21.8437 | 19.7557 | 22.4918 | 19.9339 | 19.7443 | 20.5937 | 20.9037 | 20.4795 | 21.7978 | 20.0018 | 20.0565 | 20.4689 | 19.9419 | 19.8622 | 20.2102 | 19.5677 | 20.6071 | 20.4825 | 20.1309 | 20.3308 | 19.7548 | 20.4652 | 20.6826
Pentafluorobenzene [IS1] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1-Chlorobutane (butyl chloride) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7985 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2332 | 8.5985
Carbon tetrachloride | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5146 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8827 | 8.2662
1,1-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1664 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5529 | 8.9143
Benzene | n.a./n.r. | 0.0153 | 0.011 | n.a./n.r. | n.a./n.r. | 10.1991 | 0.0129 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0754 | 8.9562
1,2-Dichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.851 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9991 | 8.8608
1,4-Difluorobenzene [IS2] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Trichloroethene | 0.0207 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.007 | 10.3556 | 0.0309 | n.a./n.r. | 0.0079 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.5393 | 9.2829
1,2-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1118 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2104 | 8.5162
Dibromomethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.1603 | 0.0137 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.2501 | 9.8472
Methyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.7662 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.1197 | 7.5466
Bromodichloromethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.3425 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 5.1432 | 2.663 | n.a./n.r. | 0.033 | n.a./n.r. | 0.028 | 0.1074 | n.a./n.r. | 0.0292 | 0.0259 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2012 | 8.5772
2-Nitropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9359 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 6.7766 | 6.7554
cis-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9685 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.8196 | 8.763
4-Methyl-2-pentanone (MIBK) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 14.6292 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.8364 | 12.0716
Toluene-d8 [SS2] | 19.0321 | 20.0576 | 19.057 | 19.6212 | 18.9839 | 20.3029 | 19.3611 | 19.405 | 19.0486 | 19.8335 | 19.4862 | 19.3435 | 19.6687 | 20.3093 | 19.3552 | 19.5584 | 19.5091 | 20.2063 | 19.2027 | 18.9432 | 19.1501 | 19.7244 | 20.1501 | 19.1933 | 19.1954 | 19.6214
Toluene | 0.0314 | 0.0209 | 0.0169 | 0.0158 | n.a./n.r. | 10.4421 | 0.0372 | 0.0153 | n.a./n.r. | 0.0595 | 0.0473 | 0.0791 | 0.0551 | n.a./n.r. | 0.0729 | 0.0309 | 0.0218 | 0.097 | 0.0856 | 0.3645 | n.a./n.r. | n.a./n.r. | 0.016 | n.a./n.r. | 8.0326 | 8.5606
trans-1,3-Dichloropropene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.711 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.7014 | 8.4639
Ethyl methacrylate | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.6789 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0241 | 8.2273
1,1,2-Trichloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.7295 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0037 | 8.28
Tetrachloroethene | n.a. | n.a. | n.a. | n.a. | n.a. | 10.156 | n.a. | n.a. | n.a./n.r. | n.a. | n.a./n.r. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a./n.r. | n.a./n.r. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | 7.9362 | 8.5523
1,3-Dichloropropane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.8173 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1838 | 8.2299
2-Hexanone | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 14.6309 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.8224 | 11.7432
Dibromochloromethane | 0.0076 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.028 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 1.4889 | 0.5135 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0757 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.536 | 8.145
1,2-Dibromoethane (EDB) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.9494 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.219 | 8.8914
Chlorobenzene-d5 [IS3] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Chlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.5769 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2235 | 8.8434
1,1,1,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.1425 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0589 | 8.9842
Ethylbenzene | n.a./n.r. | n.a./n.r. | 0.0239 | n.a./n.r. | n.a./n.r. | 11.2229 | 0.0741 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0327 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8357 | 9.4466
m/p-Xylene | 0.3046 | 0.2959 | 0.287 | 0.2917 | 0.2793 | 10.2002 | 0.3345 | 0.2937 | 0.2883 | 0.2921 | 0.2776 | 0.2786 | 0.2764 | n.a./n.r. | 0.2764 | 0.2696 | n.a./n.r. | 0.2754 | 0.2764 | 0.271 | n.a./n.r. | 0.2801 | 0.2756 | 0.2712 | 7.9512 | 8.6559
o-Xylene | 0.0267 | 0.019 | 0.0161 | n.a./n.r. | n.a./n.r. | 11.1807 | 0.0582 | 0.022 | 0.0237 | 0.0263 | n.a./n.r. | n.a./n.r. | 0.0086 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0176 | n.a./n.r. | n.a./n.r. | 8.7795 | 9.2761
Styrene | 0.228 | 0.2216 | 0.2105 | 0.2123 | 0.2038 | 10.9532 | 0.2504 | 0.2143 | 0.2036 | 0.2089 | 0.205 | n.a./n.r. | 0.1982 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.3108 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.4475 | 9.2205
Bromoform | n.a./n.r. | 0.0073 | 0.0049 | 0.0054 | n.a./n.r. | 9.9694 | 0.025 | 0.005 | n.a./n.r. | 0.2289 | 0.0525 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0015 | 0.0021 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.6677 | n.a./n.r. | 0.0017 | n.a./n.r. | 7.7933 | 8.3896
Isopropylbenzene (cumene) | 0.0458 | 0.0238 | 0.0143 | n.a./n.r. | n.a./n.r. | 10.8749 | 0.062 | 0.0198 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9143 | 8.5638
1-Bromo-4-fluorobenzene (BFB) [SS3] | 20.6693 | 19.9318 | 19.759 | 20.3532 | 19.8879 | 20.2601 | 19.8329 | 20.2378 | 19.5134 | 19.613 | 19.6734 | 19.8081 | 19.5131 | 21.5049 | 19.7748 | 19.9346 | 20.2217 | 20.8662 | 19.0739 | 19.1582 | 19.9349 | 18.5144 | 20.1977 | 20.4781 | 19.4321 | 19.0564
Bromobenzene | 0.0482 | 0.0376 | 0.0271 | 0.0256 | 0.0157 | 10.5684 | 0.0797 | 0.0292 | 0.0173 | 0.0132 | 0.0112 | 0.011 | 0.0094 | 0.0074 | 0.0072 | 0.0071 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.006 | n.a./n.r. | 0.0128 | 0.0076 | 7.8775 | 8.0648
1,1,2,2-Tetrachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.0441 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.6276 | 7.7766
1,2,3-Trichloropropane (TCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5782 | 0.0333 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.327 | 7.5336
trans-1,4-Dichloro-2-butene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.4901 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 6.9954 | 7.3061
n-Propylbenzene | 0.0799 | 0.0534 | 0.035 | 0.027 | 0.0247 | 11.4054 | 0.0971 | 0.0379 | n.a./n.r. | n.a./n.r. | 0.0153 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.347 | 8.7499
2-Chlorotoluene | 0.0562 | 0.0397 | 0.0259 | 0.0209 | 0.0177 | 11.5246 | 0.081 | 0.0285 | 0.014 | 0.0129 | 0.0114 | 0.0077 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0095 | n.a./n.r. | 8.2725 | 8.8061
1,3,5-Trimethylbenzene | 0.0472 | 0.0302 | 0.0197 | 0.0168 | n.a./n.r. | 10.7342 | 0.0646 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.894 | 8.5184
4-Chlorotoluene | 0.073 | 0.0566 | 0.0401 | 0.031 | 0.0251 | 11.7336 | 0.0993 | 0.0388 | 0.0213 | 0.0208 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0086 | 0.0095 | n.a./n.r. | n.a./n.r. | n.a./n.r. | 0.0066 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.6503 | 9.2362
tert-Butylbenzene | 0.3795 | 0.3619 | 0.3551 | n.a./n.r. | 0.3517 | 10.9253 | 0.397 | 0.3578 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.054 | 8.3867
Pentachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 12.0244 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.3213 | 10.1665
1,2,4-Trimethylbenzene | 0.0565 | 0.0403 | 0.0243 | n.a./n.r. | n.a./n.r. | 11.2477 | 0.0825 | 0.0249 | n.a./n.r. | 0.0177 | n.a./n.r. | 0.0142 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.2475 | 8.8936
1-Methylpropylbenzene (sec-butylbenzene) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.1625 | 0.0876 | n.a./n.r. | n.a./n.r. | 0.013 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.1043 | 8.6906
1,3-Dichlorobenzene | 0.0945 | 0.0759 | 0.0532 | 0.0437 | 0.0359 | 11.5022 | 0.1229 | 0.0553 | 0.0343 | 0.0252 | 0.0241 | 0.0176 | 0.0151 | 0.0117 | 0.0142 | 0.0106 | 0.0123 | n.a./n.r. | 0.0089 | 0.0079 | 0.0094 | 0.007 | 0.0173 | 0.0123 | 8.454 | 9.1287
4-Isopropyltoluene (p-cymene) | 0.0848 | 0.0542 | 0.0282 | 0.0226 | 0.0169 | 11.3451 | 0.0865 | 0.0306 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0871 | 9.2445
1,4-Dichlorobenzene-d4 [IS4] | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
1,4-Dichlorobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 11.8007 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.8021 | 9.3429
n-Butylbenzene | 0.128 | 0.0921 | 0.062 | 0.0405 | 0.0387 | 10.766 | 0.1175 | 0.0497 | 0.028 | 0.0243 | 0.0239 | 0.0172 | 0.0159 | 0.0122 | 0.0108 | 0.0101 | 0.0096 | 0.01 | 0.0066 | 0.0074 | 0.0059 | 0.0076 | 0.0156 | 0.0099 | 7.3534 | 8.3225
1,2-Dichlorobenzene | 0.0858 | 0.0681 | 0.0464 | 0.0391 | 0.0299 | 11.736 | 0.1165 | 0.0464 | 0.0259 | 0.0228 | 0.0221 | 0.0173 | 0.0171 | 0.0121 | 0.0138 | 0.0112 | 0.0122 | 0.0113 | 0.0107 | 0.0089 | 0.0104 | 0.0089 | 0.0194 | 0.0139 | 8.4968 | 9.2742
Hexachloroethane | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 10.7315 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.9149 | 8.4021
1,2-Dibromo-3-chloropropane (DBCP) | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 9.5236 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 7.5727 | 7.8281
Nitrobenzene | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.796 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 6.47 | 6.6153
1,2,4-Trichlorobenzene | 0.1424 | 0.1303 | 0.0763 | 0.0642 | 0.0502 | 10.7014 | 0.1586 | 0.0796 | 0.0467 | 0.0431 | 0.032 | 0.0263 | 0.0246 | 0.0169 | 0.0202 | 0.0185 | 0.0208 | 0.0116 | 0.0135 | 0.0142 | 0.0144 | 0.0107 | 0.0202 | 0.0135 | 7.9614 | 8.66
Hexachloro-1,3-butadiene | 0.1205 | 0.0811 | 0.0488 | 0.0351 | 0.0274 | 9.856 | 0.0941 | 0.0496 | 0.0206 | 0.0209 | 0.0202 | 0.0103 | 0.0105 | 0.008 | 0.0081 | n.a./n.r. | 0.0069 | 0.0094 | 0.0065 | 0.0073 | 0.0057 | n.a./n.r. | 0.0154 | 0.0075 | 7.0091 | 8.0152
Naphthalene | 0.2415 | 0.2336 | 0.2178 | 0.2072 | 0.2009 | 10.176 | 0.3059 | 0.2187 | 0.1969 | 0.198 | 0.1945 | 0.1961 | 0.1951 | 0.1847 | 0.1911 | 0.1843 | n.a./n.r. | 0.1892 | 0.1894 | 0.1833 | n.a./n.r. | n.a./n.r. | n.a./n.r. | n.a./n.r. | 8.0947 | 8.4611
1,2,3-Trichlorobenzene | 0.0982 | 0.0874 | 0.0539 | 0.0441 | 0.035 | 9.618 | 0.1411 | 0.0554 | 0.0334 | 0.0261 | 0.0237 | 0.0178 | 0.0179 | n.a./n.r. | 0.0114 | 0.0106 | 0.01 | 0.0078 | 0.0093 | n.a./n.r. | n.a./n.r. | 0.0074 | 0.0117 | 0.0094 | 7.1279 | 7.688

## Tent
| Peak | Ret.Time | 1st Hit | Library Compound | Library | 2nd Hit | LIbrary Compound | Library | 3rd Hit | Library Compound | Library | Area | Estimated | 20 ppb (TIC) | RT | Area
| No. | min | SI |  |  | SI |  |  | SI |  |  | counts*min | Concentration | Pentafluorobenzene [IS1] | 5.43 | 568124
Sample | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC | TIC |  | (ppb) | 1,4-Difluorobenzene [IS2]
12-8-KDP-9 | 5 | 3.41 | 931 | Cyclopentane | mainlib | 903 | 1-Pentene | mainlib | 873 | 2-Pentene, (E)- | mainlib | 838307.461 | 29.5114 | Chlorobenzene-d5 [IS3]
1,4-Dichlorobenzene-d4 [IS4]